## Barco
1 | [11 empty] | Revision 1.6-20220607 | [10 empty] | Print Area |  |  |  |  | Test frequency
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | Page1,Page2 |  |  |  |  | Standard monitor |  | Tomosynthesis monitor
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina | [10 empty] | No QAWS | QAWS | No QAWS | QAWS
4 | [22 empty] | All: | Page1,Page2 |  |  | Autocalibration | Semiannual | Semiannual | Semiannual | Semiannual
5 | [5 empty] | Mammography Monitor Compliance Inspection | [10 empty] | Mammography Monitor Compliance Inspection | [9 empty] | Compliance Test | Weekly | Weekly | Weekly | Monthly
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year | Display Test |  |  | NA | Weekly
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  | Change | Page1,Page2 | Rapid Frame | NA | NA | Quarterly | Quarterly
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [5 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah
10 |  | Facility: |  |  |  |  | Site Number: | [6 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility:
11 |  | Department: |  |  |  |  | Date of Installation: | [6 empty] | Department: |  |  |  | Date of Installation: |  |  |  |  | Department:
12 |  | Area/Division: |  |  |  |  | Workstation ID: | [6 empty] | Area/Division: |  |  |  | Workstation ID: |  |  |  |  | Area/Division:
13 |  | Survey ID: |  |  |  |  | Left monitor SN: | [6 empty] | Survey ID: |  |  |  | Left monitor SN: |  |  |  |  | Survey ID:
14 |  | Room Number: |  |  |  |  | Right monitor SN: | [6 empty] | Room Number: |  |  |  | Right monitor SN: |  |  |  |  | Room Number:
15 |  | QC Manual version: | Barco K5905277 v16 | [10 empty] | QC Manual version: | Barco K5905277 v16 | Barco K5905277 v16 | [6 empty] | QC Manual version: |  | Change | Barco K5905277 v16
16 | [22 empty] | Site Number:
17 | [10 empty] | Medical University of South Carolina | [5 empty] | Inspection Results | [5 empty] | Date of Installation:
18 | [10 empty] | Mammography Monitor Compliance Inspection | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Workstation ID:
19 |  |  |  |  | Comments |  |  |  | Test Equipment | [5 empty] | Comments |  |  |  | Test Equipment |  |  |  | Left monitor SN:
20 |  | Monitor surface is clean: | [6 empty] | Meter: | Piranha CB2-19020491 |  |  | Monitor surface is clean | [5 empty] | Meter: | Piranha CB2-19020491 | Piranha CB2-19020491 |  | Right monitor SN:
21 |  | Viewing conditions: | [6 empty] | Calibration Date: | 2023-06-12 |  |  | Viewing conditions | [5 empty] | Calibration Date: | 2023-06-12 | 2023-06-12
22 |  | TG18-QC test pattern: | [6 empty] | Calibration Due: | 2025-06-12 |  |  | TG18-QC test pattern | [5 empty] | Calibration Due: | 2025-06-12 | 2025-06-12 |  | Test Equipment
23 |  | Rapid Frame acceptable: | [10 empty] | Rapid Frame acceptable | [9 empty] | Meter: |  | Change | Piranha CB2-19020491
24 |  | Display Test: | [20 empty] | Calibration Date: |  | Change | 2023-06-12
25 | [11 empty] | Room Conditions | [5 empty] | Frequency: | Monthly |  |  |  | Calibration Due: |  | Change | 2025-06-12
26 | Room Conditions | [10 empty] | Ambient Illum: |  | lux |  | Amb Luminance: |  | cd/m^2
27 | Ambient Illum: |  | lux |  | Criteria: | Ambient light measured at the surface of the monitor(s) must be less than 20 lux | [5 empty] | Previous: |  | lux |  | Previous: |  | cd/m^2 |  |  |  |  | Ambient Illum:
28 | Ambient Lum |  | cd/m^2 | [9 empty] | Criteria: | Ambient light illuminance measured at the surface of the monitor(s) must be less than 20 lux | [8 empty] | Ambient Lum:
29 | Comments: | [11 empty] | Comments: | [9 empty] | Auto Cal Lum Ratio L:
30 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | Auto Cal Lum Ratio R:
31 | [12 empty] | New: | [9 empty] | TG-18 LN01 L:
32 | [22 empty] | TG-18 LN01 R:
33 | Auto Calibration | [10 empty] | Auto Calibration | [5 empty] | Frequency: | Semi-annually |  |  |  | TG-18 LN18 L:
34 | Date: |  |  | Lmin | Lmax | Luminance Ratio | [5 empty] | Date: |  |  | Left () | Right () | [6 empty] | TG-18 LN18 R:
35 |  |  | Left () | [10 empty] | Lmin (cd/m^2) | [8 empty] | TG-18 Lum Ratio L:
36 |  |  | Right () | [10 empty] | Lmax (cd/m^2) | [8 empty] | TG-18 Lum Ratio R:
37 | Comments: | [12 empty] | Luminance Ratio | [8 empty] | Ambience Ratio L:
38 | [22 empty] | Ambience Ratio R:
39 | [12 empty] | Comments: | [9 empty] | Uniformity Dev UNL10 L:
40 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | Uniformity Dev UNL10 R:
41 | Compliance Test | [11 empty] | New: | [9 empty] | Uniformity Dev UNL80 L:
42 | Date: |  |  | Left () | Right () | [17 empty] | Uniformity Dev UNL80 R:
43 |  |  | DICOM GSDF curve acceptable: | [8 empty] | Compliance Test | [5 empty] | Frequency: | Monthly (if using QA Web Server) |  |  |  | LUDM UNL 10 L:
44 | Comments: | [10 empty] | Date: | [6 empty] | Weekly (if not using QA Web Server) |  |  |  | LUDM UNL 10 R:
45 | [11 empty] | Enter 1 for YES, 2 for NO, 3 for NA |  |  | Left () | Right () | [6 empty] | LUDM UNL 18 L:
46 | [13 empty] | DICOM GSDF curve acceptable: | [8 empty] | LUDM UNL 18 R:
47 | [12 empty] | Comments: | [9 empty] | QUALITY CONTROL visibility:
48 | TG-18 Test Patterns | [11 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | Black L:
49 | TG-18 LN01 | [5 empty] | TG-18 LN18 | [5 empty] | New: | [9 empty] | Gray L:
50 |  | Left () |  | cd/m^2 |  |  | Left () |  | cd/m^2 | [13 empty] | White L:
51 |  | Right () |  | cd/m^2 |  |  | Right () |  | cd/m^2 |  |  | Display Test | [5 empty] | Frequency: | Weekly (if using QA Web Server) |  |  |  | Black R:
52 | [11 empty] | Date: | [6 empty] | Not available if not using QA Web Server |  |  |  | Gray R:
53 |  |  | Left () | Right () | [8 empty] | Comments: | [9 empty] | White R:
54 |  | Luminance Ratio: | [10 empty] | Do not exceed 390 chars |  | 0
55 |  | Ambience Ratio: | [10 empty] | New: | [9 empty] | Comments
56 | [22 empty] | Room Conditions:
57 | Comments: | [10 empty] | TG-18 Test Patterns | [10 empty] | Auto Calibration:
58 | [11 empty] | TG-18 Black/White |  | Left () | Right () |  | Previous | Left () | Right () |  |  |  | Compliance Test:
59 | [12 empty] | LN-01 (cd/m^2): |  |  |  | LN-01 (cd/m^2): | [5 empty] | Display Test:
60 | [12 empty] | LN-18 (cd/m^2): |  |  |  | LN-18 (cd/m^2): | [5 empty] | TG-18 Black/White:
61 | [12 empty] | Luminance ratio: |  |  |  | Luminance ratio: | [5 empty] | TG-18 QC:
62 | [12 empty] | Ambience ratio: |  |  |  | Ambience ratio: | [5 empty] | TG-18 Uniformity:
63 | [12 empty] | Comments:
64 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | Comments
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | New:
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina | TG-18 QC | [5 empty] | Frequency: | Annually
2 | [10 empty] | Mammography Monitor Compliance Inspection | Enter 1 for YES, 2 for NO, 3 for NA |  |  | Left () | Right ()
3 | [13 empty] | 5%/95% patches visible:
4 | TG-18 Uniformity | [12 empty] | Resolution acceptable:
5 |  | UNL-10 (cd/m^2) |  | UNL-80 (cd/m^2) |  |  | Visual Assessment | [6 empty] | Display is free of artifacts/noise:
6 |  | Left () | Right () | Left () | Right () |  | TBD | Free from gross non-uniformities | [5 empty] | QUALITY CONTROL visibility: | Enter as much of “QUALITY CONTROL” that is visible |  |  |  | Previous Year
7 | Average: | [5 empty] | TBD | Free from non-symmetric non-uniformities | [5 empty] | Black
8 | SD: | [5 empty] | TBD | Free from significant luminance variation
9 | Deviation (%) | [12 empty] | Gray
10 | LUDM (%): | [5 empty] | Criteria: | Luminance deviation is < 30%
11 | [7 empty] | Luminance deviation from medium (LUDM) is < 30% | [5 empty] | White
12
13 | Comments: | [12 empty] | Criteria: | There should be no significant changes to how much of QUALITY CONTROL is visible
14
15 | [12 empty] | Comments:
16 | [12 empty] | Do not exceed 390 chars |  | 0
17 | TG-18 QC | [5 empty] | QUALITY CONTROL visibility | [5 empty] | New:
18 |  |  |  | Left () | Right () |  | Left () |  | Right ()
19 |  |  | 5%/95% patches visible: |  |  | Black | [5 empty] | TG-18 Uniformity | [6 empty] | Visual assessment
20 |  |  | Resolution acceptable: |  |  | Gray | [7 empty] | UNL-10 (cd/m^2) |  | UNL-80 (cd/m^2) |  |  | Enter 1 for YES, 2 for NO, 3 for NA
21 |  |  | Display is free of artifacts/noise: |  |  | White | [7 empty] | Left () | Right () | Left () | Right () |  |  | Free from gross non-uniformities |  |  | LUDM UNL 10 |  | LUDM UNL 80
22 | Comments: | [11 empty] | Region 1: | [6 empty] | Free from non-symmetric non-uniformities |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
23 | [12 empty] | Region 2: | [6 empty] | Free from significant luminance variation |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
24 | [12 empty] | Region 3: | [9 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE!
25 | [12 empty] | Region 4: | [9 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE!
26 | Comments: | [11 empty] | Region 5: | [6 empty] | Average | SD |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
27 | [12 empty] | Region 6: | [5 empty] | UNL-10 (cd/m^2) | #DIV/0! | #DIV/0! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
28 | [12 empty] | Region 7: | [5 empty] | UNL-80 (cd/m^2) | #DIV/0! | #DIV/0! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
29 | [12 empty] | Region 8: | [9 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE!
30 | [12 empty] | Region 9: | [9 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE!
31 | [12 empty] | Average (cd/m^2):
32 | [12 empty] | SD (cd/m^2):
33 | [12 empty] | Luminance Deviation (%): |  |  |  |  | Criteria: | Luminance deviation is < 30%
34 | [12 empty] | Lum Dev from Median (%): | [5 empty] | Luminance deviation from median is < 30%
35 | [12 empty] | Comments:
36 | [12 empty] | Do not exceed 390 chars |  | 0
37 | [12 empty] | New:
38
39
40 | [16 empty] | Comments Page 1
41 | [15 empty] | Previous Year Comments
42 | [12 empty] | Comments:
43 | [12 empty] | Do not exceed 390 chars |  | 0
44 | [12 empty] | Additional Comments:
45 | [12 empty] | Do not exceed 390 chars |  | 0
46 | [12 empty] | Additional Comments:
47 | [12 empty] | Do not exceed 390 chars |  | 0
48 | [12 empty] | Additional Comments:
49 | [12 empty] | Do not exceed 390 chars |  | 0
50 | [12 empty] | Additional Comments:
51 | [12 empty] | Do not exceed 390 chars |  | 0
52 | [12 empty] | Additional Comments:
53 | [12 empty] | Do not exceed 390 chars |  | 0
54 | [12 empty] | Additional Comments:
55 | [12 empty] | Do not exceed 390 chars |  | 0
56 | [12 empty] | Additional Comments:
57 | [12 empty] | Do not exceed 390 chars |  | 0
58 | [12 empty] | Additional Comments:
59 | [12 empty] | Do not exceed 390 chars |  | 0
60 | [12 empty] | Additional Comments:
61 | [12 empty] | Do not exceed 390 chars |  | 0
62 | [12 empty] | Additional Comments:
63 | [12 empty] | Do not exceed 390 chars |  | 0
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
Note on Q59: Eugene Mah: Total Luminance = Measured luminance + ambient luminance
Note on Q88: Eugene Mah: Total Luminance = Measured luminance + ambient luminance

## Sheet1
Measurements
Test Pattern | Light (cd/m²)
Lamb
LN01
LN01
LN18
LN18
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN10
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
UN80
Waveforms