## Sheet1
1 | [9 empty] | Revision 001-20260415 (Has 48 data records) | [10 empty] | PRINT AREAS
2 |  |  |  |  | Medical University of South Carolina | [8 empty] | Medical University of South Carolina | [6 empty] | FIRST,EXP,BLD,KV_REPRO,LIN_HVL,AEC_KAP,LAST
3 |  |  |  |  | Charleston, South Carolina | [8 empty] | Charleston, South Carolina | [6 empty] | ALL: | FIRST,COMPG2,COMPG3,EXP,BLD,KV_REPRO,LIN_HVL,AEC_KAP,LAST
4 | [22 empty] | REM: | (1) Go to last year's spreadsheet (see "**Note", below) and copy from ITS "THIS YEAR" range
5 |  |  |  |  | EOS Imaging System Compliance Inspection | [7 empty] | EOS Imaging System Compliance Inspection | [10 empty] | (2) Come back to this sheet and paste (use paste special, "values") into "LAST YEAR" range
6 | [23 empty] | LAST | Cell AD6 | THIS
7 | [22 empty] | Measurement Parameter | YEAR |  | YEAR | (Locked)
8 |  |  |  |  | System Information | [5 empty] | Date: |  |  |  | System Information |  |  | Inspector: | Eugene Mah |  |  |  | Print Area: |  | Change | FIRST,EXP,BLD,KV_REPRO,LIN_HVL,AEC_KAP,LAST | X2
9 | [10 empty] | Last Year: | [6 empty] | Last Year: |  |  |  |  | Date: |  |  |  | N8
10 |  | Location | [8 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  |  | Inspector: |  | Change | Eugene Mah | U8
11 |  |  | Facility: |  |  |  | Site: |  |  |  |  | Facility: |  |  |  | Site Number: | [6 empty] | Facility: |  |  |  | O11
12 |  |  | Department: |  |  |  | Authorized Use: |  |  |  |  | Department: |  |  |  | Authorized Use: | [6 empty] | Department: |  |  |  | O12
13 |  |  | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division: |  |  |  | Date of Installation: | [6 empty] | Area/Division: |  |  |  | O13
14 |  |  | Survey ID: |  |  |  | Number of X-ray Tubes: | 2 |  |  |  | Survey ID: |  |  |  | Number of X-ray Tubes: | 2 | 2 |  |  |  |  | Survey ID: |  |  |  | O14
15 | [11 empty] | Room Number: |  |  |  | Accession Number: | [6 empty] | Room Number: |  |  |  | O15
16
17 |  | X-Ray Generator | [8 empty] | X-Ray Generator | [11 empty] | X-Ray Generator
18 |  |  | Manufacturer: |  |  |  | Date of Manufacture: |  |  |  |  | Manufacturer: |  |  |  | Date of Manufacture: | [6 empty] | Manufacturer: |  |  |  | O18
19 |  |  | Model: |  |  |  | Serial Number: |  |  |  |  | Model: |  |  |  | Serial Number: | [6 empty] | Model: |  |  |  | O19
20 |  |  | Max. kVp: |  |  |  | Max. mA: |  |  |  |  | Max. kVp: |  |  |  | Max. mA: | [6 empty] | Max. kVp: |  |  |  | O20
21 |  |  | Software Version: | [8 empty] | Software Version: | [10 empty] | Software Version: |  |  |  | O21
22
23 |  | X-Ray Tube 1 (Frontal) | [5 empty] | Focal Spot Sizes |  |  | X-Ray Tube 1 (Frontal) | [5 empty] | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube 1
24 |  |  | Tube Designation/Use: |  |  |  | Large: |  |  |  |  | Tube Designation/Use: |  |  |  | Large: | [6 empty] | Tube Designation/Use: |  |  |  | O24
25 |  |  | Date of Manufacture: |  |  |  | Small: |  |  |  |  | Date of Manufacture: |  |  |  | Small: | [6 empty] | Date of Manufacture: |  |  |  | O25
26 | [6 empty] | Micro: | [8 empty] | Micro:
27 |  |  | Insert Manufacturer: |  |  |  |  | Filtration |  |  |  | Insert Manufacturer: |  |  |  |  | Filtration | [5 empty] | Insert Manufacturer: |  |  |  | O27
28 |  |  | Insert Model: |  |  |  | Inherent: |  |  |  |  | Insert Model: |  |  |  | Inherent: | [6 empty] | Insert Model: |  |  |  | O28
29 |  |  | Insert Serial Number: |  |  |  | Added: |  |  |  |  | Insert Serial Number: |  |  |  | Added: | [6 empty] | Insert Serial Number: |  |  |  | O29
30
31 |  |  | Housing Manufacturer: | [8 empty] | Housing Manufacturer: | [10 empty] | Housing Manufacturer: |  |  |  | O31
32 |  |  | Housing Model: | [8 empty] | Housing Model: | [10 empty] | Housing Model: |  |  |  | O32
33 |  |  | Housing Serial Number: | [8 empty] | Housing Serial Number: | [10 empty] | Housing Serial Number: |  |  |  | O33
34
35
36 |  | X-Ray Tube 2 (Lateral) | [5 empty] | Focal Spot Sizes |  |  | X-Ray Tube 2 (Lateral) | [5 empty] | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube 2
37 |  |  | Tube Designation/Use: |  |  |  | Large: |  |  |  |  | Tube Designation/Use: |  |  |  | Large: | [6 empty] | Tube Designation/Use: |  |  |  | O37
38 |  |  | Date of Manufacture: |  |  |  | Small: |  |  |  |  | Date of Manufacture: |  |  |  | Small: | [6 empty] | Date of Manufacture: |  |  |  | O38
39 | [6 empty] | Micro: | [8 empty] | Micro:
40 |  |  | Insert Manufacturer: |  |  |  |  | Filtration |  |  |  | Insert Manufacturer: |  |  |  |  | Filtration | [5 empty] | Insert Manufacturer: |  |  |  | O40
41 |  |  | Insert Model: |  |  |  | Inherent: |  |  |  |  | Insert Model: |  |  |  | Inherent: | [6 empty] | Insert Model: |  |  |  | O41
42 |  |  | Insert Serial Number: |  |  |  | Added: |  |  |  |  | Insert Serial Number: |  |  |  | Added: | [6 empty] | Insert Serial Number: |  |  |  | O42
43
44 |  |  | Housing Manufacturer: | [8 empty] | Housing Manufacturer: | [10 empty] | Housing Manufacturer: |  |  |  | O44
45 |  |  | Housing Model: | [8 empty] | Housing Model: | [10 empty] | Housing Model: |  |  |  | O45
46 |  |  | Housing Serial Number: | [8 empty] | Housing Serial Number: | [10 empty] | Housing Serial Number: |  |  |  | O46
47
48 | Date: | [6 empty] | Inspector: | Eugene Mah | [13 empty] | Site Number: |  |  |  | S11
49 | Room Number: | [6 empty] | Survey ID: | [6 empty] | Inspection Results |  |  |  |  | Medical University of South Carolina |  |  | Authorized Use: |  |  |  | S12
1 | [8 empty] | Medical University of South Carolina | [10 empty] | EOS Imaging System Compliance Inspection |  |  | Date of Installation: |  |  |  | S13
2 | [8 empty] | EOS Imaging System Compliance Inspection | Enter 1 for YES, 2 for NO, 3 for NA | [12 empty] | Number of X-ray Tubes: |  | Change | 2 | S14
3 | [9 empty] | 2 | 2 | Is this a new installation? | [10 empty] | Accession Number: |  |  |  | S15
4 | Rule Number |  |  |  | Labels, Notices, Postings |  |  |  | Compliance
5 | SCDES RHB 2.5.1.1 |  | SCDES registration sticker is present, clearly visible and legible. | [10 empty] | Labels, Notices, Postings | [8 empty] | Date of Manufacture: |  |  |  | S18
6 | SCDES RHB 11.2.3 |  | SCDES form SC-RHA-20 "Notice to Employees" posted or referenced. | [8 empty] | SCDES registration sticker is present, clearly visible and legible. | [10 empty] | Serial Number: |  |  |  | S19
7 | SCDES RHB 4.2.7 |  | Pregnancy Warning sign is posted | [8 empty] | SCDES form SC-RHA-20 "Notice to Employees" posted or referenced. | [10 empty] | Max. mA: |  |  |  | S20
8 | SCDES RHB 4.3.1 |  | Radiation warning label present on the generator control panel. | [8 empty] | Pregnancy Warning sign is posted
9 | SCDES RHB 4.2.6 |  | Technique chart is present and accurate. |  |  |  | DATE: |  |  |  |  | Radiation warning label present on the generator control panel. | [10 empty] | Focal Spot Sizes
10 | JCAHO |  | Patient exposure/dose measurements are posted. |  |  |  | DATE: |  |  |  | DATE: |  | Technique chart present. | [9 empty] | Large: |  |  |  | S24
11 |  |  | Operators manuals are available. | [7 empty] | DATE: |  | Patient exposure/dose measurements posted. | [9 empty] | Small: |  |  |  | S25
12 |  |  | Monthly radiation monitoring reports are posted. | [8 empty] | Operators manuals are available. | [10 empty] | Micro: |  |  |  | S26
13 | [11 empty] | Monthly radiation monitoring reports are posted. | [10 empty] | Filtration
14 |  |  |  |  | Beam Limiting Device | [17 empty] | Inherent: |  |  |  | S28
15 | SCDES RHB 4.7.1.1 |  | A beam limiting device is provided to restrict the size of the useful beam. | [19 empty] | Added: |  |  |  | S29
16 | SCDES RHB 4.3.6 |  | Positive means is provided to ensure the minimum filtration is in beam.
17 | SCDES RHB 4.7.14 |  | Minimum field size is <= 5cm on each side. | [5 empty] | TBD | [13 empty] | Focal Spot Sizes
18 | SCDES RHB 4.3.5 |  | The measured HVL meets SCDES minimum requirements. | [5 empty] | TBD | [13 empty] | Large: |  |  |  | S37
19 |  |  | X-ray beam is level | [5 empty] | TBD | [5 empty] | COMMENTS | [7 empty] | Small: |  |  |  | S38
20 | SCDES RHB 4.7.1.5 |  | The indicated field size is accurate to within 2% of the SID. | [5 empty] | NA |  | Note: Type all comments in the shaded cells. | [11 empty] | Micro: |  |  |  | S39
21 | [13 empty] | Previous Year Comments:
22 |  |  |  |  | Generator | [5 empty] | Comments: | [11 empty] | Filtration
23 | SCDES RHB 4.3.9.1 |  | Means provided to indicate technique factors prior to exposure. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **The SCDES registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag. | [6 empty] | Inherent: |  |  |  | S41
24 |  |  | All technique indicators functional and are free of excessive wear. | [7 empty] | Additional comments: | [11 empty] | Added: |  |  |  | S42
25 |  |  | Technique overload circuits function to block excessive technique | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **A technique chart was not posted in the room. A technique chart reflecting techniques currently used with the CR system must be provided (SCDES RHB 4.2.8).
26 |  |  | Anode heat calculator/indicator is functioning properly | [7 empty] | Additional comments: | [11 empty] | New install? |  | Change | 2 | M52
27 | SCDES RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
28 | SCDES RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational | [7 empty] | Additional comments: | [11 empty] | COMMENT PAGE 1 |  |  | 0 | N71
29 | SCDES RHB 4.7.4.2.3 |  | Audible exposure end indicator present and operational | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Tube loading charts for each focal spot were not available at the time of inspection. These charts should be obtained and posted near the operator's console.
30 |  |  | Exposure delay time is not excessive. | [7 empty] | Additional comments: | [14 empty] | 0 | N72
31 | SCDES RHB 4.7.6 |  | kVp accurate to within 5% of indicated. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Operator's manuals for the imaging equipment were not available at the time of inspection. These should be obtained and made available to the operator of the equipment.
32 | SCDES RHB 4.7.5 |  | kVp and exposure reproducibility coeff. of var. < 5%. | [5 empty] | NA |  | Additional comments: | [14 empty] | 0 | N73
33 | SCDES RHB 4.7.7 |  | Exposure coefficient of linearity < 10%. | [5 empty] | NA |  | Do NOT exceed 390 |  | 0 |  | 0 | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
34 | [10 empty] | Additional comments: | [14 empty] | 0 | N74
35 |  |  |  |  | X-ray Tube | [5 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **The light field illumination failed to meet the minimum specification of 160 lux. The light source should be adjusted or replaced so as to meet this requirement.
36 |  |  | Focal spot resolution is within acceptable parameters | [5 empty] | NA | [16 empty] | 0 | N75
37 |  |  | The x-ray output @100 cm within +/-25% from previous measurements. | [5 empty] | NA
38 |  |  | Kerma area product accuracy is within +/-15% | [5 empty] | TBD | [16 empty] | 0 | N76
39 |  |  | AEC reproducibility deviation is less than 10% from the mean | [5 empty] | TBD
40 | [25 empty] | 0 | N77
41
42 | [13 empty] | DO NOT USE THIS SPACE | [8 empty] | COMMENT PAGE 2 |  |  | 0 | N106
43
44 | [25 empty] | 0 | N107
45
46 | [25 empty] | 0 | N108
47
48 | Date: | [6 empty] | Inspector: | Eugene Mah | [16 empty] | 0 | N109
49 | Room Number: | [6 empty] | Survey ID:
1 | [8 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | 0 | N110
2 | [8 empty] | EOS Imaging System Compliance Inspection | [10 empty] | EOS Imaging System Compliance Inspection
3 | [14 empty] | COMMENTS PAGE 2 | [10 empty] | 0 | N111
4 |  |  |  |  | Radiation Safety |  |  |  | Compliance |  |  |  |  | Previous Year Comments:
5 | SCDES RHB 4.7.4.2.1 |  | Continuous pressure on exposure switch required to make an exposure. | [7 empty] | Comments: | [14 empty] | 0 | N112
6 | SCDES RHB 4.7.4.2.2 |  | Exposure switch for selected tube is secured at a safe location | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
7 | SCDES RHB 4.7.4.2.4 |  | Operator has full visibility of the exposure factors and patient. | [7 empty] | Additional comments: | [14 empty] | 0 | N113
8 | SCDES RHB 4.2.12.7 |  | Lead aprons available. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
9 | SCDES RHB 4.2.12.7 |  | Lead gloves available. | [7 empty] | Additional comments: | [14 empty] | 0 | N114
10 |  |  | A properly designed and installed apron rack is present. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
11 | SCDES RHB 4.2.8 |  | Documentation of annual protective apparel integrity inspection is available. | [7 empty] | Additional comments: | [14 empty] | 0 | N115
12 |  |  | Patient restraint devices available. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
13 | SCDES RHB 4.2.12.7 |  | Gonadal shielding is present and functional. | [7 empty] | Additional comments: | [14 empty] | 0 | N116
14 |  |  | Radiation warning light at room entrance is functional with selected tube | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
15 | SCDES RHB 4.2.13.2 |  | Entrance skin exposures are within limits | [5 empty] | TBD |  | Additional comments: | [14 empty] | 0 | N117
16 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
17 |  |  |  |  | Physical Inspection | [5 empty] | Additional comments: | [14 empty] | 0 | N118
18 |  |  | System clean and free of debris. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
19 |  |  | All cables free from excessive wear or restraints. | [7 empty] | Additional comments: | [14 empty] | 0 | N119
20 |  |  | System free of missing or damaged components. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
21 |  |  | Table movements and locks acceptable | [7 empty] | Additional comments: | [14 empty] | 0 | N120
22 |  |  | Tube crane movements and locks acceptable. | [7 empty] | Do NOT exceed 390 |  | 0 |  | 0
23 |  |  | All other movements and locks functional. | [7 empty] | Additional comments: | [11 empty] | COMMENT PAGE 3 |  |  | 0 | N172
24 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
25 | Comments: | [9 empty] | Additional comments: | [14 empty] | 0 | N173
26 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
27 | [10 empty] | Additional comments: | [14 empty] | 0 | N174
28 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
29 | [10 empty] | Additional comments: | [14 empty] | 0 | N175
30 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
31 | [10 empty] | Additional comments: | [14 empty] | 0 | N176
32 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
33 | [10 empty] | Additional comments: | [14 empty] | 0 | N177
34 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
35 | [25 empty] | 0 | N178
36
37 | [25 empty] | 0 | N179
38
39 | [25 empty] | 0 | N180
40
41 | [25 empty] | 0 | N181
42
43 | [25 empty] | 0 | N182
44
45 | [25 empty] | 0 | N183
46
47 | [13 empty] | DO NOT USE THIS SPACE | [11 empty] | 0 | N184
48 | Date: | [6 empty] | Inspector: | Eugene Mah
49 | Room Number: | [6 empty] | Survey ID: | [17 empty] | 0 | N185
1 | [8 empty] | Medical University of South Carolina
2 | COMMENTS PAGE 2 | [7 empty] | EOS Imaging System Compliance Inspection | [16 empty] | 0 | N186
3 | Comments:
4 | [22 empty] | Exposure Chart
5 | [22 empty] | Patient to film distance (cm) |  | Change | 8.75 | L228
6 | [22 empty] | LF mA |  | Change | 400 | L229
7 | [22 empty] | SF mA |  | Change | 200 | L230
8
9 | [22 empty] | Procedure |  | Change | CXR | L236
10 | [24 empty] | Change | Abdomen AP | L237
11 | [24 empty] | Change | Ribs | L238
12 | [24 empty] | Change | C-Spine AP | L239
13 | [24 empty] | Change | Skull | L240
14 | [24 empty] | Change | Foot | L241
15 | [19 empty] | Medical University of South Carolina |  |  |  |  | Change | Knee | L242
16 | [19 empty] | EOS Imaging System Compliance Inspection |  |  |  |  | Change | Pelvis AP | L243
17 | [14 empty] | COMMENTS PAGE 2 | [9 empty] | Change | Shoulder | L244
18 | [13 empty] | Previous Year Comments: | [10 empty] | Change | LS Spine AP | L245
19 | [10 empty] | Comments:
20 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Grid |  | Change | Yes | M236
21 | [10 empty] | Additional comments: | [13 empty] | Change | Yes | M237
22 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | Yes | M238
23 | [10 empty] | Additional comments: | [13 empty] | Change | Yes | M239
24 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | Yes | M240
25 | [10 empty] | Additional comments: | [13 empty] | Change | No | M241
26 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | No | M242
27 | [10 empty] | Additional comments: | [13 empty] | Change | Yes | M243
28 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | Yes | M244
29 | [10 empty] | Additional comments: | [13 empty] | Change | Yes | M245
30 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0
31 | [10 empty] | Additional comments: | [11 empty] | SID (in.) |  | Change | 130 | N236
32 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 130 | N237
33 | [10 empty] | Additional comments: | [13 empty] | Change | 130 | N238
34 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 130 | N239
35 | [10 empty] | Additional comments: | [13 empty] | Change | 130 | N240
36 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 130 | N241
37 | [10 empty] | Additional comments: | [13 empty] | Change | 130 | N242
38 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 130 | N243
39 | [10 empty] | Additional comments: | [13 empty] | Change | 130 | N244
40 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 130 | N245
41 | [10 empty] | Additional comments:
42 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | thickness (cm) |  | Change | 23 | O236
43 | [10 empty] | Additional comments: | [13 empty] | Change | 23 | O237
44 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 20 | O238
45 | [10 empty] | Additional comments: | [13 empty] | Change | 13 | O239
46 | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 15 | O240
47 | [10 empty] | Additional comments: | [13 empty] | Change | 8 | O241
48 | Date: | [6 empty] | Inspector: | Eugene Mah |  | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 12 | O242
49 | Room Number: | [6 empty] | Survey ID: | [16 empty] | Change | 20 | O243
1 | [8 empty] | Medical University of South Carolina | [15 empty] | Change | 15 | O244
2 | COMMENTS PAGE 2 | [7 empty] | EOS Imaging System Compliance Inspection | [15 empty] | Change | 23 | O245
3 | Comments:
4 | [22 empty] | Field Size (in.) |  | Change | 14x17 | P236
5 | [24 empty] | Change | 14x17 | P237
6 | [24 empty] | Change | 14x17 | P238
7 | [24 empty] | Change | 10x12 | P239
8 | [24 empty] | Change | 8x10 | P240
9 | [24 empty] | Change | 8x10 | P241
10 | [24 empty] | Change | 10x12 | P242
11 | [24 empty] | Change | 14x17 | P243
12 | [13 empty] | DO NOT USE THIS SPACE | [10 empty] | Change | 10x12 | P244
13 | [24 empty] | Change | 10x12 | P245
14
15 | [22 empty] | kVp |  | Change | 120 | Q236
16 | [24 empty] | Change | 76 | Q336
17 | [24 empty] | Change | 66 | Q436
18 | [24 empty] | Change | 72 | Q137
19 | [24 empty] | Change | 72 | Q237
20 | [24 empty] | Change | 60 | Q337
21 | [24 empty] | Change | 66 | Q437
22 | [24 empty] | Change | 78 | Q138
23 | [24 empty] | Change | 74 | Q238
24 | [24 empty] | Change | 78 | Q338
25
26 | [22 empty] | mAs |  | Change | 2 | V236
27 | [24 empty] | Change | 30 | V237
28 | [24 empty] | Change | 12 | V238
29 | [19 empty] | Medical University of South Carolina |  |  |  |  | Change | 10 | V239
30 | [19 empty] | EOS Imaging System Compliance Inspection |  |  |  |  | Change | 10 | V240
31 | [10 empty] | NEW | [13 empty] | Change | 2.5 | V241
32 | [9 empty] | 8.75 | 8.75 | Enter the distance (cm) from the distal side of the patient to the image receptor for typical patient exposure conditions. | [12 empty] | Change | 8 | V242
33 | [9 empty] | 400 | 400 | Typical Large Focus mA (For technique chart) | [12 empty] | Change | 30 | V243
34 | [9 empty] | 200 | 200 | Typical Small Focus mA |  |  | Exposure/Dose Measurement | [9 empty] | Change | 14 | V244
35 | [24 empty] | Change | 25 | V245
36 | [22 empty] | Patient Entrance Exposure:
37 | [12 empty] | Patient |  |  |  | Entrance air kerma (mGy) | [5 empty] | CXR |  |  |  | T236
38 | [11 empty] | SID | thickness | Field size | [8 empty] | Abdomen AP |  |  |  | T237
39 | [9 empty] | Procedure | Grid | (in) | (cm) | (in) | kVp | mAs | [6 empty] | Ribs |  |  |  | T238
40 | [9 empty] | CXR | Yes | 130 | 23 | 14x17 | 120 | 2 | [6 empty] | C-Spine AP |  |  |  | T239
41 | [9 empty] | Abdomen AP | Yes | 130 | 23 | 14x17 | 76 | 30 | [6 empty] | Skull |  |  |  | T240
42 | [9 empty] | Ribs | Yes | 130 | 20 | 14x17 | 66 | 12 | [6 empty] | Foot |  |  |  | T241
43 | [9 empty] | C-Spine AP | Yes | 130 | 13 | 10x12 | 72 | 10 | [6 empty] | Knee |  |  |  | T242
44 | [9 empty] | Skull | Yes | 130 | 15 | 8x10 | 72 | 10 | [6 empty] | Pelvis AP |  |  |  | T243
45 | [9 empty] | Foot | No | 130 | 8 | 8x10 | 60 | 2.5 | [6 empty] | Shoulder |  |  |  | T244
46 | [9 empty] | Knee | No | 130 | 12 | 10x12 | 66 | 8 | [6 empty] | LS Spine AP |  |  |  | T245
47 | [9 empty] | Pelvis AP | Yes | 130 | 20 | 14x17 | 78 | 30 | [6 empty] | Minimum Field Size:
48 | Date: | [6 empty] | Inspector: | Eugene Mah | Shoulder | Yes | 130 | 15 | 10x12 | 74 | 14 | [6 empty] | Frontal
49 | Room Number: | [6 empty] | Survey ID: |  | LS Spine AP | Yes | 130 | 23 | 10x12 | 78 | 25 | [6 empty] | Min. field-Longitudinal |  |  |  | P284
1 | [8 empty] | Medical University of South Carolina | [13 empty] | Min. field-Transverse |  |  |  | Q284
2 | Exposure/Dose Measurement | [7 empty] | EOS Imaging System Compliance Inspection |  | Comments: | [11 empty] | Lateral
3 | [10 empty] | Do NOT exceed 390 |  | 0 | [9 empty] | Min. field-Longitudinal |  |  |  | P285
4 | [10 empty] | New | [11 empty] | Min. field-Transverse |  |  |  | Q285
5 |  |  |  | Patient |  |  |  | Entrance air kerma (mGy) | [14 empty] | Beam Levelness:
6 |  |  | SID | thickness | Field size |  |  |  | KAP | [13 empty] | Source to Table Top (cm) |  | Change | 121.25 | L294
7 | Procedure | Grid | (cm) | (cm) | (in) | kVp | mAs |  | (Gy-cm^2) | Note: 8.75 | cm distance between patient and image receptor | [11 empty] | Table top to cassette (cm) |  | Change | 8.75 | L295
8 | CXR | Yes | 130 | 23 | 14x17 | 120 | 2 | [15 empty] | Frontal Alignment Error (Degrees): |  | Change | TBD | Q298
9 | Abdomen AP | Yes | 130 | 23 | 14x17 | 76 | 30 | [5 empty] | New Values | [9 empty] | Lateral Alignment Error (Degrees): |  | Change | TBD | Q299
10 | Ribs | Yes | 130 | 20 | 14x17 | 66 | 12 | [5 empty] | Patient |  |  |  | Entrance air kerma (mGy) |  | ESE within limits |  |  |  | kVp |  |  |  | N303
11 | C-Spine AP | Yes | 130 | 13 | 10x12 | 72 | 10 |  |  |  |  | SID | thickness | Field size |  |  |  | Percent Change |  | KAP |  |  | mA |  |  |  | P303
12 | Skull | Yes | 130 | 15 | 8x10 | 72 | 10 |  |  | Procedure | Grid | (cm) | (cm) | (in) | kVp | mAs |  |  |  | (Gy-cm^2) |  |  | Time |  |  |  | R303
13 | Foot | No | 130 | 8 | 8x10 | 60 | 2.5 |  |  | CXR | Yes | 130 | 23 | 14x17 | 120 | 2 | [6 empty] | New mAs value |  | Change | 0.5 | T303
14 | Knee | No | 130 | 12 | 10x12 | 66 | 8 |  |  | Abdomen AP | Yes | 130 | 23 | 14x17 | 76 | 30 | [6 empty] | Indicated field size
15 | Pelvis AP | Yes | 130 | 20 | 14x17 | 78 | 30 |  |  | Ribs | Yes | 130 | 20 | 14x17 | 66 | 12 | [6 empty] | Transverse 1 |  |  |  | L310
16 | Shoulder | Yes | 130 | 15 | 10x12 | 74 | 14 |  |  | C-Spine AP | Yes | 130 | 13 | 10x12 | 72 | 10 | [6 empty] | Longitudinal 1 |  |  |  | M310
17 | LS Spine AP | Yes | 130 | 23 | 10x12 | 78 | 25 |  |  | Skull | Yes | 130 | 15 | 8x10 | 72 | 10 | [6 empty] | Transverse 2 |  |  |  | L311
18 | Note: 8.75 cm distance between patient and image receptor | [8 empty] | Foot | No | 130 | 8 | 8x10 | 60 | 2.5 | [6 empty] | Longitudinal 2 |  |  |  | M311
19 | [9 empty] | Knee | No | 130 | 12 | 10x12 | 66 | 8
20 | [9 empty] | Pelvis AP | Yes | 130 | 20 | 14x17 | 78 | 30 | [6 empty] | Spatial Resolution
21 | [9 empty] | Shoulder | Yes | 130 | 15 | 10x12 | 74 | 14 | [6 empty] | Frontal plane
22 | [9 empty] | LS Spine AP | Yes | 130 | 23 | 10x12 | 78 | 25 | [6 empty] | mA/mAs |  | Change | mA | O323
23 | [17 empty] | ESE Limits Acceptable | TBD |  |  |  | kVp |  |  |  | N324
24 | [22 empty] | mA |  |  |  | O324
25 | [19 empty] | Medical University of South Carolina |  |  | Time |  | #REF! | #REF! | P324
26 | [14 empty] | Measurement Data |  |  |  |  | EOS Imaging System Compliance Inspection |  |  | SID |  | Change | 130 | Q324
27 | [22 empty] | OFD |  | Change | 9 | R324
28 | [9 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [12 empty] | Last year Lrg FS resolution |  |  | 0 | T324
29 | [14 empty] | Beam Limiting Device | [7 empty] | Previous 1 yr Lrg FS resolution |  |  |  | U324
30 | [11 empty] | A beam limiting device is provided to restrict the size of the useful beam. | [10 empty] | Lateral plane
31 | [11 empty] | Positive means is provided to ensure the minimum filtration is in beam. | [10 empty] | mA/mAs |  | Change | mA | O328
32 | [22 empty] | kVp |  |  |  | N329
33 | [10 empty] | Comments: | [11 empty] | mA |  |  |  | O329
34 | [10 empty] | Do NOT exceed 390 |  | 0 | [9 empty] | Time |  | #REF! | #REF! | P329
35 | [10 empty] | New | [11 empty] | SID |  | Change | 130 | Q329
36 | [22 empty] | OFD |  | Change | 9 | R329
37 | [10 empty] | Minimum |  |  | Measurements (cm) | [8 empty] | Last year Sm FS resolution |  |  | 0 | T329
38 | [10 empty] | Field Size |  |  | Longitudinal | Transverse | [7 empty] | Previous 1 yr Sm FS resolution |  |  |  | U329
39 | [12 empty] | Frontal Plane: |  |  |  |  | Acceptable: | TBD
40 | [12 empty] | Lateral Plane: | [9 empty] | Test Equipment Utilized: | Piranha 2 CB2-19020491 |  | Piranha 2 CB2-19020491 | N337
41 | [10 empty] | Last Year: |  | Frontal Plane:
42 | [12 empty] | Lateral Plane: | [9 empty] | Half Value Layer
43 | [10 empty] | Criteria: | Each dimension of the minimum field size at an SID of 100 cm shall be <= 5 cm | [10 empty] | Frontal plane: |  |  |  | P405
44 | [10 empty] | Comments: | [11 empty] | Lateral plane: |  |  |  | U405
45 | [10 empty] | Do NOT exceed 390 |  | 0
46 | [10 empty] | New | [11 empty] | AEC Thickness Tracking
47 | [22 empty] | Frontal plane
48 | Date: | [6 empty] | Inspector: | Eugene Mah | Beam Levelness | [11 empty] | 10 cm | kV |  |  |  | M429
49 | Room Number: | [6 empty] | Survey ID: |  | 121.25 | 121.25 | (cm) Source to panel |  |  |  | IMPORTANT!! Select below if radiation field is | [6 empty] | mAs |  |  |  | N429
1 | [8 empty] | Medical University of South Carolina | 8.75 | 8.75 | (cm) Table top to detector |  |  |  | centered to upper or lower bead of the test object. | [6 empty] | mGy |  |  |  | O429
2 | [8 empty] | EOS Imaging System Compliance Inspection |  |  |  | Measured Separation |  | Error |  | Upper |  |  | Upper |  | 20 cm | kV |  |  |  | M430
3 | [12 empty] | Of Steel Balls (mm) |  | (Degrees) | Last Year: |  |  |  | Lower |  |  | mAs |  |  |  | N430
4 | Minimum Field Size | [10 empty] | Frontal: |  |  | TBD | [7 empty] | mGy |  |  |  | O430
5 |  |  | Measurements (cm) | [8 empty] | Lateral: |  |  | TBD |  | Acceptable: | TBD |  |  |  | 30 cm | kV |  |  |  | M431
6 |  |  | Longitudinal | Transverse | [6 empty] | Criteria: | X-ray beam is level to within 0.5 degrees | [10 empty] | mAs |  |  |  | N431
7 |  | Frontal Plane: |  |  |  |  | Acceptable: | TBD | [14 empty] | mGy |  |  |  | O431
8 |  | Lateral Plane: | [7 empty] | Technique: | [6 empty] | Last Year: | [5 empty] | Lateral plane
9 | Criteria: | Each dimension of the minimum field size at an SID of 100 cm shall be <= 5 cm | [8 empty] | kVp: |  | mA: |  | time: |  | mAs: | 0.5 |  |  |  | 10 cm | kV |  |  |  | Q429
10 | [10 empty] | New: |  | New: |  | New: |  | New: | [5 empty] | mAs |  |  |  | R429
11 | Beam Levelness | [21 empty] | mGy |  |  |  | S429
12 |  |  | Measured Separation |  | Error |  |  |  |  | Radiation field size | [11 empty] | 20 cm | kV |  |  |  | Q430
13 |  |  | Of Steel Balls (mm) |  | (Degrees) |  |  |  |  | Indicated Field Size @SID |  | Meas Radiation (Frontal) |  | Meas Radiation (Lateral) |  | Frontal |  | Lateral |  |  |  |  | mAs |  |  |  | R430
14 |  | Frontal Plane: |  |  | TBD |  | Acceptable: | TBD |  | (cm) |  | (cm) |  | (cm) |  | Error (% SID) |  | Error (% SID) |  |  |  |  | mGy |  |  |  | S430
15 |  | Lateral Plane: |  |  | TBD |  |  |  |  | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal |  |  | 30 cm | kV |  |  |  | Q431
16 | Criteria: | X-ray beam is level to within 0.5 degrees | [13 empty] | NA | NA | NA | NA |  |  |  | mAs |  |  |  | R431
17 | [15 empty] | NA | NA | NA | NA |  |  |  | mGy |  |  |  | S431
18 | Radiation field size | [8 empty] | New Values
19 | kVp: |  | mA: |  | mAs: | 0.5 | SID: | 130 |  | Transverse | Longitudinal | [5 empty] | Acceptable | [5 empty] | AEC Reproducibility
20 | [13 empty] | Indicated field size (cm) |  | Frontal |  | Lateral |  |  |  |  | Frontal plane
21 | [13 empty] | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal |  |  |  | kV |  |  |  | M437
22 | Frontal plane | [14 empty] | NA | NA | NA | NA |  |  |  | mAs |  |  |  | N437
23 | Indicated Field Size @SID |  | Measured Radiation | [12 empty] | NA | NA | NA | NA |  |  |  | mGy |  |  |  | O437
24 | (cm) |  | (cm) |  | Error (% SID) |  | Acceptable | [15 empty] | kV |  |  |  | M438
25 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | [14 empty] | mAs |  |  |  | N438
26 |  |  |  |  | NA | NA | NA | NA | [6 empty] | X-ray Tube |  |  | 1 Year | 2 Year |  |  |  | mGy |  |  |  | O438
27 |  |  |  |  | NA | NA | NA | NA |  | Focal Spot Resolution | [6 empty] | Measured | Previous | Previous |  |  |  | kV |  |  |  | M439
28 | Lateral plane | [8 empty] | Frontal tube |  |  | mA |  |  |  | Resolution | Resolution | Resolution |  |  |  | mAs |  |  |  | N439
29 | [11 empty] | kVp | mA | SID* | ODD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  |  |  | mGy |  |  |  | O439
30 | Indicated Field Size @SID |  | Measured Radiation | [10 empty] | 130 | 9 | 1.0744 | [6 empty] | Lateral plane
31 | (cm) |  | (cm) |  | Error (% SID) |  | Acceptable |  |  |  | New Values: |  |  | 130 | 9 | [7 empty] | kV |  |  |  | R437
32 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | [8 empty] | Measured | Previous | Previous |  |  |  | mAs |  |  |  | S437
33 |  |  |  |  | NA | NA | NA | NA |  | Lateral tube |  |  | mA |  |  |  | Resolution | Resolution | Resolution |  |  |  | mGy |  |  |  | T437
34 |  |  |  |  | NA | NA | NA | NA |  |  |  | kVp | mA | SID* | ODD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  |  |  | kV |  |  |  | R438
35 | Criteria: | The indicated field length/width are accurate to within 2% of the SID. | [11 empty] | 130 | 9 | 1.0744 | [6 empty] | mAs |  |  |  | S438
36 |  | X-ray field length/width are congruent within 2% SID. | [8 empty] | New Values: |  |  | 130 | 9 | [7 empty] | mGy |  |  |  | T438
37 | [7 empty] | 1 Year | 2 Year |  |  | * SID = Source to Image (detector) Distance. | [10 empty] | kV |  |  |  | R439
38 | Focal Spot Resolution | [5 empty] | Measured | Previous | Previous |  |  | ** ODD = Object (resolution pattern) to Detector Distance. | [10 empty] | mAs |  |  |  | S439
39 | [6 empty] | Resolution | Resolution | Resolution | Enter 1 for YES, 2 for NO, 3 for NA | [12 empty] | mGy |  |  |  | T439
40 |  | kVp | mA | SID* | ODD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) | 3 | Focal spot resolution is within acceptable parameters
41 | Frontal tube |  |  | 130 | 9 | 1.0744 | [5 empty] | Criteria: | No significant change in resolution since previous measurement. | [9 empty] | Frontal Plane Output
42 | Lateral tube |  |  | 130 | 9 | 1.0744 | [16 empty] | SDD(cm) |  | Change | 119 | M340
43 |  | * SID = Source to Image (film) Distance. |  |  |  | Acceptable: | NA |  |  |  | Test Equipment Used: | Piranha 2 CB2-19020491 |  |  | Calibration Date:
44 |  | ** OFD = Object (resolution pattern) to Film Distance. | [8 empty] | New: | Piranha 2 CB2-19020491 |  |  | Calibration Due: | [7 empty] | kVp |  |  |  | O472
45 | Criteria: | No significant change in resolution since previous measurement. | [7 empty] | kV Accuracy and Output | [16 empty] | O473
46 | [9 empty] | SDD* (cm): | 119 | Focal spot: | Large | [13 empty] | O474
47 | [9 empty] | Frontal Tube |  |  |  |  | Lateral Tube | [11 empty] | O475
48 | Date: | [6 empty] | Inspector: | Eugene Mah | Set values |  | Measured values |  |  | Set values |  | Measured values | [9 empty] | O476
49 | Room Number: | [6 empty] | Survey ID: |  | kV | mA | kV (avg) | time (sec) | mGy | kV | mA | kV (avg) | time(SEC) | mGy | [7 empty] | O477
1 | [8 empty] | Medical University of South Carolina | 50 | 400 |  |  |  | 50 | 400 | [10 empty] | O478
2 | GENERATOR PERFORMANCE | [7 empty] | EOS Imaging System Compliance Inspection | 60 | 400 |  |  |  | 60 | 400 | [10 empty] | O479
3 |  | Test Equipment Used: | Piranha 2 CB2-19020491 |  | Cal Date: |  | Cal Due: |  |  | 70 | 400 |  |  |  | 70 | 400 | [10 empty] | O480
4 | [9 empty] | 80 | 400 |  |  |  | 80 | 400 | [10 empty] | O481
5 | kVp Accuracy | [8 empty] | 90 | 200 |  |  |  | 90 | 200
6 |  | SDD: | 119 | mA: | 400 |  |  |  |  | 100 | 400 |  |  |  | 100 | 400 | [6 empty] | mGy/mA |  |  |  | P472
7 | [9 empty] | 110 | 400 |  |  |  | 110 | 400 | [6 empty] | @ 119 cm |  |  |  | P473
8 |  |  | Frontal Tube |  | Lateral Tube |  |  |  |  | 120 | 400 |  |  |  | 120 | 400 | [10 empty] | P474
9 |  | Indicated kV | Measured kV |  | Measured kV |  |  |  |  | 130 | 400 |  |  |  | 130 | 400 | [10 empty] | P475
10 |  |  |  | % error |  | % error |  |  |  | 140 | 400 |  |  |  | 140 | 400 | [10 empty] | P476
11 |  | 50 | [7 empty] | *SDD=Source to detector distance | [16 empty] | P477
12 |  | 60 | [9 empty] | Acceptable: | TBD |  |  |  | Acceptable: | TBD | [8 empty] | P478
13 |  | 70 | [7 empty] | Criteria: | kVp accurate to within 5% of indicated. | [15 empty] | P479
14 |  | 80 | [24 empty] | P480
15 |  | 90 | [7 empty] | kV, Exposure, Timer Reproducibility | [16 empty] | P481
16 |  | 100 | [7 empty] | Frontal Tube |  |  |  |  | Lateral Tube
17 |  | 110 | [7 empty] | Set values |  | Measured values |  |  | Set values |  | Measured values | [5 empty] | Lateral Plane Output
18 |  | 120 | [7 empty] | kV | mA | kV (avg) | time (sec) | mGy | kV | mA | kV (avg) | time(SEC) | mGy |  |  |  | SDD(cm) |  | Change | 119 | N1411
19 |  | 130 | [7 empty] | 80 | 400 |  |  |  | 80 | 400
20 |  | 140 | [7 empty] | 80 | 400 |  |  |  | 80 | 400 | [6 empty] | kVp |  |  |  | O487
21 |  |  | Acceptable: | TBD | Acceptable: | TBD |  |  |  | 80 | 400 |  |  |  | 80 | 400 | [10 empty] | O488
22 | Criteria: | kVp accurate to within 5% of indicated. | [7 empty] | 80 | 400 |  |  |  | 80 | 400 | [10 empty] | O489
23 | [9 empty] | 80 | 400 |  |  |  | 80 | 400 | [10 empty] | O490
24 | kVp and Exposure Reproducibility | [9 empty] | Mean: |  |  |  |  | Mean: | [10 empty] | O491
25 |  |  | Frontal Tube |  | Lateral Tube | [5 empty] | SD |  |  |  |  | SD | [10 empty] | O492
26 |  | Indicated kV | Measured |  | Measured | [5 empty] | Coeff of Var: |  |  |  |  | Coeff of Var: | [10 empty] | O493
27 |  |  | kV | mGy | kV | mGy |  |  |  |  | Acceptable: | NA | NA | NA |  | Acceptable: | NA | NA | NA | [7 empty] | O494
28 |  | 80 | [7 empty] | Criteria: | Coefficient of variation (std dev./mean) <= 5%. | [15 empty] | O495
29 | [26 empty] | O496
30 | [9 empty] | Radiation Output Linearity
31 | [9 empty] | Frontal Tube |  |  |  |  | Lateral Tube | [7 empty] | mGy/mAs |  |  |  | P487
32 | [9 empty] | Set values |  | Measured values |  |  | Set values |  | Measured values | [5 empty] | @ 119 cm |  |  |  | P488
33 |  | Mean: | [7 empty] | kV | mA | kV (avg) | mGy | mGy/mA | kV | mA | kV (avg) | mGy | mGy/mA | [7 empty] | P489
34 |  | SD | [7 empty] | 80 | 400 |  |  |  | 80 | 400 | [10 empty] | P490
35 |  | Coeff of Var: | [7 empty] | 80 | 500 |  |  |  | 80 | 500 | [10 empty] | P491
36 |  | Acceptable: | NA | NA | NA | NA |  |  |  | 80 | 50 |  |  |  | 80 | 50 | [10 empty] | P492
37 | Criteria: | Coefficient of variation (std dev./mean) <= 5%. | [7 empty] | 80 | 720 |  |  |  | 80 | 720 | [10 empty] | P493
38 | [9 empty] | 80 | 250 |  |  |  | 80 | 250 | [10 empty] | P494
39 | [12 empty] | Coeff of Linearity: |  |  |  |  | Coeff of Linearity: | [8 empty] | P495
40 | [12 empty] | Acceptable: | NA |  |  |  | Acceptable: | NA | [7 empty] | P496
41 | [9 empty] | Criteria: | Coefficient of linearity <= 10%.
42 | [22 empty] | mGy/mAs Output
43 | [9 empty] | Half Value Layer Measurement | [12 empty] | Frontal Plane Max. |  | Change | NA | Q481
44 | [9 empty] | Frontal Tube |  |  |  |  | Lateral Tube | [7 empty] | Frontal Plane Median |  | Change | NA | R481
45 | [9 empty] | Set values |  |  | Measured values |  | Set values |  |  | Measured values |  |  |  |  | Frontal Plane Min |  | Change | NA | S481
46 | [9 empty] | kV | mA | mm Al | kV (avg) | mGy | kV | mA | mm Al | kV (avg) | mGy |  |  |  | Lateral Plane Max. |  | Change | NA | Q499
47 | [9 empty] | 80 | 400 | 0 |  |  | 80 | 400 | 0 | [5 empty] | Lateral Plane Median |  | Change | NA | R499
48 | Date: | [6 empty] | Inspector: | Eugene Mah | 80 | 400 | 0 |  |  | 80 | 400 | 0 | [5 empty] | Lateral Plane Min |  | Change | NA | S499
49 | Room Number: | [6 empty] | Survey ID: |  | 80 | 400 | 3 |  |  | 80 | 400 | 3
1 | [8 empty] | Medical University of South Carolina | 80 | 400 | 3 |  |  | 80 | 400 | 3
2 | GENERATOR PERFORMANCE | [7 empty] | EOS Imaging System Compliance Inspection | 80 | 400 | 4 |  |  | 80 | 400 | 4
3 |  | Test Equipment Used: | Piranha 2 CB2-19020491 | [6 empty] | 80 | 400 | 4 |  |  | 80 | 400 | 4
4 | [9 empty] | 80 | 400 | 5 |  |  | 80 | 400 | 5
5 | Radiation Output Linearity | [8 empty] | 80 | 400 | 5 |  |  | 80 | 400 | 5
6 |  | kV: | 80 | SDD: | 119
7 | [9 empty] | Data Processing:
8 |  |  | Frontal Tube |  |  | Lateral Tube |  |  |  | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | ln(Exp) | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | ln(Exp)
9 |  |  | Measured values |  |  | Measured values
10 |  | mA | kV (avg) | mGy | mGy/mA | kV (avg) | mGy | mGy/mA
11 |  | 400
12 |  | 500
13 |  | 50 | [10 empty] | HVL @80 kV: |  |  |  |  | HVL @80 kV:
14 |  | 720 | [10 empty] | Min HVL: |  |  |  |  | Min HVL:
15 |  | 250 | [10 empty] | Acceptable: | TBD |  |  |  | Acceptable: | TBD
16 |  |  |  | Coeff of Linearity: |  |  | Coeff of Linearity: | [5 empty] | Previous: |  |  |  |  | Previous:
17 |  |  |  | Acceptable: | NA |  | Acceptable: | NA |  | Enter 1 for YES, 2 for NO, 3 for NA
18 | Criteria: | Coefficient of linearity <= 10%.
19 | [11 empty] | Means provided to indicate technique factors prior to exposure.
20 | Half Value Layer Measurement | [10 empty] | All technique indicators functional and are free of excessive wear.
21 | [11 empty] | Technique overload circuits function to block excessive technique
22 | Frontal Tube |  |  |  | Lateral Tube | [6 empty] | Anode heat calculator/indicator is functioning properly
23 | kV: | 80 | SDD: | 119 | kV: | 80 | SDD: | 119 |  |  |  | Indicator of the individual x-ray tube selected present and operational
24 | mA: | 400 |  |  | mA: | 400 | [5 empty] | Visible "beam on" indicator present and operational
25 | Al (mm) | mGy |  | HVL | Al (mm) | mGy |  | HVL |  |  |  | Audible exposure end indicator present and operational
26 | 0 |  |  |  | 0 | [6 empty] | Exposure delay time is not excessive.
27 |  |  |  | Acceptable: |  |  |  | Acceptable:
28 |  |  |  | TBD |  |  |  | TBD
29 | [19 empty] | Medical University of South Carolina
30 | [14 empty] | Measurement Data |  |  |  |  | EOS Imaging System Compliance Inspection
31 | Criteria: | Minimum acceptable HVL at 80 kV: 2.9 mm Al
32 | [14 empty] | Automatic Exposure Control
33 | [9 empty] | Thickness Tracking
34 | [10 empty] | Frontal Tube |  |  |  | Lateral Tube
35 | [9 empty] | Thickness |  |  |  | Previous |  |  |  | Previous
36 | [9 empty] | (cm) | kV | mA | mGy | mGy | kV | mA | mGy | mGy
37 | [9 empty] | Thickness (cm): 10
38 | [9 empty] | Thickness (cm): 20
39 | [9 empty] | Thickness (cm): 30
40
41 | [9 empty] | Thickness (cm): Reproducibility
42 | [9 empty] | Thickness kV: Frontal Tube; Previous mA: Lateral Tube
43 | [9 empty] | Thickness (cm): Thickness; Previous mGy: Previous; Previous kV: Deviation | Previous | Deviation
44 | [9 empty] | (cm) | kV | mA | mGy | mGy | from mean | kV | mA | mGy | mGy | from mean
45 | [9 empty] | Thickness (cm): 20
46 | [9 empty] | Thickness (cm): 20
47 | [9 empty] | Thickness (cm): 20
48 | Date: | [6 empty] | Inspector: | Eugene Mah | Previous mGy: Acceptable:; Previous kV: TBD | Acceptable: | TBD
49 | Room Number: | [6 empty] | Survey ID: |  | Thickness (cm): Criteria:; Thickness kV: Exposure deviation is less than 11% from the mean value
1 | [8 empty] | Medical University of South Carolina
2 | [8 empty] | EOS Imaging System Compliance Inspection | Previous kV: Kerma Area Product
3 | [9 empty] | Thickness (cm): Chamber:; Thickness kV: DAP 1 2501554; Thickness mGy: Calibration date:; Previous mA: Calibration due:
4 | AEC Thickness Tracking
5 |  | Thickness | Frontal Tube |  |  | Lateral Tube |  |  |  | Thickness (cm): Frontal Tube; Previous kV: Lateral Tube
6 |  | (cm) | kV | mAs | mGy | kV | mAs | mGy |  | Thickness mA: Indicated; Thickness mGy: Measured; Previous mGy: Indicated; Previous mGy: Measured
7 |  | 10 | [7 empty] | Thickness mA: KAP; Thickness mGy: KAP; Previous mGy: KAP; Previous mGy: KAP
8 |  | 20 | [7 empty] | kV | mA | µGy·m² |  | % diff | kV | mA | µGy·m² |  | % diff
9 |  | 30 | [7 empty] | Thickness (cm): 60; Thickness kV: 100; Previous kV: 60; Previous mA: 100
10 | [9 empty] | Thickness (cm): 80; Thickness kV: 50; Previous kV: 80; Previous mA: 50
11 | AEC Reproducibility | [8 empty] | Thickness (cm): 100; Thickness kV: 50; Previous kV: 100; Previous mA: 50
12 |  | Frontal Tube |  |  |  | Lateral Tube |  |  |  | Thickness (cm): 120; Thickness kV: 25; Previous kV: 120; Previous mA: 25
13 | Thickness |  |  |  | Deviation |  |  |  | Deviation | bar: Thickness (cm) to Thickness kV; Previous kV to Previous mA
14 | (cm) | kV | mAs | mGy | from mean | kV | mAs | mGy | from mean | bar: Thickness (cm) to Thickness kV; Previous kV to Previous mA
15 | 20 | bar: Thickness (cm) to Thickness kV; Previous kV to Previous mA
16 | 20 | bar: Thickness (cm) to Thickness kV; Previous kV to Previous mA
17 | 20 | [8 empty] | Thickness mGy: Acceptable:; Previous mGy: TBD; Previous mGy: Acceptable: | TBD
18 |  |  |  | Acceptable: | TBD |  |  | Acceptable: | TBD
19 | Criteria: | Exposure deviation is less than 11% from the mean value | [7 empty] | Thickness (cm): Criteria:; Thickness kV: Measured KAP should be within +/-15% of the indicated value
20
21 | Kerma Area Product Accuracy
22 |  | DAP Chamber: | DAP 1 2501554 |  | Calibration date: |  |  | Calibration due: | [11 empty] | Medical University of South Carolina
23 | [19 empty] | EOS Imaging System Compliance Inspection
24 |  | Frontal Tube
25 |  |  |  | Indicated | Measured |  |  |  |  | Thickness (cm): Output - Frontal Plane; Previous mGy: (Using |  |  | Using
26 |  |  |  | KAP | KAP |  |  |  |  | Previous kV: Current; Previous mA: Previous; Previous mGy: Previous | (Using | (Using | Current
27 |  | kV | mA | µGy·m² |  | % diff |  |  |  | Thickness (cm): Previous* Measurements; Thickness mGy: Current Meas.; Previous kV: mGy/mA; Previous mA: mGy/mA; Previous mGy: mGy/mA | Recorded | Current | mGy/mA
28 |  | 60 | 100 |  |  |  |  | Acceptable |  | SDD (cm): |  |  | SDD: | 119 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current
29 |  | 80 | 50 |  |  |  |  | TBD |  |  | mGy/mA | mGy/mA |  | mGy/mA | kVp | kVp | Percent | 100 cm SID) | kVp) | kVp) | kVp
30 |  | 100 | 50 | [6 empty] | kVp | @SDD | @100 cm | kV | @SDD | @100 cm | @100 cm | Change | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40"
31 |  | 120 | 25
32
33
34
35
36 |  | Lateral Tube
37 |  |  |  | Indicated | Measured
38 |  |  |  | KAP | KAP
39 |  | kV | mA | µGy·m² |  | % diff
40 |  | 60 | 100 |  |  |  |  | Acceptable
41 |  | 80 | 50 |  |  |  |  | TBD |  | Thickness (cm): *From previous compliance report; Previous kV: Max; Previous mA: Median; Previous mGy: Min.
42 |  | 100 | 50 | [15 empty] | Acceptable: | NA | bar: Previous kV to Previous mGy
43 |  | 120 | 25 | [6 empty] | Previous mGy: Last Year:
44
45 | [9 empty] | Thickness (cm): Output - Lateral Plane; Previous mGy: (Using |  |  | Using
46 | [9 empty] | Previous kV: Current; Previous mA: Previous; Previous mGy: Previous | (Using | (Using | Current
47 | [9 empty] | Thickness (cm): Previous* Measurements; Thickness mGy: Current Meas.; Previous kV: mGy/mA; Previous mA: mGy/mA; Previous mGy: mGy/mA | Recorded | Current | mGy/mA
48 | Date: | [6 empty] | Inspector: | Eugene Mah | SDD (cm): |  |  | SDD: | 119 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current
49 | Room Number: | [6 empty] | Survey ID: |  |  | mGy/mA | mGy/mA |  | mGy/mA | kVp | meas. kVp | Percent | 100 cm SID) | kVp) | kVp) | kVp
1 | [8 empty] | Medical University of South Carolina | kVp | @SDD | @100 cm | kV | @SDD | @100 cm | @100 cm | Change | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40"
2 | [8 empty] | EOS Imaging System Compliance Inspection
3 | Output - Frontal Plane
4 |  |  | Current | Previous
5 |  |  | mGy/mA | mGy/mA
6 |  |  | @ Current | @ Current
7 |  |  | kVp | kVp | Percent
8 |  | kV | @100 cm | @100 cm | Change | Acceptable
9
10
11
12 | [9 empty] | Thickness (cm): *From previous compliance report; Previous kV: Max; Previous mA: Median; Previous mGy: Min.
13 | [18 empty] | Acceptable: | NA | bar: Previous kV to Previous mGy
14 | [9 empty] | Previous mGy: Last Year:
15
16 | [9 empty] | Thickness (cm): Criteria:; Thickness kV: The x-ray output @100 cm is within +/-25% from previous measurements.
17
18 | [6 empty] | Acceptable: | NA |  | Thickness mA: Frontal plane; Previous kV: Lateral plane
19 | [9 empty] | Thickness mA: Current; Previous kV: Current
20 | Output - Lateral Plane | [8 empty] | Thickness mA: mGy/mAs; Previous kV: mGy/mAs
21 |  |  | Current | Previous | [5 empty] | Thickness kV: kVp; Thickness mA: @100 cm; Previous mGy: kVp; Previous kV: @100 cm
22 |  |  | mGy/mA | mGy/mA | [5 empty] | Thickness kV: 50; Previous mGy: 50
23 |  |  | @ Current | @ Current | [5 empty] | Thickness kV: 60; Previous mGy: 60
24 |  |  | kVp | kVp | Percent |  |  |  |  | Thickness kV: 70; Previous mGy: 70
25 |  | kV | @100 cm | @100 cm | Change | Acceptable |  |  |  | Thickness kV: 80; Previous mGy: 80
26 | [9 empty] | Thickness kV: 90; Previous mGy: 90
27 | [9 empty] | Thickness kV: 100; Previous mGy: 100
28 | [9 empty] | Thickness kV: 110; Previous mGy: 110
29 | [9 empty] | Thickness kV: 120; Previous mGy: 120
30 | [9 empty] | Thickness kV: 130; Previous mGy: 130
31 | [9 empty] | Thickness kV: 140; Previous mGy: 140
32
33
34 | [9 empty] | Thickness (cm): Enter 1 for YES, 2 for NO, 3 for NA
35 | [6 empty] | Acceptable: | NA |  | Previous mGy: Radiation Safety
36 | [9 empty] | Thickness mA: Continuous pressure on exposure switch required to make an exposure with selected tube
37 | [9 empty] | Thickness mA: Exposure switch for selected tube is secured at a safe location
38 | [9 empty] | Thickness mA: Operator has full visibility of the exposure factors and patient
39 | [9 empty] | Thickness mA: Lead apron available
40 | [9 empty] | Thickness mA: Lead gloves available.
41 | [9 empty] | Thickness mA: A properly designed and installed apron rack is present.
42 | [9 empty] | Thickness mA: Documentation of annual protective apparel integrity inspection is available.
43 | [9 empty] | Thickness mA: Patient restraint devices available.
44 | [9 empty] | Thickness mA: Gonadal shielding is present and functional.
45 | [9 empty] | Thickness mA: Radiation warning light at room entrance is functional with selected tube
46
47 | [9 empty] | Previous mGy: Physical Inspection
48 | Date: | [6 empty] | Inspector: | Eugene Mah | Thickness mA: System clean and free of debris.
49 | Room Number: | [6 empty] | Survey ID: |  | Thickness mA: All cables free from excessive wear or restraints.
1 | [8 empty] | Medical University of South Carolina | Thickness mA: System free of missing or damaged components.
2 | [8 empty] | EOS Imaging System Compliance Inspection | Thickness mA: Platform movements and locks acceptable
3 | [9 empty] | Thickness mA: Tube crane movements and locks acceptable.
4 | [9 empty] | Thickness mA: All other movements and locks functional.
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48 | Date: | [6 empty] | Inspector: | Eugene Mah
49 | Room Number: | [6 empty] | Survey ID:

## Frontal Gen
Frontal tube
Measurements
Set kV (kV) | Set mA (mA) | Set mAs (mAs) | Added Filt | Focal spot | Tube voltage (kV) | Exposure time (s) | Exposure (mGy) | Exposure rate (mGy/s) | Total filtr. (mm Al) | HVL (mm Al) | Exposure (norm) (mGy) | Exposure rate (norm) (mGy/s)
50 | 400 | 10 | 0 | Large
60 | 400 | 20 | 0 | Large
70 | 400 | 10 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 3 | Large
80 | 400 | 20 | 3 | Large
80 | 400 | 20 | 4 | Large
80 | 400 | 20 | 4 | Large
80 | 400 | 20 | 5 | Large
80 | 400 | 20 | 5 | Large
80 | 500 | 25 | 0 | Large
80 | 50 | 2.5 | 0 | Large
80 | 720 | 40 | 0 | Large
80 | 250 | 12.5 | 0 | Large
90 | 200 | 10 | 0 | Large
100 | 400 | 20 | 0 | Large
110 | 400 | 10 | 0 | Large
120 | 400 | 20 | 0 | Large
130 | 400 | 10 | 0 | Large
140 | 400 | 20 | 0 | Large

## Lateral Gen
Lateral tube
Measurements
Set kV (kV) | Set mA (mA) | Set mAs (mAs) | Added Filt | Focal spot | Tube voltage (kV) | Exposure time (s) | Exposure (mGy) | Exposure rate (mGy/s) | Total filtr. (mm Al) | HVL (mm Al) | Exposure (norm) (mGy) | Exposure rate (norm) (mGy/s)
50 | 400 | 10 | 0 | Large
60 | 400 | 20 | 0 | Large
70 | 400 | 10 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 0 | Large
80 | 400 | 20 | 3 | Large
80 | 400 | 20 | 3 | Large
80 | 400 | 20 | 4 | Large
80 | 400 | 20 | 4 | Large
80 | 400 | 20 | 5 | Large
80 | 400 | 20 | 5 | Large
80 | 500 | 25 | 0 | Large
80 | 50 | 2.5 | 0 | Large
80 | 720 | 40 | 0 | Large
80 | 250 | 12.5 | 0 | Large
90 | 200 | 10 | 0 | Large
100 | 400 | 20 | 0 | Large
110 | 400 | 10 | 0 | Large
120 | 400 | 20 | 0 | Large
130 | 400 | 10 | 0 | Large
140 | 400 | 20 | 0 | Large

## AEC
Measurements
Plane | Thickness (cm) | Set kV (kV) | Set mA (mA) | Exposure (mGy) | Exposure rate (mGy/min) | Exposure time (ms)
Frontal | 10
Frontal | 20
Frontal | 20
Frontal | 20
Frontal | 30
Lateral | 10
Lateral | 20
Lateral | 20
Lateral | 20
Lateral | 30

## DAP
Measurements
Plane | Set kV (kV) | Set mA (mA) | Set time (s) | Set mAs (mAs) | KAP exposure (µGym²) | KAP exposure rate (mGycm²/s) | Indicated KAP | KAP % diff
Frontal | 60 | 100 | 0.1 | 10
Frontal | 80 | 50 | 0.2 | 10
Frontal | 100 | 50 | 0.2 | 10
Frontal | 120 | 25 | 0.4 | 10
Lateral | 60 | 100 | 0.1 | 10
Lateral | 80 | 50 | 0.2 | 10
Lateral | 100 | 50 | 0.2 | 10
Lateral | 120 | 25 | 0.4 | 10

## Tables
RTI Piranha
Piranha 1 | Piranha 1 CB2-17090320
Piranha 2 | Piranha 2 CB2-19020491
RTI DAP chamber
DAP 1 | DAP 1 2501554
DAP 2 | DAP 2 2501555
KAP Conversion Table
µGy·m²
mGy·cm² | 0.1
cGy·cm² | 1
dGy·cm² | 10
Gy·cm² | 100
µGy·m² | 1
ESE Limits - Title B Part IV Appendix D
|  |  | Film speed |  | Film speed
|  | Thickness | 200/Dig | 400 | 200/Dig | 400
Projection |  | (cm) | mR |  | mGy
PA Chest - Grid |  | 23 | 38 | 23 | 0.33288 | 0.20148
PA Chest - Non grid |  | 23 | 23 | 8 | 0.20148 | 0.07008
AP Abdomen |  | 23 | 735 | 450 | 6.4386 | 3.942
AP Lumbar spine |  | 23 | 675 | 525 | 5.913 | 4.599
AP full spine |  | 23 | 390 | 218 | 3.4164 | 1.9097
AP C-spine |  | 13 | 203 | 142 | 1.7783 | 1.2439
Lateral skull |  | 15 | 218 | 105 | 1.9097 | 0.9198
AP Ret Pyelogram |  | 23 | 893 | 893 | 7.8227 | 7.8227
AP T-spine |  | 23 | 612 | 612 | 5.3611 | 5.3611
DP Foot |  | 8 | 111 | 111 | 0.97236 | 0.97236
Cephalometric |  | 15 | 45 | 45 | 0.3942 | 0.3942
RAD / DENTAL UNITS (SCDES CRITERIA)
HVL
kVp | (mm Al)
30 | 0.3
40 | 0.4
50 | 0.5
51 | 1.3
60 | 1.5
70 | 1.8
71 | 2.5
80 | 2.9
90 | 3.2
100 | 3.6
110 | 3.9
120 | 4.3
130 | 4.7
140 | 5
150 | 5.4
Huttner Grating
1.1 | 0.5
1.2 | 0.56
1.3 | 0.63
2.1 | 0.71
2.2 | 0.8
2.3 | 0.9
3.1 | 1
3.2 | 1.12
3.3 | 1.25
4.1 | 1.4
4.2 | 1.6
4.3 | 1.8
5.1 | 2
5.2 | 2.24
5.3 | 2.5
6.1 | 2.8
6.2 | 3.15
6.3 | 3.55
7.1 | 4
7.2 | 4.5
7.3 | 5
Note on A9: Eugene Mah: Multiply by this number to get µGy·m²