## QC Test Summary
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY
Site Name |  |  |  |  | Report Date
Address |  |  |  |  | Survey Date | 0
X-Ray Unit Manufacturer |  |  |  |  | Model
Date of Installation |  |  |  |  | Room ID
Film (mfr & type) |  |  |  |  | Screen (mfr & type)
Film Processor Mfr |  |  |  |  | Model
Digital Image Receptor Mfr |  |  |  |  | Model
Medical Physicist's Name |  | Eugene Mah |  |  | Signature
Medical Physicist's QC Tests
PASS/FAIL
1. Stereotactic Breast Biopsy Unit Assembly Evaluation
2. Collimation Assessment
A. X-ray field adequately matches image receptor
B. Biopsy window generally centered over digital image receptor (NA if film used)
3. Focal Spot Performance and System Limiting Resolution
A. Focal spot performance acceptable (NA if digital used)
B. Digital system spatial resolution acceptable (NA if film used)
4. kVp Accuracy and Reproducibility
Measured average kVp within ±5% of indicated kVp
kVp coefficient of variation ≤ 0.02
5. Beam Quality Assessment (Half-Value Layer Measurement)
HVL is within acceptable lower and upper limits at all kVp values tested
6. AEC System or Manual Exposure Performance Assessment
Optical density or signal range acceptable
7. Receptor Speed Uniformity
Screen speed uniformity or digital receptor uniformity acceptable
8. Breast Entrance Exposure, Average Glandular Dose and Exposure Reproducibility
Exposure reproducibility is within acceptable limits
| Average glandular dose to a 4.2 cm thick breast is ≤ 3 mGy (300 mrad) |  |  |  |  | #N/A | mGy
9. Image Quality Evaluation
Phantom image quality is acceptable
| Phantom type: | ACR Mammography Accreditation Phantom |  |  | Mini Phantom
| Phantom image quality scores: |  | Fibers | Specks |  | Masses
10. Artifact Evaluation
Artifacts were not apparent or not significant:
Artifacts identified:
11. Localization Accuracy Test
Localization and sampling accurate/object captured
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M15: Click in boxes to use drop-down lists
Note on L37: Click in boxes to use drop-down lists

## Tech QC Eval
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY (continued)
Evaluation of Site's Technologist QC Program
|  | Frequency | PASS/FAIL
1. | Localization Accuracy Test | Daily
2. | Darkroom Cleanliness (NA if digital used) | Daily | NA
3. | Processor Quality Control (NA if digital used) | Daily | NA
4. | Phantom Images | Weekly
5. | Screen Cleanliness (NA if digital used) | Weekly | NA
6. | Viewboxes and Viewing Conditions (NA if digital used) | Weekly | NA
7. | Hardcopy Output Quality (if hardcopy produced from digital data) | Monthly | NA
8. | Visual Checklist | Monthly
9. | Analysis of Fixer Retention in Film (NA if digital used) | Quarterly | NA
10. | Compression | Semi-annually
11. | Repeat Analysis | Semi-annually
12. | Screen-Film Contact (NA if digital used) | Semi-annually | NA
13. | Darkroom Fog (NA if digital used) | Semi-annually | NA
14. | Zero Alignment Test (if required by manufacturer) | Before each patient
15. | Any additional tests required by manufacturer | As required by manufacturer
Medical Physicist's Recommendations for Quality Improvement
Comments:
Note on J8: Click in boxes to use drop-down lists

## Sheet1
1 | [11 empty] | Revision 1.2-20250904 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [22 empty] | All: | Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5 | [5 empty] | Stereotactic Breast Biopsy System Compliance Inspection | [10 empty] | Stereotactic Breast Biopsy System Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  | Change | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [5 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah | kV Set | mAs Set | Add Filt | Target | Filter | kVp | ms | Dose (mGy) | Dose rate (mGy/s) | HVL
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility: |  |  |  | 24 | 50 | 0 | Mo | Mo
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department: |  |  |  | 24 | 50 | 0 | Mo | Mo
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division: |  |  |  | 25 | 50 | 0 | Mo | Mo
13 |  |  | Survey ID: |  |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: |  |  |  |  | Survey ID: |  |  |  | 25 | 50 | 0 | Mo | Mo
14 | [13 empty] | Room Number: |  |  |  | Accession Number: |  |  |  |  | Room Number: |  |  |  | 26 | 50 | 0 | Mo | Mo
15 |  | X-Ray Generator | [20 empty] | Site Number: |  |  |  | 28 | 20 | 0 | Mo | Mo
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [9 empty] | Authorized Use: |  |  |  | 28 | 50 | 0 | Mo | Mo
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation: |  |  |  | 28 | 50 | 0 | Mo | Mo
18 |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number: |  |  |  |  | Number of X-ray Tubes: |  |  |  | 28 | 50 | 0 | Mo | Mo
19 | [13 empty] | Max kVp: |  |  |  | Max mA: |  |  |  |  | Accesssion Number: |  |  |  | 28 | 50 | 0 | Mo | Mo
20 |  | X-Ray Tube 1 | [20 empty] | X-Ray Generator |  |  |  | 28 | 100 | 0 | Mo | Mo
21 |  |  | Tube Designation/Use: |  |  |  |  | Target: | Mo |  |  |  | X-Ray Tube 1 |  |  |  |  | Target: | Mo | Mo |  |  | Manufacturer: |  |  |  | 28 | 300 | 0 | Mo | Mo
22 |  |  | Manufacture Date: | [10 empty] | Tube Designation/Use: | [8 empty] | Model: |  |  |  | 30 | 50 | 0 | Mo | Mo
23 |  | Insert | [5 empty] | Filter: | Mo |  |  |  |  | Manufacture Date: | [8 empty] | Max kVp: |  |  |  | 32 | 50 | 0 | Mo | Mo
24 |  |  | Manufacturer: | [9 empty] | Insert |  |  |  |  | Filter: | Mo | Mo |  |  | Manufacture Date: |  |  |  | 32 | 50 | 0 | Mo | Mo
25 |  |  | Model: | [10 empty] | Manufacturer: | [8 empty] | Serial Number: |  |  |  | 34 | 50 | 0 | Mo | Mo
26 |  |  | Serial Number: |  |  |  | Focal Spot Sizes (mm) | [6 empty] | Model: | [8 empty] | Max mA: |  |  |  | 34 | 50 | 0 | Mo | Mo
27 |  | Housing | [5 empty] | Large: | [5 empty] | Serial Number: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | X-Ray Tube 1
28 |  |  | Manufacturer: |  |  |  |  | Small: |  |  |  |  | Housing |  |  |  |  | Large: |  |  |  |  | Tube Designation/Use:
29 |  |  | Model: |  |  |  | Filtration | Inherent: | [5 empty] | Manufacturer: |  |  |  | Small: |  |  |  |  | Manufacture Date:
30 |  |  | Serial Number: |  |  |  |  | Added: | [5 empty] | Model: | [8 empty] | Manufacturer:
31 | [13 empty] | Serial Number: |  |  |  | Filtration |  |  |  |  | Model:
32 | [17 empty] | Inherent: |  |  |  |  | Serial Number:
33 | [5 empty] | Technique Chart | [11 empty] | Added: |  |  |  |  | Manufacturer:
34 | [22 empty] | Model:
35 | Compressed | Fatty Breast |  |  | 50% Fatty - |  |  | Dense Breast | [14 empty] | Serial Number:
36 | Breast |  |  |  | 50% Dense Breast | [17 empty] | Large:
37 | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | [6 empty] | Inspection Results | [5 empty] | Small:
38 | <3 cm | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Target 1: |  | Change | Mo
39 | 3 to 5 cm | [11 empty] | DHEC Registration sticker is present, clearly visible and legible | [9 empty] | Target 2:
40 | 5 to 7 cm | [11 empty] | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [9 empty] | Filter 1: |  | Change | Mo
41 | >7 cm | [11 empty] | Radiation warning label posted on the generator control panel | [9 empty] | Filter 2:
42 | [12 empty] | Operator manuals are available. | [9 empty] | Inherent:
43 | [5 empty] | Inspection Results | [6 empty] | Monthly radiation monitoring reports are posted. | [9 empty] | Added
44 | Rule Number | [8 empty] | Compliance | [12 empty] | Technique Chart
45 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Mammography Unit Assembly Evaluation | [5 empty] | Fatty Breast
46 | DHEC RHB 11.2.3 |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  |  | Free-standing unit is mechanically stable | [9 empty] | Target/Filter:
47 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD |  |  | All moving parts move smoothly, without obstructions to motion | [9 empty] | kVp:
48 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | All locks and detents work properly | [9 empty] | Density:
49 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | Needle holder and guides are firmly attached and support the needle | [9 empty] | Target/Filter:
50 | [5 empty] | Mammography Unit Assembly Evaluation | [6 empty] | Image receptor, compression plate and biopsy window are free from vibration | [9 empty] | kVp:
51 |  |  | Free-standing unit is mechanically stable | [6 empty] | TBD |  |  | Image receptor is held securely by assembly in any orientation | [9 empty] | Density:
52 |  |  | All moving parts move smoothly, without obstructions to motion | [6 empty] | TBD |  | 1 | Compressed breast thickness scale accurate to +0.5 cm, reproducible to +2 mm | [9 empty] | Target/Filter:
53 |  |  | All locks and detents work properly | [6 empty] | TBD |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [9 empty] | kVp:
54 |  |  | Needle holder and guides are firmly attached and support the needle | [6 empty] | TBD |  |  | Operator technique control charts are posted | [9 empty] | Density:
55 |  |  | Image receptor, compression plate and biopsy window are free from vibration | [6 empty] | TBD |  |  | Operator protected during exposure by adequate radiation shielding | [9 empty] | Target/Filter:
56 |  |  | Image receptor is held securely by assembly in any orientation | [6 empty] | TBD |  |  | All indicator lights working properly | [9 empty] | kVp:
57 |  |  | Compressed breast thickness scale accurate to 0.5 cm, reproducible to 2 mm | [6 empty] | YES |  |  | Localization accuracy test passes | [9 empty] | Density:
58 |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [6 empty] | TBD | [12 empty] | 50/50 Breast
59 |  |  | Operator technique control charts are posted | [6 empty] | TBD | [12 empty] | Target/Filter:
60 |  |  | Operator protected during exposure by adequate radiation shielding | [6 empty] | TBD | [6 empty] | MQSA Equipment Standards | [5 empty] | kVp:
61 |  |  | All indicator lights working properly | [6 empty] | TBD |  |  | Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position | [9 empty] | Density:
62 |  |  | Localization accuracy test passes | [6 empty] | TBD |  |  | Tube-receptor assembly locking mechanism does not fail during power interruption | [9 empty] | Target/Filter:
63 | [11 empty] | 3 | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [9 empty] | kVp:
64 | [12 empty] | Moving grids are present for all image receptor sizes | [9 empty] | Density:
65 | [11 empty] | 3 | Magnification can be done without grid | [9 empty] | Target/Filter:
66 | [12 empty] | Beam limiting device allows useful beam to extend beyond chest wall edge of receptor | [9 empty] | kVp:
67 | Date: | [8 empty] | Inspector: | Eugene Mah | 3 | Average illumination of the light field is at least 160 lux | [9 empty] | Density:
68 | Room Number | [8 empty] | Survey ID: |  | 3 | Magnification mode is available for non-interventional problem solving procedures | [9 empty] | Target/Filter:
1 | [10 empty] | Medical University of South Carolina | 3 | System provides magnification factor between 1.4 – 2.0 | [9 empty] | kVp:
2 | [5 empty] | Inspection Results |  |  |  |  | Stereotactic Breast Biopsy System Compliance Inspection |  | Prior indication of selected focal spot is provided | [9 empty] | Density:
3 | Rule Number |  |  |  |  | MQSA Equipment Standards | [6 empty] | Prior indication of selected target is provided | [9 empty] | Dense Breast
4 | 900.12(b)(3i) |  | Tube-receptor can be fixed into any operating position and does not move | [6 empty] | TBD |  |  | Post exposure display of target/focal spot is provided | [9 empty] | Target/Filter:
5 | 900.12(b)(3ii) |  | Tube-receptor locking mechanism does not fail during power interruption | [6 empty] | TBD |  |  | Initial power driven compression can be activated hands-free from both sides | [9 empty] | kVp:
6 | 900.12(b)(4i) |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [6 empty] | NA |  |  | Fine adjustment of compression can be activated from both sides | [9 empty] | Density:
7 | 900.12(b)(4ii) |  | Moving grids are present for all image receptor sizessizes | [6 empty] | TBD |  |  | Compression paddles match the sizes of all image receptors | [9 empty] | Target/Filter:
8 | 900.12(b)(4iii) |  | Magnification can be done without grid | [6 empty] | NA |  |  | Compression paddles do not flex by more than 1 cm | [9 empty] | kVp:
9 | 900.12(b)(5i) |  | Beam limiting device allows useful beam to extend beyond chest wall edge | [6 empty] | TBD |  |  | Compression paddles are flat and parallel to the breast support table | [9 empty] | Density:
10 | 900.12(b)(5ii) |  | Average illumination of the light field is at least 160 lux | [6 empty] | NA |  |  | Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor | [9 empty] | Target/Filter:
11 | 900.12(b)(6i) |  | Magnification mode available for non-interventional problem solving | [6 empty] | NA |  |  | Chest wall edge may be bent upwards for patient comfort and does not appear in the image | [9 empty] | kVp:
12 | 900.12(b)(6ii) |  | System provides magnification factor between 1.4 – 2.0 | [6 empty] | NA |  |  | Manual selection of mAs or mA and/or time is available | [9 empty] | Density:
13 | 900.12(b)(7i) |  | Prior indication of selected focal spot is provided | [6 empty] | TBD |  |  | Prior indication of selected technique factors is provided | [9 empty] | Target/Filter:
14 | 900.12(b)(7i) |  | Prior indication of selected target is provided | [6 empty] | TBD |  |  | Post exposure display of technique factors is provided for AEC exposures | [9 empty] | kVp:
15 | 900.12(b)(7iii) |  | Post exposure display of target/focal spot is provided | [6 empty] | TBD |  | 3 | AEC mode is operable in all combinations of equipment configuration | [9 empty] | Density:
16 | 900.12(b)(8)(iB) |  | Initial power driven compression can be activated hands-free from both sides | [6 empty] | TBD |  | 3 | AEC sensor positioning is flexible
17 |  |  | Fine adjustment of compression can be activated from both sides | [6 empty] | TBD |  | 3 | AEC sensor position and size is indicated on the compression paddle | [9 empty] | MGD: |  |  | #N/A
18 | 900.12(b)(8)(iiA) |  | Compression paddles match the sizes of all image receptors | [6 empty] | TBD |  | 3 | Selected position of the AEC sensor is clearly indicated | [9 empty] | mGy/mAs:
19 | 900.12(b)(8)(iiB) |  | Compression paddles do not flex by more than 1 cm | [6 empty] | TBD |  |  | AEC density control setting is provided | [9 empty] | mGy/s:
20 | 900.12(b)(8)(iiC) |  | Compression paddles are flat and parallel to the breast support table | [6 empty] | TBD |  | 3 | X-ray film used is appropriate for mammography | [9 empty] | Res Par 512:
21 | 900.12(b)(8)(iiD) |  | Chest wall edge of the compression paddle is straight and parallel to receptor | [6 empty] | TBD |  | 3 | Intensifying screens used are appropriate for mammography and matched to the film used | [9 empty] | Res Perp 512:
22 | 900.12(b)(8)(iiE) |  | Chest wall edge may be bent upwards and does not appear in the image | [6 empty] | TBD |  | 3 | Film processing chemicals are appropriate for the film being used | [9 empty] | Res Par 1024:
23 | 900.12(b)(9i) |  | Manual selection of mAs or mA and/or time is available | [6 empty] | TBD |  |  | Lights for film illumination i.e. hot lights are available | [9 empty] | Res Perp 1024:
24 | 900.12(b)(9ii) |  | Prior indication of selected technique factors is provided | [6 empty] | TBD |  |  | Film masking devices are available to all interpreting physicians
25 | 900.12(b)(9iii) |  | Post exposure display of technique factors is provided for AEC exposures | [6 empty] | TBD | [12 empty] | Image Quality (Acq Unit)
26 | 900.12(b)(10i) |  | AEC mode is operable in all combinations of equipment configuration | [6 empty] | NA | [12 empty] | kV: |  |  | 28
27 | 900.12(b)(10ii) |  | AEC sensor positioning is flexible | [6 empty] | NA | [6 empty] | Mammography Phototimer Technique Chart | [5 empty] | mA: |  |  | 0
28 |  |  | AEC sensor position and size is indicated on the compression paddle | [6 empty] | NA | [12 empty] | Exp Ind: |  |  | 0
29 |  |  | Selected position of the AEC sensor is clearly indicated | [6 empty] | NA |  | Compressed | Fatty Breast |  |  | 50% Fatty - |  |  | Dense Breast |  |  |  | Fibers:
30 | 900.12(b)(10iii) |  | AEC density control setting is provided | [6 empty] | TBD |  | Breast |  |  |  | 50% Dense Breast | [6 empty] | Specks:
31 | 900.12(b)(11) |  | X-ray film used is appropriate for mammography | [6 empty] | NA |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  | Masses:
32 | 900.12(b)(12) |  | Intensifying screens used are appropriate for mammography and matched to film | [6 empty] | NA |  | <3 cm | [10 empty] | Average:
33 | 900.12(b)(13) |  | Film processing chemicals are appropriate for the film being used | [6 empty] | NA |  | 3 to 5 cm | [10 empty] | SD:
34 | 900.12(b)(14) |  | Lights for film illumination i.e. hot lights are available | [6 empty] | TBD |  | 5 to 7 cm | [10 empty] | kV: |  |  | 28
35 | 900.12(b)(15) |  | Film masking devices are available to all interpreting physicians | [6 empty] | TBD |  | >7 cm | [10 empty] | mA: |  |  | 138.9
36 | [22 empty] | Exp Ind: |  |  | 0
37 | [11 empty] | Compressed | Fatty Breast |  |  | 50% Fatty - |  |  | Dense Breast |  |  |  | Fibers:
38 | [11 empty] | Breast |  |  |  | 50% Dense Breast | [6 empty] | Specks:
39 | [11 empty] | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  | Masses:
40 | [11 empty] | <3 cm | [10 empty] | Average:
41 | [11 empty] | 3 to 5 cm | [10 empty] | SD:
42 | [11 empty] | 5 to 7 cm
43 | [11 empty] | >7 cm
44
45 | [16 empty] | Inspection Results | [5 empty] | Comment Page 1
46 | [11 empty] | Compression Force Indicator
47 | [13 empty] | Indicated | Measured | After 20s | Slip
48 | [12 empty] | Power
49 | [12 empty] | Manual
50 | [12 empty] | Criteria: | Minimum compression force – 25 lbs
51 | [13 empty] | Maximum compression force – 45 lbs
52
53 | [11 empty] | Compression Thickness Indicator
54 | [12 empty] | BR-12 (cm) | Indicated | Difference
55 | [12 empty] | 1
56 | [12 empty] | 2
57 | [12 empty] | 4 |  |  | Mean:
58 | [12 empty] | 4 |  |  | StDev:
59 | [12 empty] | 4
60 | [12 empty] | 4 |  |  |  | Criteria: | Difference between indicated and actual thickness is <= 2mm
61 | [12 empty] | 6
62 | [12 empty] | 8
63
64 | [12 empty] | Paddle flex is less than 1 cm |  |  |  | Compression thickness indicator:
65 | [12 empty] | Overall unit assembly
66
67 | Date: | [8 empty] | Inspector: | Eugene Mah | Artifact Evaluation
68 | Room Number | [8 empty] | Survey ID: |  |  |  |  | kV: | 28
1 | [10 empty] | Medical University of South Carolina |  | Target: | Mo | mAs:
2 | [5 empty] | Inspection Results |  |  |  |  | Stereotactic Breast Biopsy System Compliance Inspection |  | Filter: | Mo | EI:
3 | Compression Force Indicator | [11 empty] | Focal spot: | Large | Acceptable:
4 |  |  | Indicated | Measured | After 20s | Slip
5 |  | Power | [10 empty] | Comments:
6 |  | Manual | [5 empty] | Acceptable: |  |  |  |  | Do not exceed 390 chars
7 |  | Criteria: | Minimum compression force – 25 lbs | [9 empty] | New:
8 |  |  | Maximum compression force – 45 lbs
9 | [11 empty] | X-ray Field/Image Receptor Alignment
10 | Compression Thickness Indicator | [10 empty] | 65 | SID (cm)
11 |  | BR-12 (cm) | Indicated | Difference |  |  |  | Paddle flex is less than 1 cm: | TBD |  |  |  | Enter deviation in cm
12 |  | 1 | [5 empty] | Overall unit assembly: | TBD | [5 empty] | % Deviation
13 |  | 2 | [5 empty] | Compression thickness indicator: | TBD |  |  |  | Left:
14 |  | 4 |  |  | Mean: | [7 empty] | Right:
15 |  | 4 |  |  | StDev: | [7 empty] | Anterior:
16 |  | 4 |  |  | Criteria: | Difference between indicated and actual thickness is <= 2mm | [6 empty] | Chest Wall:
17 |  | 4
18 |  | 6 | [10 empty] | Criteria: | X-ray field does not extend beyond image receptor by more than 5 mm on any side
19 |  | 8 | [11 empty] | Biopsy window is centered in image
20 | [12 empty] | Biopsy window centered in image
21 | Artifact Evaluation
22 | [11 empty] | System Resolution |  |  | 512 |  | 1024
23 |  | Target: | Mo | [11 empty] | Par | Perp | Par | Perp
24 |  | Filter: | Mo | [8 empty] | 28 | kV | lp/mm:
25 |  | Focal spot: | Large | [8 empty] | 65 | mAs | Previous:
26 |  | kV: | 28 | [10 empty] | Acceptable:
27 |  | mAs: | [11 empty] | Criteria: | Note any significant degradation from previous measurement
28 |  | EI:
29 |  | Acceptable: | [9 empty] | AEC Thickness Tracking
30 |  | Criteria: | No visible artifacts in the images | [8 empty] | AEC Mode: | Auto-time |  | AEC Position:
31 |  | Comments: | [10 empty] | Thickness |  |  |  |  | Reported
32 | [12 empty] | (cm) | Filter | kV | mAs | ROI | MGD (mGy) | Deviation
33 | [12 empty] | 2
34 | [12 empty] | 4
35 | X-ray Field/Image Receptor Alignment | [11 empty] | 4 | 0 | 0
36 | [12 empty] | 4 | 0 | 0
37 |  | % Deviation | [10 empty] | 4 | 0 | 0
38 | Left: |  |  |  | Biopsy window is centered in image | [7 empty] | 6
39 | Right: |  |  | Criteria: | X-ray field does not extend beyond image receptor by more than 5 mm on any side | [7 empty] | 8
40 | Anterior: |  |  |  | Biopsy window is centered in image | [10 empty] | Average:
41 | Chest Wall: | [11 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
42 | Acceptable:
43 | System Resolution |  | 512 |  | 1024
44 |  |  | Perp | Par | Perp | Par | [5 empty] | AEC Density Control
45 |  | lp/mm: |  |  |  |  | Acceptable: |  |  |  |  | AEC Mode: | Auto-time |  | AEC Position: | 0 | Reported
46 |  | Criteria: | Note any significant degradation from previous measurement | [9 empty] | Density | kV | mAs | ROI | MGD (mGy) | Change |  |  | Low | High
47 | Detector Uniformity | [11 empty] | -3 | [7 empty] | 0.5 | 0.61
48 | kV: | 28 | mA: | 0 | [8 empty] | -2 | [7 empty] | 0.63 | 0.77
49 |  | ROI | SD | SNR | Deviation | [7 empty] | -1 | [7 empty] | 0.77 | 0.94
50 | Center | [11 empty] | 0
51 | ROI 1 | [11 empty] | 1 | [7 empty] | 1.04 | 1.27
52 | ROI 2 | [11 empty] | 2 | [7 empty] | 1.17 | 1.43
53 | ROI 3 | [11 empty] | 3 | [7 empty] | 1.31 | 1.6
54 | ROI 4 | [11 empty] | 4 | [7 empty] | 1.44 | 1.76
55 | Average | [11 empty] | Criteria: | Change in pixel value at each density step is within specified ranges
56 | SD
57 | [11 empty] | Detector Uniformity
58 | Criteria: | SNR values are within 15% of the SNR in the center | [10 empty] | kV: | 28
59 | [12 empty] | mAs: | 0
60 | [13 empty] | ROI | SD | SNR | Deviation
61 | [12 empty] | Center
62 | [12 empty] | ROI 1
63 | [12 empty] | ROI 2
64 | [12 empty] | ROI 3
65 | [12 empty] | ROI 4
66 | [12 empty] | Average
67 | Date: | [8 empty] | Inspector: | Eugene Mah |  | SD
68 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina |  | Criteria: | SNR values are within 15% of the SNR in the center
2 | [5 empty] | Inspection Results |  |  |  |  | Stereotactic Breast Biopsy System Compliance Inspection
3 | AEC Thickness Tracking
4 | AEC Mode: | Auto-time |  | AEC Position:
5 |  | Thickness |  |  |  |  | Reported |  |  |  |  | Mean Glandular Dose
6 |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | Phantom: |  |  | kV: | 28
7 |  | 2 | [10 empty] | Target/Filter: | Mo/Mo |  | mA:
8 |  | 4 | [13 empty] | Reported |  | Air KERMA
9 |  | 4 | 0 | 0 | [9 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
10 |  | 4 | 0 | 0 | [15 empty] | ESE (mGy):
11 |  | 4 | 0 | 0 | [15 empty] | DGN (mrad/R): | #N/A
12 |  | 6 | [17 empty] | MGD (mGy): | #N/A
13 |  | 8 | [17 empty] | Prev MGD (mGy):
14 |  |  |  |  | Average: | [7 empty] | Average: | [6 empty] | % Change: | #N/A
15 | [12 empty] | Coeff of Var:
16 |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
17 | [12 empty] | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A
18 | AEC Density Control | [12 empty] | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | #N/A
19 | AEC Mode: | Auto-time |  | AEC Position: | 0 |  | kV:
20 |  |  |  |  | Reported | [6 empty] | kVp Accuracy and Output
21 |  | Density | mAs | EI | MGD (mGy) | Change | Acceptable |  |  |  |  | Meter: |  |  | Calibration Date:
22 |  | -3 | [9 empty] | Serial #: |  |  | Calibration Due:
23 |  | -2
24 |  | -1 | [14 empty] | Measured Values
25 |  | 0 | [10 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
26 |  | 1 | [10 empty] | W | Ag | 24 | 50
27 |  | 2 | [10 empty] | W | Ag | 25 | 50
28 |  | 3 | [10 empty] | W | Ag | 26 | 50
29 |  | 4 | [10 empty] | W | Ag | 28 | 50
30 |  | Criteria: | Change in pixel value at each density step is within specified ranges | [9 empty] | W | Ag | 30 | 50
31 | [12 empty] | W | Ag | 32 | 50
32 | Mean Glandular Dose | [11 empty] | W | Ag | 34 | 50
33 |  | Phantom: |  | kV: | 28
34 |  | Target/Filter: | Mo/Mo | mA: | [8 empty] | Criteria: | kV accuracy is <= 5%
35 |  |  |  |  | Reported |  | Air KERMA |  |  |  |  | kV, Exposure, and Output Reproducibility
36 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
37 | [8 empty] | ESE (mGy): | [7 empty] | Measured Values
38 | [8 empty] | DGN (mrad/R): | #N/A |  |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
39 | [8 empty] | MGD (mGy): | #N/A |  |  | W | Ag | 28 | 50
40 | [8 empty] | Prev MGD (mGy): |  |  |  | W | Ag | 28 | 50
41 |  | Average: | [6 empty] | % Change: | #N/A |  |  | W | Ag | 28 | 50
42 |  | Coeff of Var: | [10 empty] | W | Ag | 28 | 50
43 |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A | [5 empty] | Mean:
44 |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | #N/A | [5 empty] | Std Dev:
45 | [15 empty] | Coeff of Var:
46 | Signal to Noise Ratio | [14 empty] | Previous:
47 |  | kV: |  |  |  | Current | Previous | Change | Acceptable |  |  |  | Criteria: | Coefficient of variation <= 5%
48 |  | mA: |  |  | SNR: | [8 empty] | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
49 |  | Target: |  |  | CNR:
50 |  | Filter: | [9 empty] | Exposure Linearity
51 |  | Criteria: | SNR must be >= 40 | [13 empty] | Measured Values
52 |  |  | CNR does not change by more than 15% | [9 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
53 | [12 empty] | W | Ag | 28 | 20
54 | Image Quality Evaluation | [11 empty] | W | Ag | 28 | 50
55 | [12 empty] | W | Ag | 28 | 100
56 |  |  | Acq Workstation | [9 empty] | W | Ag | 28 | 320
57 |  |  | Normal | [9 empty] | Criteria: | Coefficient of linearity is <= 10% |  |  |  |  | Coeff of Linearity:
58 |  | kV: | 28
59 |  | mA: | 0 | [9 empty] | Unit Can Maintain Output for 3 sec at 28 kVp
60 |  | Exposure Index: | 0
61 |  | Fibers: | [9 empty] | Half Value Layer
62 |  | Specks: | [10 empty] | Target/Filter | W/Ag | W/Ag | W/Ag | W/Ag | W/Ag
63 |  | Masses: | [10 empty] | kV | 24 | 25 | 28 | 32 | 34
64 | [12 empty] | Al (mm) | Air Kerma (mGy)
65 | [12 empty] | 0
66 | [12 empty] | 0.2
67 | Date: | [8 empty] | Inspector: | Eugene Mah |  | 0.3
68 | Room Number | [8 empty] | Survey ID: |  |  | 0.4
1 | [10 empty] | Medical University of South Carolina |  | 0.5
2 | [5 empty] | Inspection Results |  |  |  |  | Stereotactic Breast Biopsy System Compliance Inspection |  | 0
3 | Meter: |  | SN: |  |  | Calibration Date: | [6 empty] | HVL
4 | [5 empty] | Calibration Due: | [6 empty] | 0 mm
5 | kVp Accuracy |  |  |  |  | kVp and Output Reproducibility | [6 empty] | Min | 0.24 | 0.25 | 0.28 | 0.32 | 0.34
6 | Target: | W | Filter: | Ag |  | Target: | W | Filter: | Ag |  |  |  | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.46
7 | mAs: | 50 |  |  |  | mAs: | 50
8 | Indicated | Measured |  |  |  | Indicated | Measured | [5 empty] | Criteria: | HVL is between the minimum and maximum limits
9 | kVp | kVp | Error (%) |  |  | kVp | kVp | mGy | mGy/mAs | mGy/s
10 | 24 |  |  |  |  | 28 | [5 empty] | Image Quality Evaluation
11 | 25 | [10 empty] | Acquisition Unit |  | 1 | Mini phantom used? |  |  | Previous
12 | 26 | [12 empty] | 512 | 1024 |  |  | 512 | 1024
13 | 28 | [11 empty] | kV: | 28 | 28
14 | 30 |  |  |  |  | Mean: | [6 empty] | mAs: | 0 | 138.9
15 | 32 |  |  |  |  | Std Dev: | [6 empty] | Exposure Index: | 0 | 0
16 | 34 |  |  |  |  | Coeff of Var: | [6 empty] | Fibers:
17 |  | Acceptable: |  |  |  | Acceptable: | [6 empty] | Specks:
18 | [5 empty] | Criteria: | Coefficient of variation <= 2% | [5 empty] | Masses:
19 | [6 empty] | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target | [5 empty] | Average PV:
20 |  | Criteria: | kVp accurate to within 5% of indicated |  |  |  | Minimum acceptable output is 2 mGy/s at 28 kVp for W target | [5 empty] | SD:
21 | [12 empty] | Criteria: | Must see at least 3 fibers, 3 specks, 2.5 masses
22 | Exposure Linearity | [14 empty] | Previous Year Comments
23 | Target: | W | Filter: | Ag | [8 empty] | Comments:
24 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
25 | Set mAs | kVp | mGy | mGy/mAs | mGy/s | [7 empty] | Additional Comments:
26 | 20 | [11 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
27 | 50 | [11 empty] | Additional Comments:
28 | 100 | [11 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
29 | 320 | [11 empty] | Additional Comments:
30 |  |  | Coeff of Linearity: | [9 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
31 |  | Criteria: | Coefficient of linearity is <= 10% | [9 empty] | Additional Comments:
32 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
33 | Half Value Layer | [11 empty] | Additional Comments:
34 | Target/Filter | W/Ag | W/Ag | W/Ag | W/Ag | [7 empty] | Do not exceed 390 chars |  | 0
35 | kV | 24 | 25 | 28 | 32 | [7 empty] | Additional Comments:
36 | HVL | [11 empty] | Do not exceed 390 chars |  | 0
37 | 0 mm
38 | Min | 0.24 | 0.25 | 0.28 | 0.32
39 | Max | 0.36 | 0.37 | 0.4 | 0.44
40 | Acceptable:
41 |  | Criteria: | HVL is between the minimum and maximum limits
42
43 | [5 empty] | Comments
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:

## Sheet2
Measurements
Set kV (kV) | Set mAs (mAs) | Target | Filter | Calibration | Tube voltage (kV) | Exposure time (ms) | Exposure (mGy) | Exposure rate (mGy/s) | HVL (mm Al) | Calc mA
24 | 50 | W | Ag | Affirm Prone W/Ag
25 | 50 | W | Ag | Affirm Prone W/Ag
26 | 50 | W | Ag | Affirm Prone W/Ag
28 | 20 | W | Ag | Affirm Prone W/Ag
28 | 50 | W | Ag | Affirm Prone W/Ag
28 | 50 | W | Ag | Affirm Prone W/Ag
28 | 50 | W | Ag | Affirm Prone W/Ag
28 | 50 | W | Ag | Affirm Prone W/Ag
28 | 100 | W | Ag | Affirm Prone W/Ag
28 | 320 | W | Ag | Affirm Prone W/Ag
30 | 50 | W | Ag | Affirm Prone W/Ag
32 | 50 | W | Ag | Affirm Prone W/Ag
34 | 50 | W | Ag | Affirm Prone W/Ag
36 | 50 | W | Ag | Affirm Prone W/Ag
38 | 50 | W | Ag | Affirm Prone W/Ag
28 | 50 | W | Al | Affirm Prone W/Al
30 | 50 | W | Al | Affirm Prone W/Al
32 | 50 | W | Al | Affirm Prone W/Al
34 | 50 | W | Al | Affirm Prone W/Al
36 | 50 | W | Al | Affirm Prone W/Al
38 | 50 | W | Al | Affirm Prone W/Al

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
| kVp | [9 empty] | kVp | [9 empty] | kVp
HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
0.23 | 116 | [8 empty] | 0.23 | [9 empty] | 0.23
0.24 | 121 | 124 | [7 empty] | 0.24 | [9 empty] | 0.24
0.25 | 126 | 129 | 131 | [6 empty] | 0.25 | [9 empty] | 0.25
0.26 | 130 | 133 | 135 | 138 | [5 empty] | 0.26 | [9 empty] | 0.26
0.27 | 135 | 138 | 140 | 142 | 143 |  |  |  |  | 0.27 | [9 empty] | 0.27
0.28 | 140 | 142 | 144 | 146 | 147 | 149 |  |  |  | 0.28 |  |  | 149 | 151 | 154 |  |  |  |  | 0.28 |  |  | 150 | 155 | 159
0.29 | 144 | 146 | 148 | 150 | 151 | 153 | 154 |  |  | 0.29 |  |  | 154 | 156 | 158 | 159 |  |  |  | 0.29 |  |  | 155 | 160 | 164 | 168
0.3 | 149 | 151 | 153 | 155 | 156 | 157 | 158 | 159 |  | 0.3 |  |  | 158 | 160 | 162 | 162 | 163 |  |  | 0.3 |  |  | 160 | 164 | 168 | 172 | 176
0.31 | 154 | 156 | 157 | 159 | 160 | 161 | 162 | 163 | 164 | 0.31 |  |  | 163 | 164 | 166 | 166 | 167 | 167 |  | 0.31 |  |  | 165 | 168 | 172 | 174 | 180 | 182
0.32 | 158 | 160 | 162 | 163 | 164 | 166 | 167 | 168 | 168 | 0.32 |  |  | 167 | 169 | 171 | 171 | 171 | 172 | 172 | 0.32 |  |  | 169 | 173 | 177 | 181 | 184 | 186 | 188
0.33 | 163 | 165 | 166 | 168 | 169 | 170 | 171 | 173 | 173 | 0.33 |  |  | 171 | 173 | 175 | 176 | 176 | 176 | 176 | 0.33 |  |  | 174 | 178 | 181 | 185 | 188 | 190 | 192
0.34 | 168 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 0.34 |  |  | 176 | 178 | 179 | 179 | 180 | 180 | 180 | 0.34 |  |  | 179 | 183 | 186 | 190 | 193 | 195 | 196
0.35 |  | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 0.35 |  |  | 180 | 181 | 183 | 183 | 184 | 185 | 185 | 0.35 |  |  | 184 | 187 | 190 | 194 | 197 | 199 | 201
0.36 |  |  | 179 | 181 | 182 | 183 | 184 | 185 | 185 | 0.36 |  |  | 185 | 186 | 187 | 187 | 188 | 188 | 189 | 0.36 |  |  | 189 | 192 | 195 | 198 | 201 | 204 | 205
0.37 |  |  |  | 185 | 186 | 187 | 188 | 189 | 190 | 0.37 |  |  | 189 | 190 | 191 | 191 | 192 | 193 | 193 | 0.37 |  |  | 193 | 196 | 199 | 202 | 205 | 207 | 209
0.38 |  |  |  |  | 190 | 191 | 192 | 193 | 194 | 0.38 |  |  | 193 | 194 | 196 | 196 | 197 | 197 | 197 | 0.38 |  |  | 198 | 201 | 204 | 207 | 209 | 211 | 213
0.39 |  | 174 178: 196; 175 179: 197; 176 180: 198; 177 181: 198; 0.34 0.35: 0.39 |  |  | 198 | 199 | 200 | 200 | 201 | 201 | 202 | 0.39 |  |  | 203 | 206 | 208 | 211 | 214 | 216 | 217
0.4 |  | 175 179: 201; 176 180: 202; 177 181: 203; 0.34 0.35: 0.4 |  |  | 202 | 203 | 204 | 204 | 205 | 205 | 206 | 0.4 |  |  | 208 | 211 | 213 | 216 | 218 | 220 | 221
0.41 |  | 176 180: 206; 177 181: 207; 0.34 0.35: 0.41 |  |  | 206 | 207 | 208 | 208 | 209 | 209 | 210 | 0.41 |  |  | 213 | 215 | 217 | 220 | 222 | 224 | 225
0.42 |  | 177 181: 211; 0.34 0.35: 0.42 |  |  | 211 | 211 | 212 | 212 | 213 | 213 | 214 | 0.42 |  |  | 218 | 220 | 222 | 224 | 226 | 228 | 229
Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast | [12 empty] | Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
| kVp | [12 empty] | kVp
HVL | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | HVL | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39
0.3 | 152 | 157 | 163 | 166 | 170 | 173 | 175 | 177 | 179 | 182 | 184 | 187 | 0.3 | 222 | 226 | 229 | 231 | 234 | 236 | 239 | 241 | 244 | 246 | 248 | 250 | 252
0.325 | 163 | 169 | 174 | 177 | 181 | 183 | 186 | 188 | 190 | 192 | 195 | 197 | 0.325 | 233 | 236 | 239 | 242 | 244 | 246 | 248 | 251 | 253 | 256 | 258 | 260 | 262
0.35 | 175 | 180 | 185 | 188 | 191 | 194 | 196 | 198 | 200 | 202 | 205 | 207 | 0.35 | 244 | 247 | 249 | 252 | 254 | 256 | 258 | 260 | 263 | 265 | 267 | 269 | 271
0.375 | 186 | 191 | 196 | 199 | 202 | 205 | 207 | 209 | 211 | 213 | 215 | 218 | 0.375 | 254 | 257 | 260 | 262 | 264 | 266 | 268 | 270 | 273 | 275 | 277 | 279 | 281
0.4 | 198 | 203 | 207 | 210 | 213 | 215 | 217 | 219 | 221 | 223 | 226 | 228 | 0.4 | 265 | 267 | 270 | 272 | 274 | 276 | 278 | 280 | 282 | 284 | 286 | 288 | 290
0.425 | 209 | 214 | 218 | 221 | 224 | 226 | 228 | 230 | 232 | 234 | 236 | 238 | 0.425 | 275 | 278 | 280 | 282 | 284 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 300
0.45 | 221 | 226 | 230 | 232 | 235 | 237 | 238 | 240 | 242 | 244 | 246 | 248 | 0.45 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 299 | 301 | 303 | 305 | 307 | 309
0.475 | 233 | 237 | 241 | 243 | 245 | 247 | 249 | 251 | 253 | 254 | 256 | 258 | 0.475 | 296 | 298 | 3090 | 302 | 304 | 305 | 307 | 309 | 311 | 313 | 315 | 317 | 318
0.5 | 244 | 248 | 252 | 254 | 256 | 258 | 260 | 261 | 263 | 265 | 267 | 269 | 0.5 | 306 | 308 | 310 | 312 | 313 | 315 | 317 | 319 | 320 | 322 | 324 | 326 | 328
0.525 | 256 | 260 | 263 | 265 | 267 | 269 | 270 | 272 | 273 | 275 | 277 | 279 | 0.525 | 316 | 318 | 320 | 322 | 323 | 325 | 326 | 328 | 330 | 332 | 333 | 335 | 337
0.55 | 267 | 271 | 274 | 276 | 278 | 279 | 281 | 282 | 284 | 285 | 287 | 289 | 0.55 | 326 | 328 | 330 | 331 | 333 | 334 | 336 | 338 | 339 | 341 | 343 | 344 | 346
0.575 | 279 | 282 | 285 | 287 | 288 | 290 | 291 | 292 | 294 | 296 | 297 | 299 | 0.575 | 336 | 338 | 339 | 341 | 342 | 344 | 345 | 347 | 349 | 350 | 352 | 354 | 355
0.6 | 290 | 293 | 296 | 297 | 299 | 300 | 301 | 303 | 304 | 306 | 308 | 310 | 0.6 | 346 | 348 | 349 | 350 | 352 | 353 | 355 | 356 | 358 | 359 | 361 | 363 | 364
0.625 | 301 | 304 | 306 | 308 | 309 | 310 | 312 | 313 | 315 | 316 | 318 | 320 | 0.625 | 356 | 357 | 358 | 360 | 361 | 362 | 364 | 365 | 367 | 368 | 370 | 372 | 373
0.65 | 312 | 314 | 317 | 318 | 320 | 321 | 322 | 323 | 325 | 326 | 328 | 330 | 0.65 | 365 | 367 | 368 | 369 | 370 | 372 | 373 | 375 | 376 | 378 | 379 | 381 | 382
0.675 | 322 | 325 | 327 | 328 | 330 | 331 | 333 | 333 | 335 | 336 | 338 | 340 | 0.675 | 374 | 376 | 377 | 378 | 379 | 381 | 382 | 383 | 385 | 386 | 388 | 390 | 391
0.7 | 333 | 335 | 337 | 339 | 340 | 341 | 342 | 343 | 345 | 346 | 348 | 350 | 0.7 | 384 | 385 | 386 | 387 | 388 | 390 | 391 | 392 | 394 | 395 | 397 | 398 | 400
0.725 | 342 | 345 | 347 | 348 | 349 | 351 | 352 | 353 | 354 | 356 | 358 | 360 | 0.725 | 393 | 394 | 395 | 396 | 397 | 399 | 400 | 401 | 403 | 404 | 406 | 407 | 408
0.75 | 352 | 355 | 357 | 358 | 359 | 360 | 361 | 363 | 364 | 366 | 368 | 369 | 0.75 | 402 | 403 | 404 | 405 | 406 | 407 | 409 | 410 | 411 | 413 | 414 | 416 | 417
0.775 | 361 | 365 | 367 | 368 | 369 | 370 | 371 | 372 | 374 | 375 | 377 | 379 | 0.775 | 410 | 411 | 412 | 413 | 415 | 416 | 417 | 418 | 420 | 421 | 423 | 424 | 425
0.8 | 369 | 374 | 376 | 377 | 378 | 379 | 380 | 382 | 383 | 385 | 387 | 389 | 0.8 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 427 | 428 | 429 | 431 | 432 | 434
Siemens Inspiration
| PMMA thickness (mm) | 170 174: Equiv Breast thickness (mm); 171 175: HVL (mm Al)
|  |  | 0.25 | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 | Slope | Intercept
g-factor | 45 | 53 | 0.13 | 0.155 | 0.177 | 0.198 | 0.22 | 0.245 | 0.272 | 0.295 | 0.4681 | 0.01256
c-factor | 45 | 53 |  | 1.109 | 1.105 | 1.102 | 1.099 | 1.096 | 1.091 | 1.088 | -0.069286 | 1.1297
g-factor | 50 | 60 | 0.112 | 0.135 | 0.154 | 0.172 | 0.192 | 0.214 | 0.236 | 0.261 | 0.41619 | 0.007619
c-factor | 50 | 60 |  | 1.164 | 1.16 | 1.151 | 1.15 | 1.144 | 1.139 | 1.134 | -0.099286 | 1.1935
| Mo/Mo | 170 174: Mo/Rh; 171 175: W/Rh
s-factor | 1 | 170 174: 1.017; 171 175: 1.042
kV Accuracy/Output |  | 170 174: Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV^2 |  |  | Radcal kV correction for Hologic Selenia/Dimension
Target | kV | 170 174: kV^2; 171 175: mGy/mAs; 172 176: mGy/s |  |  |  | Selenia |  | Dimension
W | 24 | 170 174: 576 |  |  | Measured | W/Rh | W/Ag | W/Rh | W/Ag | W/Al
Filter | 25 | 170 174: 625
Ag | 26 | 170 174: 676
| 28 | 170 174: 784 |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| 30 | 170 174: 900 |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| 32 | 170 174: 1024 |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| 34 | 170 174: 1156 |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | Slope: | #DIV/0! | #DIV/0!
|  | Intercept: | #DIV/0! | #DIV/0!
[13 empty] | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Half Value Layer | [12 empty] | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Target | kV | 170 174: HVL; 175 179: Target; 176 180: Filter; 177 181: C |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
W | 24 | 175 179: Mo; 176 180: Mo; 177 181: 0.12 |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Filter | 25 | 175 179: Mo; 176 180: Rh; 177 181: 0.19
Ag | 28 | 175 179: Rh; 176 180: Rh; 177 181: 0.22
| 32 | 175 179: Critiera:; 176 180: Min HVL > kV/100 (no paddle)
| Slope: | 170 174: #DIV/0!; 176 180: Min HVL > kV/100 + 0.03 (with paddle)
| Intercept: | 170 174: #DIV/0!; 176 180: Max HVL < kV/100 + C
DGN | #VALUE! | [11 empty] | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Mo/Mo | Mo/Rh | 170 174: Rh/Rh; 171 175: W/Rh; 172 176: W/Ag; 173 177: Siemens Inspiration
#N/A | #N/A | 170 174: #N/A; 171 175: #N/A; 172 176: #N/A; 173 177: #VALUE!
Radcal kV Correction for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Rh | < | 27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| >= | 27.586 | -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Ag | < | 30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| >= | 30.1 | -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
Radcal kV Correction for Hologic Dimension
| kV | 171 175: kV^0; 172 176: kV^1; 173 177: kV^2; 174 178: kV^3; 175 179: kV^4
W/Rh | < | 170 174: 26.9; 171 175: 138.8867; 172 176: -10.7264; 173 177: 0.26216; 174 178: -0.00082
| >= | 26.9 | -5009.7752 | 605.732 | -27.3019 | 0.54671 | -0.0040986
W/Ag | kV | 171 175: kV^0; 172 176: kV^1; 173 177: kV^2; 174 178: kV^3; 175 179: kV^4
| < | 170 174: 28.7; 171 175: 296.3419; 172 176: -31.6292; 173 177: 1.1803; 174 178: -0.01417
| >= | 170 174: 28.7; 171 175: 4.8345; 172 176: 0.91924
W/Al |  | 171 175: kV^0; 172 176: kV^1; 173 177: kV^2; 174 178: kV^3; 175 179: kV^4
| < | 170 174: 28.7; 171 175: 49.3111; 172 176: -2.9302; 173 177: 0.07379
|  | 28.7 < kV < 30.1 | -24.875 | 1.8032
| >= | 170 174: 30.1; 171 175: -4.8346; 172 176: 1.1571
Note on A86: Radcal kV correction for W anodes uses a 5th order polynomial divided into two domains.
Note on A98: Radcal kV correction for W anodes uses a 5th order polynomial divided into two domains.

## DataPage
Sheet | SBB

## Corrected kV
GENERAL
Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
Read the true corrected kV in the adjacent cell to the right.
If that cell is light green, the corrected kV is accurate.
If the cell is red, the entered kV is out of range and the correction should not be used.
If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.
Tungsten-Rhodium (W / Rh)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| Radcal kV | True kV
| 25.03 | 21.8081
Tungsten-Silver (W / Ag)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
| Radcal kV | True kV
| 37.1 | 38.9675
27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5