## Barco
1					Revision 1.1-201500513					Print Area					Test frequency
2					Medical University of South Carolina					Medical University of South Carolina					Standard monitor		Tomosynthesis monitor
3					Charleston, South Carolina					Charleston, South Carolina					No QAWS	QAWS	No QAWS	QAWS
4					All:					Autocalibration	Semiannual	Semiannual	Semiannual	Semiannual
5					Mammography Monitor Compliance Inspection					Mammography Monitor Compliance Inspection					Compliance Test	Weekly	Weekly	Weekly	Monthly
6					Measurement Parameter	Last Year		This Year		Display Test			NA	Weekly
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area					Rapid Frame	NA	NA	Quarterly	Quarterly
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10			Facility:					Site Number:					Facility:				Site Number:					Facility:
11			Department:					Date of Installation:					Department:				Date of Installation:					Department:
12			Area/Division:					Workstation ID:					Area/Division:				Workstation ID:					Area/Division:
13			Survey ID:					Left monitor SN:					Survey ID:				Left monitor SN:					Survey ID:
14			Room Number:					Right monitor SN:					Room Number:				Right monitor SN:					Room Number:
15			QC Manual version:					QC Manual version:					QC Manual version:
16					Site Number:
17					Medical University of South Carolina					Inspection Results					Date of Installation:
18					Mammography Monitor Compliance Inspection		Enter 1 for YES, 2 for NO, 3 for NA					Workstation ID:
19					Corrective Actions & Comments					Comments					Left monitor SN:
20			Monitor surface is clean					Monitor surface is clean					Right monitor SN:
21			Viewing conditions					Viewing conditions
22			TG18-QC test pattern					TG18-QC test pattern					Ambient Light:
23			Rapid Frame acceptable					Rapid Frame acceptable					Auto Cal Lum Ratio L:
24			Display Test					Auto Cal Lum Ratio R:
25					Room Conditions					Uniformity Dev UNL10 L:
26		Room Conditions					Ambient Light:		lux					Uniformity Dev UNL10 R:
27		Ambient light:		lux		Criteria:	Ambient light measured at the surface of the monitor(s) must be less than 20 lux					Previous:		lux	Criteria:	Ambient light measured at the surface of the monitor(s) must be less than 20 lux					Uniformity Dev UNL80 L:
28		Comments:					Comments:					Uniformity Dev UNL80 R:
29					Do not exceed 390 chars		0
30					New:					Comments
31					Room Conditions:
32		Auto Calibration					Auto Calibration					Auto Calibration:
33		Date:			Lmin	Lmax	Luminance Ratio					Date:			Left ()	Right ()					Compliance Test:
34				Left ()					Lmin (cd/m^2)					Display Test:
35				Right ()					Lmax (cd/m^2)					TG-18 Black/White:
36		Comments:					Luminance Ratio					TG-18 QC:
37					TG-18 Uniformity:
38					Comments:
39					Do not exceed 390 chars		0					Comments
40		Compliance Test					New:
41		Date:			Left ()	Right ()
42				DICOM GSDF curve acceptable:					Compliance Test
43		Comments:					Date:
44					Enter 1 for YES, 2 for NO, 3 for NA			Left ()	Right ()
45					DICOM GSDF curve acceptable:
46					Comments:
47		TG-18 Test Patterns					Do not exceed 390 chars		0
48		TG-18 LN01					TG-18 LN18					New:
49		Left ()		Right ()				Left ()		Right ()
50		Comments:					Display Test
51					Date:
52					Comments:
53					Do not exceed 390 chars		0
54		TG-18 QC					New:
55					Left ()	Right ()
56				5%/95% patches visible:					TG-18 Test Patterns
57				Resolution acceptable:					TG-18 Black/White		Left ()	Right ()
58				Display is free of artifacts/noise:					LN-01 (cd/m^2)
59		Comments:					LN-18 (cd/m^2)
60					Comments:
61					Do not exceed 390 chars		0
62					New:
63
64					TG-18 QC
65		Date:					Inspector:	Eugene Mah		Enter 1 for YES, 2 for NO, 3 for NA			Left ()	Right ()
66		Room Number					Survey ID:					5%/95% patches visible:
1					Medical University of South Carolina				Resolution acceptable:
2					Mammography Monitor Compliance Inspection				Display is free of artifacts/noise:
3					Comments:
4		TG-18 Uniformity					Do not exceed 390 chars		0
5			UNL-10 (cd/m^2)		UNL-80 (cd/m^2)			Visual Assessment					New:
6			Left ()	Right ()	Left ()	Right ()		TBD	Free from gross non-uniformities
7		Average:					TBD	Free from non-symmetric non-uniformities					TG-18 Uniformity					Visual assessment
8		SD:					TBD	Free from significant luminance variation					UNL-10 (cd/m^2)		UNL-80 (cd/m^2)			Enter 1 for YES, 2 for NO, 3 for NA
9		Deviation (%)					Left ()	Right ()	Left ()	Right ()			Free from gross non-uniformities
10					Criteria:	Luminance deviation is < 30%					Top left					Free from non-symmetric non-uniformities
11		Comments:					Top right					Free from significant luminance variation
12					Center
13					Bottom left
14					Bottom right
15		Comments:					Average (cd/m^2):
16					SD (cd/m^2):
17					Luminance Deviation (%):					Criteria:	Luminance deviation is < 30%
18					Comments:
19					Do not exceed 390 chars		0
20					New:
21
22
23					Comments Page 1
24					Previous Year Comments
25					Comments:
26					Do not exceed 390 chars		0
27					Additional Comments:
28					Do not exceed 390 chars		0
29					Additional Comments:
30					Do not exceed 390 chars		0
31					Additional Comments:
32					Do not exceed 390 chars		0
33					Additional Comments:
34					Do not exceed 390 chars		0
35					Additional Comments:
36					Do not exceed 390 chars		0
37					Additional Comments:
38					Do not exceed 390 chars		0
39					Additional Comments:
40					Do not exceed 390 chars		0
41					Additional Comments:
42					Do not exceed 390 chars		0
43					Additional Comments:
44					Do not exceed 390 chars		0
45					Additional Comments:
46					Do not exceed 390 chars		0
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID: