## Sheet1
1					Revision 1.0-20140806					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					CT System Compliance Inspection					CT System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1
8					System Information					Previous Date:					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Accession Number:					Survey ID:				Accession Number:					Survey ID:
14					Room Number:					Room Number:
15			CT Scanner					Site Number:
16				Manufacturer:					Manufacture Date:					CT Scanner					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18					Model:				Serial Number:					Accesssion Number:
19			X-Ray Tube					Focal Spot Sizes (mm)
20				Tube Designation/Use:					Large:					X-Ray Tube					Focal Spot Sizes (mm)					X-Ray Generator
21				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Manufacturer:
22			Insert					Micro:					Manufacture Date:				Small:					Model:
23				Manufacturer:					Insert					Micro:					Manufacture Date:
24				Model:					Filtration					Manufacturer:					Serial Number:
25				Serial Number:					Inherent:					Model:				Filtration
26			Housing					Added:					Serial Number:				Inherent:					X-Ray Tube 1
27				Manufacturer:					Housing					Added:					Tube Designation/Use:
28				Model:					Manufacturer:					Manufacture Date:
29				Serial Number:					Model:					Manufacturer:
30					Serial Number:					Model:
31					Serial Number:
32		Rule Number			Labels, Notices, Postings					Compliance					Manufacturer:
33		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Inspection Results					Model:
34		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
35				Pregnancy Warning sign is posted					TBD					Large:
36		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Labels, Notices, Postings					Small:
37				Operator manuals are available.					TBD				DHEC Registration sticker is present, clearly visible and legible					Micro:
38				Monthly radiation monitoring reports are posted.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Inherent:
39					Pregnancy Warning sign is posted					Added:
40					Radiation Safety					Radiation warning label posted on the generator control panel					Measurement Protocols
41				Radiation warning light at room entrance is functional with selected tube					TBD				Operator manuals are available.					Adult Abdomen
42				Lead aprons available.					TBD				Monthly radiation monitoring reports are posted.					kVp
43				A properly designed and installed apron rack is present.					TBD					mA
44				Documentation of annual protective apparel integrity inspection is available.					TBD					Radiation Safety					mAs
45				Patient restraint devices available.					TBD				Radiation warning light at room entrance is functional with selected tube					Eff mAs
46					Lead aprons available.					Time/rotation (s)
47				Reconstructed images are free of artifacts					TBD				A properly designed and installed apron rack is present.					Scan FOV (cm)
48				Acquisition station display performance is adequate					TBD				Documentation of annual protective apparel integrity inspection is available.					Display FOV (cm)
49				Daily QC performed by technologist					TBD				Patient restraint devices available.					Reconstruction Algorithm
50				Interlock systems function to prevent scanning while door is open					TBD					Axial/Helical
51		DHEC RHB 4.11.1.1		Tomographic plane is indicated by a light source					TBD				Reconstructed images are free of artifacts					Z-Axis collimation
52		DHEC RHB 4.11.1.3.1		Visual indication of x-ray production is present and functional					TBD				Acquisition station display performance is adequate					# of channels
53		DHEC RHB 4.11.1.6.1		Positioning with the alignment light is accurate to within 2mm					TBD				Daily QC performed by technologist					Table Increment/speed
54		DHEC RHB 4.11.1.6.3		Table travel is accurate to within 1mm					TBD				Interlock systems function to prevent scanning while door is open					Pitch
55		DHEC RHB 4.11.2.5.1		Technologist has a clear view of the patient during the scan					TBD				Tomographic plane is indicated by a light source					Recon scan width (mm)
56				Patient communication system is operational					TBD				Visual indication of x-ray production is present and functional					Recon scan interval (mm)
57					Positioning with the alignment light is accurate to within 2mm					Dose reduction techniques
58					Radiation Dose Measurements					Table travel is accurate to within 1mm					Indiated CTDIvol (mGy)
59				Routine head protocol CTDI is below ACR reference level of 75 mGy					NA				Technologist has a clear view of the patient during the scan					Adult Chest
60				Routine body protocol CTDI is below ACR reference level of 25 mGy					NA				Patient communication system is operational					kVp
61				Pediatric body protocol CTDI is below ACR reference level of 15 mGy					NA					mA
62				Pediatric head protocol CTDI is below ACR reference level of 35 mGy					NA					Radiation Dose Measurements					mAs
63					3	Routine head protocol CTDI is below ACR reference level of 75 mGy					Eff mAs
64					3	Routine body protocol CTDI is below ACR reference level of 25 mGy					Time/rotation (s)
65		Date:					Inspector:	Eugene Mah		3	Pediatric body protocol CTDI is below ACR reference level of 15 mGy					Scan FOV (cm)
66		Room Number					Survey ID:			3	Pediatric head protocol CTDI is below ACR reference level of 35 mGy					Display FOV (cm)
1					Medical University of South Carolina					Reconstruction Algorithm
2					Inspection Results					CT System Compliance Inspection					Axial/Helical
3					Measurement Protocols					Measurement Protocols					Z-Axis collimation
4					Adult			Pediatric					Adult			Pediatric					# of channels
5					Abdomen	Chest	Head	Abdomen	Head					Abdomen	Chest	Head	Abdomen	Head					Table Increment/speed
6				kVp					kVp					Pitch
7				mA					mA					Recon scan width (mm)
8				mAs					mAs					Recon scan interval (mm)
9				Eff mAs					Eff mAs					Dose reduction techniques
10				Time/rotation (s)					Time/rotation (s)					Indiated CTDIvol (mGy)
11				Scan FOV (cm)					Scan FOV (cm)					Adult Head
12				Display FOV (cm)					Display FOV (cm)					kVp
13				Reconstruction Algorithm					Reconstruction Algorithm					mA
14				Axial/Helical					Axial/Helical					mAs
15				Z-Axis collimation					Z-Axis collimation					Eff mAs
16				# of channels					# of channels					Time/rotation (s)
17				Table Increment/speed					Table Increment/speed					Scan FOV (cm)
18				Pitch					Pitch					Display FOV (cm)
19				Recon scan width (mm)					Recon scan width (mm)					Reconstruction Algorithm
20				Recon scan interval (mm)					Recon scan interval (mm)					Axial/Helical
21				Dose reduction techniques					Dose reduction techniques					Z-Axis collimation
22				Indiated CTDIvol (mGy)					Indiated CTDIvol (mGy)					# of channels
23					Adult			Pediatric					Table Increment/speed
24		CTDI Measurements					Abdomen	Chest	Head	Abdomen	Head					Pitch
25					Adult			Pediatric					kVp					Recon scan width (mm)
26					Abdomen	Chest	Head	Abdomen	Head	Criteria:					mA					Recon scan interval (mm)
27				Indicated CTDIvol (mGy):					CTDIvol is less than ACR Pass/Fail dose					mAs					Dose reduction techniques
28			CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					Reference	Pass/Fail			Eff mAs					Indiated CTDIvol (mGy)
29					(mGy)	(mGy)			Time/rotation (s)					Ped Abdomen
30					Adult Body:	25	30			Scan FOV (cm)					kVp
31				Average CTDIvol:					Adult Head:	75	80			Display FOV (cm)					mA
32				Phantom size indicated (cm)					Pediatric Body:	15	20			Reconstruction Algorithm					mAs
33				Phantom size used (cm)					Pediatric Head:	35	40			Axial/Helical					Eff mAs
34				Acceptable:					Z-Axis collimation					Time/rotation (s)
35					# of channels					Scan FOV (cm)
36				CTDIvol/mAs:					Table Increment/speed					Display FOV (cm)
37					Pitch					Reconstruction Algorithm
38		Beam Profile					Recon scan width (mm)					Axial/Helical
39					Recon scan interval (mm)					Z-Axis collimation
40					Dose reduction techniques					# of channels
41				Channels (N)					Indiated CTDIvol (mGy)					Table Increment/speed
42				Detector width (mm)					Pitch
43				Table velocity (mm/s)					Recon scan width (mm)
44				Beam width (mm)					Inspection Results					Recon scan interval (mm)
45				Measured beam width (mm)					CTDI Measurements					Dose reduction techniques
46					Adult			Pediatric			Criteria:				Indiated CTDIvol (mGy)
47					Abdomen	Chest	Head	Abdomen	Head		CTDIvol is less than ACR Pass/Fail dose				Ped Head
48					Indicated CTDIvol (mGy)					Reference	Pass/Fail		kVp
49				Acceptable:					CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					(mGy)	(mGy)		mA
50					Criteria:	Beam profile is within 3mm or 30% of specified value					Adult Body:	25	30		mAs
51					Adult Head:	75	80		Eff mAs
52					Average CTDIvol:					Pediatric Body:	15	20		Time/rotation (s)
53					Phantom size indicated (cm)					Pediatric Head:	35	40		Scan FOV (cm)
54					Phantom size used (cm)					Display FOV (cm)
55					Previous year:					Reconstruction Algorithm
56					Acceptable:					Axial/Helical
57					Z-Axis collimation
58					CTDIvol/100 mAs:					# of channels
59					Prev CTDIvol/100 mAs:					Table Increment/speed
60					Pitch
61					Beam Profile					Recon scan width (mm)
62					Channels (N)					Recon scan interval (mm)
63					Detector width (mm)					Criteria:	Beam profile is within 3mm or 30% of specified value					Dose reduction techniques
64					Table velocity (mm/s)					Indiated CTDIvol (mGy)
65		Date:					Inspector:	Eugene Mah			Beam width (mm)					CTDI Measurements
66		Room Number					Survey ID:				Measured beam width (mm)					Adult Abdomen
1					Medical University of South Carolina					Adult Chest
2					Inspection Results					CT System Compliance Inspection					Adult Head
3		Alignment Accuracy					Average (mm):					Ped Abdomen
4			Landmark position:		mm		Slice thickness:		mm					Acceptable:					Ped Head
5					CTDIvol/100 mAs
6					Foot					Head					Adult Abdomen
7				Location of alignment slice:		mm			Location of alignment slice:		mm					Inspection Results					Adult Chest
8				All markers visible					All markers visible					Alignment Accuracy					Adult Head
9				Central wire visible					Table movement accuracy					Landmark position:		mm	Slice thickness		mm					Ped Abdomen
10					Ped Head
11		CT Number Accuracy					Foot				Head
12			kVp	80	100	110	120	130	140	150					Location of alignment slice:		mm		Location of alignment slice:		mm
13			mAs					All markers visible				All markers visible
14			Polyethylene					Central wire visible				Table movement accuracy					CT Beam Profile – Siemens
15			Water					Enter 1 for YES, 2 for NO				Enter 1 for YES, 2 for NO					NxT	Definition Flash			NxT	Definition AS			NxT	Sensation 64		NxT	Sensation 16		NxT	Emotion 16
16			Water Std Dev					CT Number Accuracy					Z-sharp	Nominal (mm)	Tolerance)		Z-sharp	Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)
17			Acrylic					kVp	80	100	110	120	130	140	150		Acceptable HU			2x1	Off	3.5	1.0	2x1	Off	4.7	1.0	30x0.6	22.5	1.5	16x1.5	30.5	1.0	4x0.6	4.9	1.0
18			Bone					mAs					Low	High		1x5	Off	6.0	1.0	1x5	Off	6.6	1.0	24x1.2	33.5	1.5	16x0.75	18.5	1.0	12x0.6	12.6	1.0
19			Air					Polyethylene					Polyethylene	-107	-87		10x0.6	On	10.5	1.5	10x0.6	On	11.2	1.0	12x1.2	19.5	1.5	12x1.5	24.5	1.0	16x0.6	14.5	1.0
20					Acceptable:					Water					Water	-7	7		8x0.6	On	11.5	1.5	8x0.6	Off	8.5	1.0	1x10	11.5	1.5	12x0.75	15.5	1.0	2x5	12.6	1.0
21					Critiera:	Water HU is 0 +/- 5 at all kVp					Water Std Dev					Acrylic	110	130		1x10	Off	11.1	1.5	1x10	Off	11.6	1.0	1x5	6.5	1.5	2x5	12.5	1.0	12x1.2	18.5	1.0
22					CT number for inserts using the adult abdomen protocol is within the acceptable range					Acrylic					Bone	850	970		20x0.6	On	18.8	1.5	8x0.6	On	11.5	1.0	2x1	4.0	1.0	2x1	3.5	1.0	2x8	18.5	1.0
23					Bone					Air	-1005	-970		12x1.2	Off	19.4	1.5	6x1.2	Off	11.8	1.0	6x0.6	7.0	1.0	2x0.6	3.0	1.0	16x1.2	24.2	1.0
24		CT Slice Thickness Accuracy					Air					32x0.6	Off	24.4	2.0	12x1.2	Off	19.8	1.0
25				Slice thickness (mm):					Acceptable:		Enter 1 for YES, 2 for NO					32x0.6	On	27.4	2.0	60x0.6	Off	41.0	1.5
26				Top (mm):					Critiera:	Water HU is 0 +/- 5 at all kVp					32x1.2	Off	44.0	2.0	64x0.6	On	45.5	1.5
27				Bottom (mm):					CT number for inserts using the adult abdomen protocol is within the acceptable range					64x0.6	On	45.5	2.0	32x1.2	Off	43.5	1.5
28				Average (mm):					30x0.6	Off	23.5	1.0
29				Water HU:					CT Slice Thickness Accuracy					32x0.6	On	28.0	1.0
30				Water HU Std Dev:					Slice thickness (mm):					20x0.6	Off	18.0	1.0
31				Acceptable:					Top (mm):					16x1.2	Off	25.0	1.0
32					Criteria:	Slice thickness is within 1.5mm of the set thickness					Bottom (mm):					20x0.6	On	20.0	1.0
33					CT number for water is 0 +/- 5					Average (mm):
34					Water HU:
35		Low Contrast Performance					Water HU Std Dev:
36				Adult		Pediatric					Acceptable:
37				Abdomen	Head	Abdomen	Head					Criteria:	Slice thickness is within 1.5mm of the set thickness
38			Inside					6mm rods visible:					CT number for water is 0 +/- 5
39			Adjacent					Acceptable:
40			Adjacent StDev					Low Contrast Performance
41			Contrast					Adult		Pediatric
42			CNR					Abdomen	Head	Abdomen	Head			Enter 1 for YES, 2 for NO, 3 for Not Applicable
43					Criteria:	All four 6mm rods are visible					Inside					6mm rods visible:		Acceptable:
44					CNR > 1.0 for adult abdomen and head protocols					Adjacent					Criteria:	All four 6mm rods are visible
45					CNR > 0.4 for pediatric body protocol					Adjacent StDev					CNR > 1.0 for adult abdomen and head protocols
46					CNR > 0.7 for pediatric head protocol					Contrast					CNR > 0.4 for pediatric body protocol
47					CNR					CNR > 0.7 for pediatric head protocol
48		Uniformity
49				Mean	St Dev				Acceptable					Uniformity
50			Center:				Center-Top:					Mean	St Dev				Acceptable
51			Top:				Center-Bottom:					Center:				Center-Top:
52			Bottom:				Center-Right:					Top:				Center-Bottom:
53			Right:				Center-Left:					Bottom:				Center-Right:
54			Left:					Right:				Center-Left:
55					Criteria:	CT number of all ROIs is within +/- 5 HU of the central region					Left:				Criteria:	CT number of all ROIs is within +/- 5 HU of the central region
56
57		High Contrast Resolution					High Contrast Resolution
58				lp/cm	Acceptable					lp/cm	Acceptable
59			Abdomen				Critiera:	5 lp/cm for abdomen					Abdomen				Critiera:	5 lp/cm for abdomen
60			Hi-res Chest					6 lp/cm for high resolution					Hi-res Chest					6 lp/cm for high resolution
61
62		Acquisition Console Monitor					Acquisition Console Monitor
63			SMPTE 5%/95% visible			Display free of artifacts			Display uniformity
64			Lmin (cd/m^2)			Lmax (cd/m^2)					SMPTE 5%/95% visible			Lmin (cd/m^2)
65		Date:					Inspector:	Eugene Mah			Display free of artifacts			Lmax (cd/m^2)
66		Room Number					Survey ID:				Display uniformity