## Gen_form
1 | [10 empty] | Revision 0.1-20200916 (Has 42 data records) | [10 empty] | PRINT AREAS
2 |  |  |  |  | Medical University of South Carolina | [9 empty] | Medical University of South Carolina | [6 empty] | FIRST,COMPG1,COMPG2,EXP_BLD,FS_KV,TIM,REPRO,LIN,HVL_IMGQUAL,OUTPUT,CTPG1,CTPG2
3 |  |  |  |  | Charleston, South Carolina | [9 empty] | Charleston, South Carolina | [6 empty] | ALL: | FIRST,COMPG1,COMPG2,EXP_BLD,FS_KV,TIM,REPRO,LIN,HVL_IMGQUAL,OUTPUT,CTPG1,CTPG2
4 | [23 empty] | REM: | (1) Go to last year's spreadsheet (see "**Note", below) and copy from ITS "THIS YEAR" range
5 |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | [8 empty] | Radiation Oncology IGRT Compliance Inspection | [10 empty] | (2) Come back to this sheet and paste (use paste special, "values") into "LAST YEAR" range
6 | [24 empty] | LAST | Cell AD6 | THIS
7 | [23 empty] | Measurement Parameter | YEAR |  | YEAR | (Locked) | RANGES: | LASTYEAR (Data, starting with cell AB8)
8 |  |  |  |  | System Information | [6 empty] | Date: |  |  |  | System Information |  |  | Inspector: | Eugene Mah |  |  |  | Print Area: |  | Change | FIRST,COMPG1,COMPG2,EXP_BLD,FS_KV,TIM,REPRO,LIN,HVL_IMGQUAL,OUTPUT,CTPG1,CTPG2 | Y2 |  | THISYEAR (Data, starting with cell AD8)
9 | [11 empty] | Last Year: | [6 empty] | Last Year: |  |  |  |  | Date: |  |  |  | O8
10 |  | Location | [9 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  |  | Inspector: |  | Change | Eugene Mah | V8
11 |  |  | Facility: |  |  |  | Site: | [5 empty] | Facility: |  |  |  | Site Number: | [6 empty] | Facility: |  |  |  | P11
12 |  |  | Department: |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [6 empty] | Department: |  |  |  | P12
13 |  |  | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [6 empty] | Area/Division: |  |  |  | P13
14 |  |  | Survey ID: |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: | [6 empty] | Survey ID: |  |  |  | P14
15 | [12 empty] | Room Number: |  |  |  | Accession Number: | [6 empty] | Room Number: |  |  |  | P15
16
17 |  | X-Ray Generator | [9 empty] | X-Ray Generator | [11 empty] | X-Ray Generator
18 |  |  | Manufacturer: |  |  |  | Date of Manufacture: | [5 empty] | Manufacturer: |  |  |  | Date of Manufacture: | [6 empty] | Manufacturer: |  |  |  | P18
19 |  |  | Model: |  |  |  | Serial Number: | [5 empty] | Model: |  |  |  | Serial Number: | [6 empty] | Model: |  |  |  | P19
20 |  |  | Max. kVp: |  |  |  | Max. mA: | [5 empty] | Max. kVp: |  |  |  | Max. mA: | [6 empty] | Max. kVp: |  |  |  | P20
21 |  |  | Software Version: | [9 empty] | Software Version: | [10 empty] | Software Version: |  |  |  | P21
22 | [23 empty] | Date of Manufacture: |  |  |  | T18
23 |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes |  |  |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes (mm) | [5 empty] | Serial Number: |  |  |  | T19
24 |  |  | Tube Designation/Use: |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Large: | [6 empty] | Max. mA: |  |  |  | T20
25 |  |  | Date of Manufacture: |  |  |  | Small: | [5 empty] | Date of Manufacture: |  |  |  | Small:
26 | [16 empty] | Micro: | [6 empty] | X-Ray Tube 1
27 |  |  | Insert Manufacturer: |  |  |  |  | Filtration |  |  |  |  | Insert Manufacturer: |  |  |  |  | Filtration | [5 empty] | Tube Designation/Use: |  |  |  | P24
28 |  |  | Insert Model: |  |  |  | Inherent: | [5 empty] | Insert Model: |  |  |  | Inherent: | [6 empty] | Date of Manufacture: |  |  |  | P25
29 |  |  | Insert Serial Number: |  |  |  | Added: | [5 empty] | Insert Serial Number: |  |  |  | Added:
30 | [23 empty] | Insert Manufacturer: |  |  |  | P27
31 |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor |  |  |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor | [5 empty] | Insert Model: |  |  |  | P28
32 |  |  | Housing Model: |  |  |  | Film: | [5 empty] | Housing Model: |  |  |  | Film: | [6 empty] | Insert Serial Number: |  |  |  | P29
33 |  |  | Housing Serial Number: |  |  |  | Screen: | [5 empty] | Housing Serial Number: |  |  |  | Screen:
34 | [23 empty] | Housing Manufacturer: |  |  |  | P31
35 | [23 empty] | Housing Model: |  |  |  | P32
36 | [23 empty] | Housing Serial Number: |  |  |  | P33
37
38 | [23 empty] | Site Number: |  |  |  | T11
39 | [23 empty] | Authorized Use: |  |  |  | T12
40 | [23 empty] | Date of Installation: |  |  |  | T13
41 | [23 empty] | Number of X-ray Tubes: |  |  |  | T14
42 | [23 empty] | Accession Number: |  |  |  | T15
43
44 | [23 empty] | Focal Spot Sizes
45 | [23 empty] | Large: |  |  |  | T24
46 | [23 empty] | Small: |  |  |  | T25
47 | [23 empty] | Micro: |  |  |  | T26
48 | [23 empty] | Filtration
49 |  |  |  |  | Inspection Results | [10 empty] | Inspection Results | [7 empty] | Inherent: |  |  |  | T28
50 | [23 empty] | Added: |  |  |  | T29
51 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA
52 | Rule Number |  |  |  | Labels, Notices, Postings |  |  |  | Compliance |  | 2 | 2 | Is this a new installation? | [10 empty] | Image Receptor
53 | [23 empty] | Film: |  |  |  | T32
54 | DHEC RHB 2.5.1.1 |  | DHEC registration sticker is present, clearly visible and legible. | [5 empty] | TBD | [5 empty] | Labels, Notices, Postings |  |  | 1 | 1 | Tube Number** |  |  |  | Screen: |  |  |  | T33
55 | DHEC RHB 10.2.1 |  | DHEC form SC-RHA-20 "Notice to Employees" posted or referenced. | [5 empty] | TBD | [10 empty] | for this room or
56 | DHEC RHB 4.2.7 |  | Pregnancy Warning sign is posted | [5 empty] | NA |  |  |  | DHEC registration sticker is present, clearly visible and legible. | [6 empty] | generator |  |  |  | New install? |  | Change | 2 | M52
57 | DHEC RHB 4.3.1 |  | Radiation warning label present on the generator control panel. | [5 empty] | TBD |  |  |  | DHEC form SC-RHA-20 "Notice to Employees" posted or referenced. | [5 empty] | **Note: Input only one of the |  |  |  |  | Tube Designation |  | Change | 1 | U54
58 | DHEC RHB 4.2.6 |  | Technique chart is present and accurate. |  |  |  | DATE: |  | NA |  | 3 |  | Pregnancy Warning sign is posted | [5 empty] | following numerals or names:
59 | JCAHO |  | Patient exposure/dose measurements are posted. |  |  |  | DATE: |  | NA |  |  |  | Radiation warning label present on the generator control panel. | [5 empty] | 1 | 2 |  |  |  | Table Bucky
60 |  |  | Tube loading charts for each focal spot are posted. | [5 empty] | NA |  | 3 | DATE: |  | Technique chart present. |  |  |  |  | AP | 3 |  |  |  | TB used? |  | Change | 1 | M409
61 |  |  | Tube loading charts are correct for installed tube(s). | [5 empty] | NA |  | 3 | DATE: |  | Patient exposure/dose measurements posted. |  |  |  |  | Front | Lat |  |  |  | Beam alignment on table top? |  | Change | 2 | M410
62 |  |  | Operators manuals are available. | [5 empty] | TBD |  | 3 |  | Tube loading charts for each focal spot are posted. | [5 empty] | Frontal | Lateral
63 |  |  | Monthly radiation monitoring reports are posted. | [5 empty] | TBD |  | 3 |  | Tube loading charts are correct for installed tube(s). | [5 empty] | Ceph | Pan |  |  |  | Source to Table Top (cm) |  | Change | 100 | M420
64 | [12 empty] | Operators manuals are available. | [5 empty] | Mobile | BCM |  |  |  | Table top to cassette (cm) |  | Change | 50 | M421
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  |  | Monthly radiation monitoring reports are posted.
66 | Room Number: | [6 empty] | Survey ID: | [15 empty] | Table bucky cm/in |  | Change | cm | V423
1 | [9 empty] | Medical University of South Carolina | [8 empty] | NEW VALUES |  |  |  |  | Bucky SID |  | Change | 150 | U423
2 |  |  |  |  | Inspection Results |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | [23 empty] | Indicated field size
4 | [23 empty] | Transverse 1 |  | Change | 5 | M192
5 |  |  |  |  | Generator |  |  |  | Compliance | [14 empty] | Longitudinal 1 |  | Change | 5 | N192
6 | DHEC RHB 4.3.9.1 |  | Means are provided to indicate technique factors prior to exposure. | [5 empty] | TBD | [14 empty] | Transverse 2 |  |  |  | M193
7 |  |  | All technique indicators functional and are free of excessive wear. | [5 empty] | TBD | [14 empty] | Longitudinal 2 |  |  |  | N193
8 |  |  | Technique overload circuits function to block excessive technique. | [5 empty] | TBD
9 |  |  | Anode heat calculator/indicator is functioning properly. | [5 empty] | NA | [14 empty] | X-RAY TUBE
10 | DHEC RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational. | [5 empty] | TBD | [6 empty] | COMMENTS | [6 empty] | Large Focus: | mA/mAs |  | Change | mAs | P231
11 | DHEC RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational. | [5 empty] | TBD |  |  | Note: Type all comments in the shaded cells. | [11 empty] | kVp |  |  |  | O233
12 | DHEC RHB 4.7.4.2.4 |  | Audible exposure end indicator present and operational. | [5 empty] | TBD | [5 empty] | Previous Year Comments: | [8 empty] | mA |  |  |  | P233
13 |  |  | Exposure delay time is not excessive. | [5 empty] | TBD |  |  | Comments: | [11 empty] | Time |  |  |  | Q233
14 | DHEC RHB 4.7.7 |  | kVp accurate to within 5% of indicated. | [5 empty] | TBD |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag. | [6 empty] | SID |  |  |  | R233
15 | DHEC RHB 4.7.5 |  | kVp and exposure reproducibility coeff. of var. < 5%. | [5 empty] | NA |  |  | Additional comments: | [11 empty] | OFD |  |  |  | S233
16 | DHEC RHB 4.7.7 |  | Timer accuracy within 10%. | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **A technique chart was not posted in the room. A technique chart reflecting techniques currently used with the CR system must be provided (DHEC RHB 4.2.8).
17 | DHEC 4.7.4.2.7 |  | Timer reproducibility coeff. of var. < 5%. | [5 empty] | NA |  |  | Additional comments: | [11 empty] | Last year Lrg FS resolution |  |  |  | U233
18 | DHEC RHB 4.7.8 |  | Exposure coefficient of linearity < 10%. | [5 empty] | TBD |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console. | [6 empty] | Previous 1 yr Lrg FS resolution |  |  |  | V233
19 | [11 empty] | Additional comments:
20 |  |  |  |  | X-ray Tube | [6 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Tube loading charts for each focal spot were not available at the time of inspection. These charts should be obtained and posted near the operator's console. | [5 empty] | Small Focus: | mA/mAs |  | Change | mAs | P236
21 |  |  | Focal spot resolution is within acceptable parameters | [5 empty] | TBD |  |  | Additional comments: | [11 empty] | kVp |  |  |  | O238
22 |  |  | The x-ray output @40" within +/-25% from previous measurements. | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **Operator's manuals for the imaging equipment were not available at the time of inspection. These should be obtained and made available to the operator of the equipment. | [6 empty] | mA |  |  |  | P238
23 | [11 empty] | Additional comments: | [11 empty] | Time |  |  |  | Q238
24 |  |  |  |  | Radiation Safety | [6 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office. | [6 empty] | SID |  |  |  | R238
25 | DHEC RHB 4.7.4.2.1 |  | Continuous pressure on exposure switch required to make an exposure. | [5 empty] | TBD |  |  | Additional comments: | [11 empty] | OFD |  |  |  | S238
26 |  |  | Spot film exposure blocked when there is no cassette | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **The light field illumination failed to meet the minimum specification of 160 lux. The light source should be adjusted or replaced so as to meet this requirement.
27 |  |  | Spot film exposure blocked when serial exposures are completed | [5 empty] | NA |  |  | Additional comments: | [11 empty] | Last year Sm FS resolution |  |  | 0 | U238
28 | 21CFR 1020.31(g)(4) |  | Spot film device provides acceptable centering to the image receptor | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Previous 1 yr Sm FS resolution |  |  |  | V238
29 | DHEC RHB 4.7.4.2.2 |  | Exposure switch for selected tube is secured at a safe location | [5 empty] | TBD |  |  | Additional comments:
30 | DHEC RHB 4.7.4.2.5 |  | Operator has full visibility of the exposure factors and patient. | [5 empty] | TBD |  |  | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Test Equipment Utilized: |  | Change | Piranha CB2-17090320 | O251
31 | DHEC RHB 4.2.14.7 |  | Lead aprons available. |  | Half: 0 |  | Full: 0 |  | NA | [14 empty] | Piranha CB2-17090320
32 | DHEC RHB 4.2.14.7 |  | Lead gloves available. |  |  | Number (Left + Right): | 0 |  | NA
33 |  |  | A properly designed and installed apron rack is present. | [5 empty] | NA | [14 empty] | LFSDD |  | Change | cm | N257
34 | DHEC RHB 4.2.14.7 |  | Documentation of annual protective apparel integrity inspection is available. | [5 empty] | NA | [14 empty] | SFSDD |  | Change | cm | N275
35 |  |  | Patient restraint devices available. | [5 empty] | TBD | [5 empty] | DO NOT USE THIS SPACE
36 | DHEC RHB 4.2.12 |  | Gonadal shielding is present and functional. | [5 empty] | NA | [14 empty] | Timer
37 |  |  | Radiation warning light at room entrance is functional with selected tube | [5 empty] | NA | [14 empty] | Generator Type |  | Change | HF | O369
38 | [23 empty] | Timer Test Valid? |  | Change | 2 | M370
39 |  |  |  |  | Physical Inspection
40 |  |  | System clean and free of debris. | [5 empty] | YES
41 |  |  | All cables free from excessive wear or restraints. | [5 empty] | YES
42 |  |  | System free of missing or damaged components. | [5 empty] | YES | [14 empty] | Large Focus Output
43 |  |  | Table movements and locks acceptable | [5 empty] | YES | [14 empty] | SDD(in.) |  | Change | 39.3701 | AB188
44 |  |  | Tube crane movements and locks acceptable. | [5 empty] | YES
45 |  |  | All other movements and locks functional. | [5 empty] | YES | [6 empty] | COMMENTS PAGE 2 | [7 empty] | kVp |  |  |  | P1397
46 | [27 empty] | P1398
47 | [14 empty] | Previous Year Comments: | [12 empty] | P1399
48 | [11 empty] | Comments: | OPERATIONAL NOTES | [14 empty] | P1400
49 | [11 empty] | Do NOT exceed 390 |  | 17 |  | 0 | [11 empty] | P1401
50 | [11 empty] | Additional comments: | Isocenter is 100 cm from focal spot | [14 empty] | P1402
51 | [11 empty] | Do NOT exceed 390 |  | 35 |  | 0 | [11 empty] | P1403
52 | [11 empty] | Additional comments: | Collimator blade locations and field sizes are referenced to isocenter. | [14 empty] | P1404
53 | [11 empty] | Do NOT exceed 390 |  | 71 |  | 0
54 | [11 empty] | Additional comments: | Use 5x5cm field size for exposure/kV measurements. | [10 empty] | mGy/mAs |  |  |  | R1397
55 | [11 empty] | Do NOT exceed 390 |  | 50 |  | 0 | [7 empty] | @ 39.4 in. |  |  |  | R1398
56 | [11 empty] | Additional comments: | For generator check, align kV/exposure detectors at isocenter. | [14 empty] | R1399
57 | [11 empty] | Do NOT exceed 390 |  | 62 |  | 0 | [11 empty] | R1400
58 | [11 empty] | Additional comments: | [15 empty] | R1401
59 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1402
60 | [11 empty] | Additional comments: | [15 empty] | R1403 |  |  | 80 kV
61 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1404 |  |  | mm Al. | Avg. Exp.
62 | [11 empty] | Additional comments: | [18 empty] | #DIV/0! | #DIV/0!
63 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Small Focus Output | [6 empty] | #DIV/0! | #DIV/0!
64 | [11 empty] | Additional comments: | [11 empty] | SDD(in.) |  | Change | 39.3701 | N1411 |  |  | #DIV/0! | #DIV/0!
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Do NOT exceed 390 |  | 0 |  | 0 | [14 empty] | #DIV/0! | #DIV/0!
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | Additional comments: | [11 empty] | kVp |  |  |  | P1414 |  |  | #DIV/0! | #DIV/0!
1 | [9 empty] | Medical University of South Carolina |  | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1415 |  |  | #DIV/0! | #DIV/0!
2 |  |  |  |  | COMMENTS PAGE 1 |  |  |  |  | Radiation Oncology IGRT Compliance Inspection |  | Additional comments: | [15 empty] | P1416
3 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1417
4 | [11 empty] | Additional comments: | [15 empty] | P1418
5 | Comments: | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1419
6 | [11 empty] | Additional comments:
7 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | mGy/mAs |  |  |  | R1414
8 | [11 empty] | Additional comments: | [11 empty] | @ 39.4 in. |  |  |  | R1415
9 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1416
10 | [11 empty] | Additional comments: | [15 empty] | R1417
11 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1418
12 | [11 empty] | Additional comments: | [15 empty] | R1419
13 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | mGy/mAs Output
14 | [11 empty] | Additional comments: | [11 empty] | Lrg Focus Max. |  | Change | NA | Q1406
15 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Lrg Focus Median |  | Change | NA | R1406
16 | [11 empty] | Additional comments: | [11 empty] | Lrg Focus Min |  | Change | NA | S1406
17 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Sml Focus Max. |  | Change | NA | Q1421
18 | [11 empty] | Additional comments: | [11 empty] | Sml Focus Median |  | Change | NA | R1421
19 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Sml Focus Min |  | Change | NA | S1421
20
21 | [23 empty] | COMMENT PAGE 1 |  |  | 0 | O170
22
23 | [26 empty] | 0 | O172
24 | [14 empty] | DO NOT USE THIS SPACE
25 | [26 empty] | 0 | O174
26
27 | [26 empty] | 0 | O176
28
29 | [26 empty] | 0 | O178
30
31 | [26 empty] | 0 | O180
32
33 | [26 empty] | 0 | O182
34
35 | [26 empty] | 0 | O184
36
37 | [26 empty] | 0 | O186
38
39 | [23 empty] | COMMENT PAGE 2 |  | Change | OPERATIONAL NOTES | O205
40
41 | [15 empty] | IMAGE RECEPTOR | [9 empty] | Change | Isocenter is 100 cm from focal spot | O207
42 | [10 empty] | Enter 1 for YES, 2 for NO
43 | [10 empty] | 1 | 1 | Is this tube used with a table bucky or other horizontal type film image receptor? | [12 empty] | Change | Collimator blade locations and field sizes are referenced to isocenter. | O209
44 | [10 empty] | 2 | 2 | Mag. Correction - Is beam alignment test being done with the film on the table top?
45 | [25 empty] | Change | Use 5x5cm field size for exposure/kV measurements. | O211
46 | [10 empty] | Beam Alignment |  |  |  |  | NOTE: | Use consistent units (cm/in.)
47 | [10 empty] | 100 | 100 | (cm) Source to Table |  |  |  | throughout |  | Last Year: | ?? | [5 empty] | Change | For generator check, align kV/exposure detectors at isocenter. | O213
48 | [10 empty] | 50 | 50 | (cm) Table top to receptor
49 | [18 empty] | Push button to select cm/in. | [7 empty] | 0 | O215
50 | [17 empty] | Bucky SID: | 150 | cm
51 | [17 empty] | Last Year | [8 empty] | 0 | O217
52 | [17 empty] | New
53 | [26 empty] | 0 | O219
54 | [12 empty] | Indicated Field Size @SID |  | Measured* Radiation |  | Field Size Indicators |  | Acceptable
55 | [10 empty] | New Values |  | (cm) |  | (cm) |  | Error (% SID) |  | Field Size Indicators | [7 empty] | 0 | O221
56 | [10 empty] | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal
57 | [10 empty] | 5 | 5 | 5 | 5 | 5 | 5 | 0% | 0% | YES | YES | [6 empty] | 0 | O223
58 | [16 empty] | NA | NA | NA | NA
59 | [10 empty] | *Dimension of highest optical density region | [15 empty] | 0 | O225
60 | [10 empty] | **Dimension between the image of the coins marking the edge of the light field
61 | [26 empty] | 0 | O227
62 | [11 empty] | Criteria: | The indicated field length/width are accurate to within 2% of the SID.
63 | [12 empty] | X-ray field and light field length/width are congruent within 2% SID. | [13 empty] | 0 | O229
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [17 empty] | 0 | O231
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [16 empty] | 0 | O233
2 |  |  |  |  | Inspection Results |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | [26 empty] | 0 | O235
4 |  |  |  |  | COMMENTS PAGE 2
5 | [26 empty] | 0 | O237
6
7 | Comments: | OPERATIONAL NOTES | [12 empty] | X-ray Tube |  |  |  |  | 1 Year | 2 Year | [5 empty] | 0 | O239
8 | [10 empty] | Focal Spot Resolution | [7 empty] | Measured | Previous | Previous
9 | [10 empty] | Large: |  | mm | mAs |  |  |  |  | Resolution | Resolution | Resolution | [5 empty] | 0 | O241
10 |  | Isocenter is 100 cm from focal spot | [10 empty] | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm)
11 | [23 empty] | COMMENTS
12 | [11 empty] | New Values: | [11 empty] | Focal Spot Resolution |  |  |  | O683
13 |  | Collimator blade locations and field sizes are referenced to isocenter. | [16 empty] | Measured | Previous | Previous |  |  | Large Focus kV accuracy |  |  |  | O703
14 | [10 empty] | Small: |  | mm | mAs |  |  |  |  | Resolution | Resolution | Resolution |  |  | Small Focus kV accuracy |  |  |  | O719
15 | [12 empty] | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  |  | Large Focus Timer Reprod-1 |  |  |  | O751
16 |  | Use 5x5cm field size for exposure/kV measurements. | [21 empty] | Large Focus Timer Reprod-2 |  |  |  | O753
17 | [11 empty] | New Values: | [11 empty] | Small Focus Timer Reprod-1 |  |  |  | P770
18 | [12 empty] | * SID = Source to Image (film) Distance. | [10 empty] | Small Focus Timer Reprod-2 |  |  |  | P772
19 |  | For generator check, align kV/exposure detectors at isocenter. | [10 empty] | ** OFD = Object (resolution pattern) to Film Distance. | [10 empty] | Exposure Timer |  |  |  | O824
20 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [12 empty] | Large Focus Linearity |  |  |  | O877
21 | [11 empty] | Focal spot resolution is within acceptable parameters | [11 empty] | Small Focus Linearity |  |  |  | O892
22 | [12 empty] | Criteria: | No significant change in resolution since previous measurement. | [9 empty] | Half Value Layer |  |  |  | O946
23
24 | [23 empty] | Tube Designation(column 1): |  | Change | 1 | U59
25 | [11 empty] | Comments: | [13 empty] | Change | AP | U60
26 | [11 empty] | Do NOT exceed 390 |  | 0 | [11 empty] | Change | Front | U61
27 | [11 empty] | New | [13 empty] | Change | Frontal | U62
28 | [15 empty] | Generator Performance | [9 empty] | Change | Ceph | U63
29 | [11 empty] | Test Equipment Used: | Piranha CB2-17090320 |  |  | Calibration Date: | [9 empty] | Change | Mobile | U64
30 | [11 empty] | New: | Piranha CB2-17090320 |  |  | Calibration Due: | [11 empty] | U65
31 | [10 empty] | Large Focus kV Accuracy and Output
32 | [23 empty] | Tube Designation(column 2): |  | Change | 2 | V59
33 | [14 empty] | Note: See focal spot data above for mA/mAs selector button | [10 empty] | Change | 3 | V60
34 | [10 empty] | Test | SDD* | Focal | Filter Pack/ |  |  | Set |  | Measured values | [6 empty] | Change | Lat | V61
35 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  |  |  |  | Change | Lateral | V62
36 | [11 empty] | 100 | Large | 50-85/Low | 60 | 20 | 0.05 | [8 empty] | Change | Pan | V63
37 | [11 empty] | 100 | Large | 70-120/Low | 80 | 20 | 0.05 | [8 empty] | Change | BCM | V64
38 | [11 empty] | 100 | Large | 70-120/Low | 100 | 20 | 0.05 | [10 empty] | V65
39 | [11 empty] | 100 | Large | 100-155/Low | 120 | 20 | 0.05
40 | [11 empty] | 100 | Large | 100-155/Low | 140 | 20 | 0.05 | [6 empty] | HVL Measured kVp: |  | Change | HVL @80 kVp (Max kV) | S950
41 | [23 empty] | HVL Measured: |  | #DIV/0! | #DIV/0! | S951
42
43 | [23 empty] | No. Half Aprons: |  |  | 0 | R1430
44 | [10 empty] | *SDD=Source to detector distance | [12 empty] | No. Full Aprons: |  |  | 0 | T1430
45 | [11 empty] | Comments: | [11 empty] | No. Lead Gloves: |  |  | 0 | R1431
46 | [11 empty] | Do NOT exceed 190 |  | 0 |  |  |  |  | Acceptable: | TBD
47 | [11 empty] | New | [11 empty] | Contact/access Info.:
48 | [10 empty] | Small Focus kV Accuracy and Output | [16 empty] | T1
49 | [27 empty] | T2
50 | [27 empty] | T3
51 | [14 empty] | Note: See focal spot data above for mA/mAs selector button | [12 empty] | T4
52 | [10 empty] | Test | SDD* | Focal | Filter Pack/ |  |  | Set |  | Measured values | [8 empty] | T5
53 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | [6 empty] | T6
54 | [11 empty] | 100 | Small | 35-60/Low | 50 | 10 | 0.1
55 | [11 empty] | 100 | Small | 50-85/Low | 70 | 10 | 0.1 | [6 empty] | Leeds Phantom
56 | [11 empty] | 100 | Small | 70-120/Low | 90 | 10 | 0.1 | [6 empty] | kVp: |  | Change | 70 | O543
57 | [11 empty] | 100 | Small | 70-120/Low | 110 | 10 | 0.1 | [6 empty] | mAs: |  |  |  | P543
58 | [11 empty] | 100 | Small | 100-155/Low | 130 | 10 | 0.1 | [6 empty] | TO.10.A |  |  |  | O546
59 | [23 empty] | TO.10.B |  |  |  | O547
60 | [10 empty] | *SDD=Source to detector distance | [12 empty] | TO.10.C |  |  |  | O548
61 | [11 empty] | Comments: | [11 empty] | TO.10.D |  |  |  | O549
62 | [11 empty] | New |  | 0 | [9 empty] | TO.10.E |  |  |  | O550
63 | [13 empty] | (Combined comments) |  |  |  |  | Acceptable: | TBD |  |  |  | TO.10.F |  |  |  | O551
64 | [11 empty] | Do NOT exceed 190 |  | 0 | [9 empty] | TO.10.G |  |  |  | O552
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Criteria: | kVp accurate to within 5% of indicated. | [10 empty] | TO.10.H |  |  |  | O553
66 | Room Number: | [6 empty] | Survey ID: | [15 empty] | TO.10.J |  |  |  | O554
1 | [9 empty] | Medical University of South Carolina | [13 empty] | TO.10.K |  |  |  | O555
2 | [9 empty] | Radiation Oncology IGRT Compliance Inspection | [13 empty] | TO.10.L |  |  |  | O556
3 |  |  |  |  | Exposure/Dose Measurement | [18 empty] | TO.10.M |  |  |  | O557
4
5 | [23 empty] | CT Protocols
6 | [23 empty] | Adult Abdomen
7 | [23 empty] | Protocol name |  |  |  | O564
8 | [15 empty] | Generator Performance | [7 empty] | kVp |  |  |  | O565
9 | [23 empty] | mA |  |  |  | O566
10 | [11 empty] | Test Equipment Used: | Piranha CB2-17090320 | [10 empty] | mAs |  |  |  | O567
11 | [23 empty] | Eff mAs |  |  |  | O568
12 | [10 empty] | Large Focus kV, Exposure, Timer Reproducibility | [12 empty] | Time/rotation (s) |  |  |  | O569
13 | [23 empty] | Scan FOV (cm) |  |  |  | O570
14 | [10 empty] | Test | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values |  |  |  |  | Display FOV (cm) |  |  |  | O571
15 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  |  | Reconstruction Algorithm |  |  |  | O572
16 | [11 empty] | 100 | Large | 70-120 | 80 | 20 | 0.05 | [6 empty] | Axial/Helical |  |  |  | O573
17 | [11 empty] | 100 | Large | 70-120 | 80 | 20 | 0.05 | [6 empty] | Z-Axis collimation |  |  |  | O574
18 | [11 empty] | 100 | Large | 70-120 | 80 | 20 | 0.05 | [6 empty] | # of channels |  |  |  | O575
19 | [11 empty] | 100 | Large | 70-120 | 80 | 20 | 0.05 | [6 empty] | Table Increment/speed |  |  |  | O576
20 | [11 empty] | 100 | Large | 70-120 | 80 | 20 | 0.05 | [6 empty] | Pitch |  |  |  | O577
21 | [23 empty] | Recon scan width (mm) |  |  |  | O578
22 | [23 empty] | Recon scan interval (mm) |  |  |  | O579
23 | [23 empty] | Dose reduction techniques |  |  |  | O580
24 | [23 empty] | Indiated CTDIvol (mGy) |  |  |  | O581
25 | [23 empty] | Adult Chest
26 | [10 empty] | Note: Type comments in the below shaded cell. | [12 empty] | Protocol name |  |  |  | P564
27 | [11 empty] | Do NOT exceed 190 | [11 empty] | kVp |  |  |  | P565
28 | [11 empty] | New |  | 0 | [9 empty] | mA |  |  |  | P566
29 | [11 empty] | Do NOT exceed 190 | [6 empty] | kVp | mGy | Time |  |  | mAs |  |  |  | P567
30 | [11 empty] | New |  | 0 |  |  |  | Acceptable: | NA | NA | NA |  |  | Eff mAs |  |  |  | P568
31 | [10 empty] | Small Focus kV, Exposure, Timer Reproducibility | [12 empty] | Time/rotation (s) |  |  |  | P569
32 | [23 empty] | Scan FOV (cm) |  |  |  | P570
33 | [10 empty] | Test | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values |  |  |  |  | Display FOV (cm) |  |  |  | P571
34 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  |  | Reconstruction Algorithm |  |  |  | P572
35 | [11 empty] | 100 | Small | 70-120/Low | 90 | 10 | 0.1 | [6 empty] | Axial/Helical |  |  |  | P573
36 | [11 empty] | 100 | Small | 70-120/Low | 90 | 10 | 0.1 | [6 empty] | Z-Axis collimation |  |  |  | P574
37 | [11 empty] | 100 | Small | 70-120/Low | 90 | 10 | 0.1 | [6 empty] | # of channels |  |  |  | P575
38 | [11 empty] | 100 | Small | 70-120/Low | 90 | 10 | 0.1 | [6 empty] | Table Increment/speed |  |  |  | P576
39 | [11 empty] | 100 | Small | 70-120/Low | 90 | 10 | 0.1 | [6 empty] | Pitch |  |  |  | P577
40 | [23 empty] | Recon scan width (mm) |  |  |  | P578
41 |  | LFS Output Slope: |  |  | SFS Output Slope: |  |  | Overall Slope: | [15 empty] | Recon scan interval (mm) |  |  |  | P579
42 | [23 empty] | Dose reduction techniques |  |  |  | P580
43 | [23 empty] | Indiated CTDIvol (mGy) |  |  |  | P581
44 | [23 empty] | Adult Head
45 |  |  |  |  | Image Receptor | [18 empty] | Protocol name |  |  |  | Q564
46 | [11 empty] | Do NOT exceed 190 | [11 empty] | kVp |  |  |  | Q565
47 | Indicated Field Size @SID |  | Measured* Radiation |  | Field Size Indicators |  | Acceptable |  |  |  |  | New |  | 0 | [9 empty] | mA |  |  |  | Q566
48 | (cm) |  | (cm) |  | Error (% SID) |  | Field Size Indicators |  |  |  |  | Do NOT exceed 190 | [11 empty] | mAs |  |  |  | Q567
49 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal |  |  |  | New |  | 0 | [5 empty] | kVp | mGy |  |  | Eff mAs |  |  |  | Q568
50 | 5 | 5 | 5 | 5 | 0% | 0% | YES | YES | [10 empty] | Acceptable: | NA | NA |  |  | Time/rotation (s) |  |  |  | Q569
51 |  |  |  |  | NA | NA | NA | NA |  |  |  | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%. | [10 empty] | Scan FOV (cm) |  |  |  | Q570
52 | *Dimension of highest optical density region | [22 empty] | Display FOV (cm) |  |  |  | Q571
53 | **Dimension between the image of the coins marking the edge of the light field | [22 empty] | Reconstruction Algorithm |  |  |  | Q572
54 | [23 empty] | Axial/Helical |  |  |  | Q573
55 | Criteria: | The indicated field length/width are accurate to within 2% of the SID. | [9 empty] | Combined Comments: | [11 empty] | Z-Axis collimation |  |  |  | Q574
56 |  | X-ray field and light field length/width are congruent within 2% SID. | [9 empty] | Large Focus: | [11 empty] | # of channels |  |  |  | Q575
57 | [23 empty] | Table Increment/speed |  |  |  | Q576
58 | [11 empty] | Small Focus: | [11 empty] | Pitch |  |  |  | Q577
59 | [23 empty] | Recon scan width (mm) |  |  |  | Q578
60 | [23 empty] | Recon scan interval (mm) |  |  |  | Q579
61 | [11 empty] | Test Equipment Used: | Piranha CB2-17090320 | [10 empty] | Dose reduction techniques |  |  |  | Q580
62 | [23 empty] | Indiated CTDIvol (mGy) |  |  |  | Q581
63 | [23 empty] | CTDI Measurements
64 | [10 empty] | Timer performance (Large Focal Spot) | [12 empty] | Adult Abdomen |  |  |  | O599
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [14 empty] | Adult Chest |  |  |  | P599
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | NOTE: Enter "Dent" for Dental, "1*" for 1-phase, "3" for 3-phase, "HF" for High Frequency, "CP" for Constant Potential | [11 empty] | Adult Head |  |  |  | Q599
1 | [9 empty] | Medical University of South Carolina |  | Generator Type: | HF | HF | (High Frequency) | [8 empty] | CTDIvol/100 mAs
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | 2 | 2 | Timer Test is Valid for this generator | [10 empty] | Adult Abdomen |  |  |  | O605
3 | [10 empty] | Test | SDD | Focal | Filter Pack/ |  | Set Values* |  |  | Measured Values* |  |  |  |  | Adult Chest |  |  |  | P605
4 |  |  |  |  | X-RAY TUBE | [5 empty] | Setup | cm | Spot | Sensitivity | kV | mA/mAs | seconds | kV (max) | kV (avg) | seconds | mGy |  |  | Adult Head |  |  |  | Q605
5 | [11 empty] | 100 | Large | 70-120/Low | 80 | 2
6 | [11 empty] | 100 | Large | 70-120/Low | 80 | 4
7 | Focal Spot Resolution | [10 empty] | 100 | Large | 70-120/Low | 80 | 8
8 | [8 empty] | 1 Year | 2 Year |  | 100 | Large | 70-120/Low | 80 | 20
9 | Large: |  | mm |  |  |  |  | Measured | Previous | Previous |  | 100 | Large | 70-120/Low | 80 | 50
10 | [7 empty] | Resolution | Resolution | Resolution |  | 100 | Large | 70-120/Low | 80 | 80
11 |  | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm)
12
13
14 | Small: |  | mm |  |  |  |  | Measured | Previous | Previous
15 | [7 empty] | Resolution | Resolution | Resolution |  | *Record time in seconds for 3-phase, HF, and CP
16 |  | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  | For 1-phase, record time in pulses up to 0.5 sec. (60 pulses), and seconds thereafter.
17
18 |  | * SID = Source to Image (film) Distance. | [9 empty] | Comments:
19 |  | ** OFD = Object (resolution pattern) to Film Distance. | [5 empty] | Acceptable: | TBD |  | Note: Type comments in the above shaded cell.
20 | [11 empty] | Additional comments:
21 |  | Criteria: | No significant change in resolution since previous measurement. | [8 empty] | New |  | 0
22 | [15 empty] | Time (sec) |  | Error | 3-Phase | 1-Phase | Dent
23 | Comments: | [14 empty] | Indicated | Measured | (Percent) | Acceptable | Acceptable | Acceptable
24
25
26
27
28
29 |  |  |  |  | GENERATOR PERFORMANCE
30 | [13 empty] | Combined Comments
31
32 |  | Test Equipment Used: | Piranha CB2-17090320 |  |  | Calibration Date: |  |  | Calibration Due:
33
34 | kVp Accuracy | [10 empty] | Timer Specifications
35 | [11 empty] | Generator |  |  |  | Generator
36 | [11 empty] | Type | time (ms) | error |  | Type | mAs | error
37 |  | Large Focal Spot |  |  |  |  | Small Focal Spot |  |  |  |  | 1f | < =4 pulses | 0 | pulses | Cine | < 2 | 15%
38 | [12 empty] | 5-24 pulses | 1 | pulse |  | 2 - 5 | 10%
39 | SDD: | 100 | cm |  |  | SDD: | 100 | cm |  |  |  |  | 25-60 pulses | 2 | pulses |  | > 5 | 5%
40 | mA: |  |  |  |  | mA: | [6 empty] | >60 pulses | 3 | pulses | mAs timer | < 2 | 20%
41 | mAs: | 20 |  |  |  | mAs: | 10 |  |  |  |  | 3f | < 2 | 25% |  |  | 2 - 5 | 15%
42 | time (s): | 0.05 |  |  |  | time (s): | 0.1 | [5 empty] | 2 - 10 | 20% |  |  | 5 - 100 | 10%
43 | *SDD=Source to detector distance |  |  |  |  | *SDD=Source to detector distance | [6 empty] | 10 - 50 | 15% |  |  | > 100 | 5%
44 | [12 empty] | 50 - 100 | 10%
45 | Indicated | Measured |  | kV (max) |  | Indicated | Measured |  | kV (max) |  |  |  | > 100 | 5%
46 | kVp | kV (max) | kV (avg) | % error |  | kVp | kV (max) | kV (avg) | % error |  |  | Dental | All | +2 | pulses
47 | 60 |  |  |  |  | 50 | [6 empty] | All | +15%
48 | 80 |  |  |  |  | 70
49 | 100 |  |  |  |  | 90
50 | 120 |  |  |  |  | 110
51 | 140 |  |  |  |  | 130 | [5 empty] | Test Equipment Used: | Piranha CB2-17090320
52
53 | [10 empty] | Large Focus Linearity
54
55
56 |  |  | Acceptable: | TBD |  |  |  | Acceptable: | TBD |  | mA/mAs | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
57 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
58 |  | Criteria: | kVp accurate to within 5% of indicated. | [7 empty] | Adjacent* | 100 | Large | 70-120/Low | 80 | 20
59 | [10 empty] | Adjacent* | 100 | Large | 70-120/Low | 80 | 25
60 | [10 empty] | Non-Adj. | 100 | Large | 70-120/High | 80 | 2.5
61 | Comments: | [9 empty] | Non-Adj. | 100 | Large | 70-120/Low | 80 | 40
62 | [10 empty] | Non-Adj. | 100 | Large | 70-120/Low | 80 | 12.5
63 | [10 empty] | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
64 | [19 empty] | Acceptable
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Comments: | [7 empty] | Adjacent
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | Do NOT exceed 390 |  | 0 | [5 empty] | TBD
1 | [9 empty] | Medical University of South Carolina |  | New | [7 empty] | Non-Adj.
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | [9 empty] | TBD
3 | [10 empty] | Small Focus Linearity
4
5 |  |  |  |  | Generator Performance | [5 empty] | mA/mAs | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
6 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
7 |  | Test Equipment Used: | Piranha CB2-17090320 | [7 empty] | Adjacent* | 100 | Small | 70-120/Low | 80 | 10
8 | [10 empty] | Adjacent* | 100 | Small | 70-120/Low | 80 | 11
9 | kVp and Exposure Reproducibility | [9 empty] | Non-Adj. | 100 | Small | 70-120/High | 80 | 25
10 | [10 empty] | Non-Adj. | 100 | Small | 70-120/High | 80 | 5
11 |  |  | Large Focal Spot |  |  |  | Small Focal Spot |  |  |  | Non-Adj. | 100 | Small | 70-120/Low | 80 | 16
12 |  |  | SDD*: | 100 | cm |  | SDD*: | 100 | cm |  | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
13 |  |  | mA: |  |  |  | mA: | [12 empty] | Acceptable
14 |  |  | mAs: | 20 |  |  | mAs: | 10 |  |  |  | Comments: | [7 empty] | Adjacent
15 |  |  | time (s): | 0.05 |  |  | time (s): | 0.1 |  |  |  | Do NOT exceed 390 |  | 0 | [5 empty] | TBD
16 | [11 empty] | New | [7 empty] | Non-Adj.
17 |  |  | Indicated | Measured |  |  | Indicated | Measured | [11 empty] | TBD
18 |  |  | kVp | kVp | mGy |  | kVp | kVp | mGy
19 |  |  | 80 |  |  |  | 90
20
21 | [11 empty] | Coefficient of Linearity | [7 empty] | Combined Focal Spot Linearity
22 | [12 empty] | Adjacent: |  | (mGy/mAs)adj1 - (mGy/mAs)adj2
23 | [14 empty] | (mGy/mAs)adj1 + (mGy/mAs)adj2 |  |  |  |  | Acceptable
24 | [19 empty] | Non-Adj.
25 | [12 empty] | Non-Adjacent: |  | (mGy/mAs)max - (mGy/mAs)min |  |  |  |  | NA
26 | [14 empty] | (mGy/mAs)max + (mGy/mAs)min
27
28 | [11 empty] | Criteria: | Coefficient of linearity <= 10% (Both adjacent and non-adjacent).
29
30 |  |  | mean: |  |  |  | mean:
31 |  |  | std dev: |  |  |  | std dev:
32 |  |  | Coeff. of var. | TBD | TBD |  | Coeff. of var. | TBD | TBD
33 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA
34 |  |  | Acceptable: | NA | NA |  | Acceptable: | NA | NA | [5 empty] | Generator Functions
35 | [12 empty] | Means provided to indicate technique factors prior to exposure.
36 |  | *SDD = Source to Detector Distance | [10 empty] | All technique indicators functional and are free of excessive wear.
37 | [12 empty] | Technique overload circuits function to block excessive technique
38 |  |  | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%. | [6 empty] | 3 |  | Anode heat calculator/indicator is functioning properly
39 | [12 empty] | Indicator of the individual x-ray tube selected present and operational
40 | [12 empty] | Visible "beam on" indicator present and operational
41 | [12 empty] | Audible exposure end indicator present and operational
42 | [12 empty] | Exposure delay time is not excessive.
43 | Comments:
44
45 | [10 empty] | 2 | Change to DHEC Standards?
46 | [10 empty] | (Type 1 to base acceptance on DHEC criteria)
47 | [17 empty] | Measured HVL based on:
48 | [10 empty] | Half Value Layer Measurement |  |  |  | Push button to select if HVL is |  |  |  | Max kV
49 | [14 empty] | based on Max kV or Avg kV
50 | [11 empty] | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
51 | [10 empty] | mm Al | cm | Spot | Sensitivity | kV | mA/mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
52 | [10 empty] | 0 | 100 | Large | 70-120/Low | 80 | 20 | 0.05
53 | [10 empty] | 0 | 100 | Large | 70-120/Low | 80 | 20 | 0.05
54 | [10 empty] | 3 | 100 | Large | 70-120/Low | 80 | 20 | 0.05
55 | [10 empty] | 3 | 100 | Large | 70-120/Low | 80 | 20 | 0.05
56
57
58
59
60
61
62
63
64 | [11 empty] | Comments:
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Do NOT exceed 390 |  | 0
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | New
1 | [9 empty] | Medical University of South Carolina | Data Processing: |  |  |  | Max kV
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp |  | HVL @80 kVp (Max kV) |  | Criteria: MUSC STANDARD |  |  | Last Year:
3 | [10 empty] | 0 |  |  |  | 80 |  | #DIV/0! |  | HVL | MIN. HVL | MAX. HVL
4 | [10 empty] | 3 | [6 empty] | kVp | (mm Al) | (mm Al) | (mm Al)
5 |  | Test Equipment Used: | Piranha CB2-17090320 | [12 empty] | Acceptable (mm Al) |  | 60 | 2.1855 | 1.9 | 2.4
6 | [13 empty] | Data Exp | LN(EXP) | LOW: | 2.6 | 70 | 2.6109 | 2.3 | 2.9
7 | [15 empty] | DESIRED: | 3 | 80 | 3.0364 | 2.6 | 3.4
8 | [5 empty] | Timer performance (Large Focal Spot) | [9 empty] | HIGH: | 3.4 | 90 | 3.4618 | 3 | 3.8
9 | [10 empty] | #DIV/0! | #DIV/0! | [5 empty] | 100 | 3.8873 | 3.4 | 4.3
10 | [15 empty] | Acceptable: | TBD | 110 | 4.3127 | 3.8 | 4.8
11 |  |  | Generator Type: | HF | (High Frequency) | [10 empty] | Meets minimum HVL criteria: | TBD | 120 | 4.7382 | 4.2 | 5.2
12 | [17 empty] | 130 | 5.1636 | 4.6 | 5.7
13 |  |  |  |  | Accuracy |  |  |  | Reproducibility
14
15 |  | SDD: | 100 | cm |  |  |  | SDD: | 100 | cm
16 |  | kVp: | 80 |  |  |  |  | kVp: | 80
17 | [7 empty] | mA: |  |  | Output - Large Focal Spot | [5 empty] | Calculated |  | (Using |  |  | Using | [9 empty] | Using
18 | [7 empty] | mAs: | 20 | [6 empty] | Current | Previous |  | Previous | (Using | (Using | Current | [8 empty] | (Using | Current
19 | [10 empty] | Previous* Measurements |  |  | Current Meas. |  | mGy/mAs | mGy/mAs |  | mGy/mAs | Recorded | Current | mGy/mAs | [8 empty] | Current | mGy/mAs
20 | Measured Values |  |  | Time (sec) |  | Error |  |  | Time (sec) |  | SDD (in.): |  |  | SDD (cm): | 100 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current | [8 empty] | Measured | @ Current
21 | kV (max) | kV (avg) | mA/mAs | Indicated | Measured | (Percent) | Acceptable |  | Indicated | Measured |  | mGy/mAs | mGy/mAs |  | mGy/mAs | kVp | meas. kVp |  | 40" SID) | kVp) | kVp) | kVp | [8 empty] | kVp) | kVp
22 |  |  | 2 | [5 empty] | 0.05 |  | kVp | @SDD | @40" | kV (max) | @SDD | @40" | @40" |  | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40" | [8 empty] | LN(kVp) | @40"
23 |  |  | 4
24 |  |  | 8
25 |  |  | 20
26 |  |  | 50
27 |  |  | 80
28
29
30
31 | [10 empty] | *From previous compliance report |  |  |  | Max | Median | Min.
32 | [14 empty] | NA | NA | NA |  | Acceptable: | NA
33 | [8 empty] | mean: |  |  |  |  | Last Year:
34 | [8 empty] | std dev: |  | Output - Small Focal Spot | [5 empty] | Calculated |  | (Using |  |  | Using
35 | [8 empty] | Coeff. of var.: | TBD | [5 empty] | Current | Previous |  | Previous | (Using | (Using | Current
36 | [10 empty] | Previous* Measurements |  |  | Current Meas. |  | mGy/mAs | mGy/mAs |  | mGy/mAs | Recorded | Current | mGy/mAs
37 | [5 empty] | Acceptable: | NA |  | Acceptable: | NA | SDD (in.): |  |  | SDD (cm): | 100 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current
38 | Comments: | [10 empty] | mGy/mAs | mGy/mAs |  | mGy/mAs | kVp | meas. kVp |  | 40" SID) | kVp) | kVp) | kVp
39 | [10 empty] | kVp | @SDD | @40" | kV (max) | @SDD | @40" | @40" |  | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40"
40
41
42
43
44
45 |  | Timer Specifications
46 |  | Generator |  |  |  | Generator |  |  |  |  | *From previous compliance report |  |  |  | Max | Median | Min.
47 |  | Type | time (ms) | error |  | Type | mAs | error | [6 empty] | NA | NA | NA |  | Acceptable: | NA
48 |  | 1f | < =4 pulses | 0 | pulses | Cine | < 2 | 15% | [5 empty] | Last Year:
49 |  |  | 5-24 pulses | 1 | pulses |  | 2 - 5 | 10%
50 |  |  | 25-60 pulses | 2 | pulses |  | > 5 | 5% |  |  | Enter 1 for YES, 2 for NO, 3 for NA
51 |  |  | >60 pulses | 3 | pulses | mAs timer | < 2 | 20% | [6 empty] | Radiation Safety
52 |  | 3f | < 2 | 25% |  |  | 2 - 5 | 15% |  |  |  |  | Continuous pressure on exposure switch required to make an exposure with selected tube
53 |  |  | 2 - 10 | 20% |  |  | 5 - 100 | 10% |  |  | 3 |  | Spot film exposure blocked when there is no cassette
54 |  |  | 10 - 50 | 15% |  |  | > 100 | 5% |  |  | 3 |  | Spot film exposure blocked when serial exposures are completed
55 |  |  | 50 - 100 | 10% | [6 empty] | 3 |  | Spot film device provides acceptable centering to the image receptor
56 |  |  | > 100 | 5% | [8 empty] | Exposure switch for selected tube is secured at a safe location | [6 empty] | NEW
57 |  | Dental | All | +2 | pulses | [7 empty] | Operator has full visibility of the exposure factors and patient | [7 empty] | Half
58 |  |  | All | +15% | [6 empty] | 3 |  | Lead apron available |  | Number: | 0 | Half | 0 | Full |  | Full
59 | [10 empty] | 3 |  | Lead gloves available. |  | Number: | 0 | (Left + Right) |  |  |  | (Left + Right)
60 | [10 empty] | 3 |  | A properly designed and installed apron rack is present.
61 | [10 empty] | 3 |  | Documentation of annual protective apparel integrity inspection is available.
62 | [12 empty] | Patient restraint devices available.
63 | [10 empty] | 3 |  | Gonadal shielding is present and functional.
64 | [10 empty] | 3 |  | Radiation warning light at room entrance is functional with selected tube
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID: | [6 empty] | Physical Inspection
1 | [9 empty] | Medical University of South Carolina | 1 |  | System clean and free of debris.
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | 1 |  | All cables free from excessive wear or restraints.
3 | [10 empty] | 1 |  | System free of missing or damaged components.
4 | [10 empty] | 1 |  | Table movements and locks acceptable
5 |  | Test Equipment Used: | Piranha CB2-17090320 | [7 empty] | 1 |  | Tube crane movements and locks acceptable.
6 | [10 empty] | 1 |  | All other movements and locks functional.
9
8 | Large Focus Linearity |  |  | kVp: | 80 | kV filter: | 70-120/Low
9 |  |  |  | SDD: | 100 | cm |  |  | Large Focus |  | IMAGE QUALITY
10 | Test |  | Set |  | Measured values |  |  | mGy/mAs | Coeff. of
11 | Setup | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | @ SDD | Linearity | Acceptable | Leeds Phantom |  |  |  |  | Detector Uniformity
12 | Adjacent* | 20 | [7 empty] | Adjacent | [7 empty] | Dev from mean
13 | Adjacent* | 25 | [7 empty] | TBD | TO.16 | kV | mAs |  |  |  | Mean
14 | Non-Adj. | 2.5 | [9 empty] | 70 |  |  |  | ROI 1:
15 | Non-Adj. | 40 | [7 empty] | Non-Adj. |  | 70 |  |  |  | ROI 2:
16 | Non-Adj. | 12.5 | [7 empty] | TBD | [5 empty] | ROI 3:
17 | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X. | [9 empty] | Row | Disks | Previous |  |  | ROI 4:
18 | [10 empty] | A |  |  |  |  | ROI 5:
19 | Comments: | [9 empty] | B |  |  |  |  | Average:
20 | [10 empty] | C |  |  |  |  | Max deviation:
21 | [10 empty] | D |  |  |  |  | Acceptable:
22 | [10 empty] | E
23 | Small Focus Linearity |  |  | kVp: | 80 | kV filter: | 70-120/Low |  |  |  | F |  |  |  |  | Criteria:
24 |  |  |  | SDD: | 100 | cm |  |  | Small Focus |  | G |  |  |  |  | Measure with 1.5 mm Cu filter in the beam.
25 | Test |  | Set |  | Measured values |  |  | mGy/mAs | Coeff. of |  | H
26 | Setup | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | @ SDD | Linearity | Acceptable | J
27 | Adjacent* | 10 | [7 empty] | Adjacent | K
28 | Adjacent* | 11 | [7 empty] | TBD | L
29 | Non-Adj. | 25 | [8 empty] | M
30 | Non-Adj. | 5 | [7 empty] | Non-Adj.
31 | Non-Adj. | 16 | [7 empty] | TBD | COMPUTED TOMOGRAPHY
32 | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
33 | [10 empty] | Measurement Protocols
34 | Comments:
35 | [13 empty] | Abdomen | Chest | Head | Abdomen | Chest | Head
36 | [12 empty] | Protocol name:
37 | [12 empty] | kVp:
38 | Combined Focal Spot Linearity | [11 empty] | mA:
39 |  |  |  | Combined | [8 empty] | mAs:
40 |  | mGy/mAs |  | Coeff. of | [8 empty] | Eff mAs:
41 |  | @ SDD |  | Linearity |  | Acceptable | [6 empty] | Rotation time (s):
42 |  | Max.: | 0 | Non-Adj. |  | Non-Adj. | [6 empty] | Scan FOV (cm):
43 |  | Min.: | 0 | NA |  | NA | [6 empty] | Display FOV (cm):
44 | [12 empty] | Reconstruction algorithm:
45 |  | Coefficient of Linearity | [10 empty] | Axial (A)/Helical (H):
46 |  |  | Adjacent: |  | (mGy/mAs)adj1 - (mGy/mAs)adj2 | [7 empty] | Z-Axis collimation (T):
47 |  |  |  |  | (mGy/mAs)adj1 + (mGy/mAs)adj2 | [7 empty] | # of channels (N):
48 | [12 empty] | Table Increment/speed (I):
49 |  |  | Non-Adjacent: |  | (mGy/mAs)max - (mGy/mAs)min | [7 empty] | Pitch (p):
50 |  |  |  |  | (mGy/mAs)max + (mGy/mAs)min | [7 empty] | Recon scan width (mm):
51 | [12 empty] | Recon scan interval (mm):
52 |  | Criteria: | Coefficient of linearity <= 10% (Both adjacent and non-adjacent). | [9 empty] | Dose reduction techniques:
53 | [12 empty] | Indiated CTDIvol (mGy):
54 | Rule Number |  |  |  | Generator Functions |  |  |  | Compliance
55 | DHEC RHB 4.3.9.1 |  | Means are provided to indicate technique factors prior to exposure. | [5 empty] | TBD |  | CTDI Measurements
56 |  |  | All technique indicators functional and are free of excessive wear. | [5 empty] | TBD |  |  | Test Equipment Used: | Piranha CB2-17090320
57 |  |  | Technique overload circuits function to block excessive technique. | [5 empty] | TBD |  |  | Chamber length (mm): | 100
58 |  |  | Anode heat calculator/indicator is functioning properly. | [5 empty] | NA |  |  | Chamber correction factor: | 1
59 | DHEC RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational. | [5 empty] | TBD
60 | DHEC RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational. | [5 empty] | TBD |  |  |  | Abdomen | Chest | Head |  |  |  |  | Reference | Pass/Fail
61 | DHEC RHB 4.7.4.2.4 |  | Audible exposure end indicator present and operational. | [5 empty] | TBD |  |  | Indicated CTDIvol (mGy): | [7 empty] | (mGy) | (mGy)
62 |  |  | Exposure delay time is not excessive. | [5 empty] | TBD |  |  | Center (mGy) | [6 empty] | Adult Abdomen | 25 | 30
63 | [18 empty] | Adult Head: | 75 | 80
64 | [18 empty] | Ped Abd (16cm) | 15 | 20
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Center CTDI (mGy): | [6 empty] | Ped Abd (32cm) | 7.5 | 10
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | Periphery (mGy) | [6 empty] | Pediatric Head: | 35 | 40
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 |  |  |  |  | HALF VALUE LAYER MEASUREMENT | [6 empty] | Periphery CTDI (mGy):
4 | [11 empty] | CTDIw (mGy):
5 | [11 empty] | CTDIvol (mGy):
6 |  |  | Large Focal Spot | [8 empty] | Phantom size indicated (cm) | 32 | 32 | 16
7 |  | SDD*: | 100 | cm |  |  | mA: |  |  |  |  | Phantom size used (cm) | 32 | 32 | 16
8 |  | Set kVp: | 80 |  |  |  | mAs: | 20 |  |  |  | Previous year:
9 |  |  |  |  | Set |  | time (s): | 0.05 |  |  |  | Variation from prev year:
10 |  | Al (mm) | mGy | mGy | kVp | [6 empty] | Variation from indicated:
11 |  | 0 |  |  | 80 |  | kVp (Max kV) | 80 |  |  |  | Acceptable:
12 |  | 3 |  |  |  |  | Meas. HVL: | #DIV/0! |  |  |  | CTDIvol/100 mAs:
13 | [6 empty] | Meets min HVL: | TBD |  |  |  | Prev CTDIvol/100 mAs:
14 | [11 empty] | Criteria: | CTDI is within 5% of previous year.
15 | [6 empty] | Criteria: MUSC STANDARD | [5 empty] | Measured CTDI is within 20% of indicated CTDI
16 | [6 empty] | HVL | MIN. HVL | MAX. HVL |  |  |  | CTDIvol is less than ACR Pass/Fail dose
17 |  | * SDD = Source to Detector Distance. |  |  |  | kVp | (mm Al) | (mm Al) | (mm Al)
18 | [5 empty] | 60 | 2.1855 | 1.9 | 2.4 |  | Beam Profile
19 | [5 empty] | 70 | 2.6109 | 2.3 | 2.9 |  |  | Channels (N):
20 |  |  | **ACCEPTABLE RESULTS (mm Al) |  |  | 80 | 3.0364 | 2.6 | 3.4 |  |  | Detector width T (mm):
21 |  |  |  | Min. acceptable HVL at 80 kvp: | 2.6 | 90 | 3.4618 | 3 | 3.8 |  |  | NxT Beam width (mm):
22 |  |  |  | Desired HVL at meas. 80 kVp: | 3 | 100 | 3.8873 | 3.4 | 4.3 |  |  | Nominal beam width (mm):
23 |  |  |  | Max. acceptable HVL at 80 kvp: | 3.4 | 110 | 4.3127 | 3.8 | 4.8 |  | Measured beam width (mm)
24 | [5 empty] | 120 | 4.7382 | 4.2 | 5.2
25 | [5 empty] | 130 | 5.1636 | 4.6 | 5.7
26 | Comments: | [10 empty] | Average (mm):
27 | [11 empty] | Acceptable:
28 | [11 empty] | Criteria: | Beam profile is within 3mm or 30% of specified value
29
30 | [10 empty] | Artifact Evaluation
31 | TO.16 Contrast Resolution | [5 empty] | Detector Uniformity |  |  |  |  | Water phantom is free from rings, streaks, lines, cupping artifacts
32 |  |  |  | Contrast Detail | Threshold Index |  |  | Mean
33 | Row | Diameter (mm) | sqrt(Area) |  |  |  | ROI 1: |  |  |  | CT Number Accuracy
34 | A | 11.1 | 9.8371 |  |  |  | ROI 2: |  |  |  | kVp: | 80 | 100 | 110 | 120 | 130 | 140 |  |  | Acceptable HU
35 | B | 8 | 7.0898 |  |  |  | ROI 3: |  |  |  | mAs: | [8 empty] | Low | High
36 | C | 5.6 | 4.9629 |  |  |  | ROI 4: |  |  |  | Polyethylene: | [7 empty] | Polyethylene | -107 | -84
37 | D | 4 | 3.5449 |  |  |  | ROI 5: |  |  |  | Water: | [7 empty] | Water | -7 | 7
38 | E | 2.8 | 2.4814 |  |  |  | Average: |  |  |  | Water Std Dev: | [7 empty] | Acrylic | 110 | 135
39 | F | 2 | 1.7725 |  |  |  | Max deviation: |  |  |  | Acrylic: | [7 empty] | Bone | 850 | 970
40 | G | 1.4 | 1.2407 |  |  |  | Acceptable: |  |  |  | Bone: | [7 empty] | Air | -1005 | -970
41 | H | 1 | 0.88623 | [7 empty] | Air:
42 | J | 0.7 | 0.62036 | [10 empty] | Acceptable: |  | Enter 1 for YES, 2 for NO
43 | K | 0.5 | 0.44311 | [8 empty] | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
44 | L | 0.35 | 0.31018 | [9 empty] | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
45 | M | 0.25 | 0.22156
46 | [10 empty] | CT Slice Thickness Accuracy
47 | [11 empty] | Slice thickness (mm):
48 | [11 empty] | Top (mm):
49 | [11 empty] | Bottom (mm):
50 | [11 empty] | Average (mm):
51 | [11 empty] | Water HU:
52 | [11 empty] | Water HU Std Dev:
53 | [11 empty] | Acceptable:
54 | [12 empty] | Criteria: | Slice thickness is within 1.5mm of the set thickness
55 | [13 empty] | CT number for water is 0 +/- 5
56
57 | [10 empty] | Low Contrast Performance
58 | [11 empty] | Adult
59 | [11 empty] | Abdomen | Head |  |  |  |  | Enter 1 for YES, 2 for NO, 3 for Not Applicable
60 | [10 empty] | Inside | [5 empty] | 6mm rods visible:
61 | [10 empty] | Adjacent | [5 empty] | Acceptable:
62 | [10 empty] | Adj StDev | [5 empty] | Criteria: | All four 6mm rods are visible
63 | [10 empty] | Contrast | [6 empty] | CNR > 1.0 for adult abdomen and head protocols
64 | [10 empty] | CNR
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID: |  |  | Uniformity
1 | [9 empty] | Medical University of South Carolina |  | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | Center:
3 | Output - Large Focal Spot |  |  | Current | Previous | [5 empty] | Top:
4 |  |  |  | mGy/mAs | mGy/mAs | [5 empty] | Bottom:
5 |  |  | Current | @ Current | @ Current | [5 empty] | Right:
6 |  |  | Meas. | kVp | kVp | Percent |  |  |  |  | Left:
7 |  |  | kVp | @40" | @40" | Change | Acceptable |  |  |  |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region
8 | [11 empty] | Uniformity image is free from artifacts
9
10 | [10 empty] | High Contrast Resolution
11 | [11 empty] | lp/cm | Acceptable
12 | [10 empty] | Abdomen | [5 empty] | Critiera: | 6 lp/cm for abdomen
13 | [10 empty] | Hi-res | [6 empty] | 8 lp/cm for high resolution
14
15 | [7 empty] | Acceptable: | NA
16 | Output - Small Focal Spot |  |  | Current | Previous
17 |  |  |  | mGy/mAs | mGy/mAs
18 |  |  | Current | @ Current | @ Current
19 |  |  | Meas. | kVp | kVp | Percent
20 |  |  | kVp | @40" | @40" | Change | Acceptable
21
22
23
24
25
26 | [7 empty] | Acceptable: | NA
27
28 |  | Criteria: | The x-ray output @40" within +/-25% from previous measurements.
29 | Comments:
30
31
32
33 |  |  | Lrg. Focus |  |  |  |  | Sm. Focus
34 |  |  | Current |  |  |  |  | Current
35 |  |  | mGy/mAs |  |  |  |  | mGy/mAs
36 |  | kVp | @40" |  |  |  | kVp | @40"
37 |  | 55 |  |  |  |  | 45
38 |  | 60 |  |  |  |  | 55
39 |  | 70 |  |  |  |  | 65
40 |  | 80 |  |  |  |  | 75
41 |  | 90 |  |  |  |  | 85
42 |  | 100 |  |  |  |  | 95
43 |  | 110 |  |  |  |  | 105
44 |  | 120 |  |  |  |  | 115
45 |  | 130 |  |  |  |  | 125
46 |  | 140 |  |  |  |  | 135
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | Protocols |  | Abdomen | Chest | Head
4 |  | kVp
5 |  | mA
6 |  | mAs
7 |  | Eff mAs
8 |  | Time/rotation (s)
9 |  | Scan FOV (cm)
10 |  | Display FOV (cm)
11 |  | Reconstruction Algorithm
12 |  | Axial/Helical
13 |  | Z-Axis collimation
14 |  | # of channels
15 |  | Table Increment/speed
16 |  | Pitch
17 |  | Recon scan width (mm)
18 |  | Recon scan interval (mm)
19 |  | Dose reduction techniques
20 |  | Indiated CTDIvol (mGy)
21
22 | CTDI Measurements
23 |  | Test Equipment Used: | Piranha CB2-17090320
24
25
26 |  |  |  | Abdomen | Chest | Head |  |  | Reference | Pass/Fail
27 |  |  | Indicated CTDIvol (mGy): | [5 empty] | (mGy) | (mGy)
28 |  |  | Center (mGy): |  |  |  |  | Adult Abdomen | 25 | 30
29 |  |  | Periphery (mGy): |  |  |  |  | Adult Head: | 75 | 80
30 |  |  | Average CTDIvol: |  |  |  |  | Ped Abd (16cm) | 15 | 20
31 |  |  | Phantom size indicated (cm): | 32 | 0 | 0 |  | Ped Abd (32cm) | 7.5 | 10
32 |  |  | Phantom size used (cm): | 32 | 0 | 0 |  | Pediatric Head: | 35 | 40
33 |  |  | Variation from previous year:
34 |  |  | Variation from indicated CTDI:
35 |  |  | Acceptable:
36 |  |  | CTDIvol/mAs:
37 |  |  | Criteria: | CTDI is within 5% of previous year.
38 |  |  |  | Measured CTDI is within 20% of indicated CTDI
39 |  |  |  | CTDIvol is less than ACR Pass/Fail dose
40
41
42 | Beam Profile
43 |  | Channels (N)
44 |  | Detector width T (mm)
45 |  | NxT Beam width (mm)
46 |  | Nominal beam width (mm)
47 |  | Average beam width (mm)
48
49 |  | Acceptable:
50 |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value
51
52 | Artifact Evaluation
53 |  | Water phantom is free from rings, streaks, lines, cupping artifacts
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | CT Number Accuracy
4 |  | kVp | 80 | 100 | 110 | 120 | 130 | 140
5 |  | mAs
6 |  | Polyethylene
7 |  | Water
8 |  | Water Std Dev
9 |  | Acrylic
10 |  | Bone
11 |  | Air
12 |  |  |  |  | Acceptable:
13 |  |  | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
14 |  |  |  | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
15
16 | CT Slice Thickness Accuracy
17 |  |  | Slice thickness (mm):
18 |  |  | Top (mm):
19 |  |  | Bottom (mm):
20 |  |  | Average (mm):
21 |  |  | Water HU:
22 |  |  | Water HU Std Dev:
23 |  |  | Acceptable:
24 |  |  |  | Criteria: | Slice thickness is within 1.5mm of the set thickness
25 |  |  |  |  | CT number for water is 0 +/- 5
26
27 | Low Contrast Performance
28 |  |  | Adult
29 |  |  | Abdomen | Head
30 |  | Inside | [5 empty] | 6mm rods visible:
31 |  | Adjacent | [5 empty] | Acceptable:
32 |  | Adjacent StDev
33 |  | Contrast
34 |  | CNR
35 |  |  | Criteria: | All four 6mm rods are visible
36 |  |  |  | CNR > 1.0 for adult abdomen and head protocols
37
38
39 | Uniformity
40 |  | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable
41 | Center:
42 | Top:
43 | Bottom:
44 | Right:
45 | Left:
46 |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region
47 |  | Uniformity image is free from artifacts
48
49 | High Contrast Resolution
50 |  |  | lp/cm | Acceptable
51 |  | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen
52 |  | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
Note on U464: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on H609: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.

## Tables
3 | Enter 1 - Manual, 2 - Radcal, 3 - Barracuda |  |  |  | Manual |  |  |  | Radcal System 2000 |  |  | Barracuda MPD | [6 empty] | Huttner Grating |  |  |  | RADIOGRAPHIC UNITS (MUSC CRITERIA) |  |  |  |  | RAD / DENTAL UNITS (DHEC CRITERIA)
[7 empty] | Exp |  |  | Exp | [8 empty] | 1.1 | 0.5 |  |  | Criteria: |  |  |  |  | Criteria:
kVp Set | mA Set | Time | mAs Set | Add Filt | kVp Eff | kVp Max | Time | Exposure (mGy) | Exposure (mGy) | Time | kVp Max | kVp | ms | Dose (mGy) | Dose rate (mGy/s) | Total filtration (mm Al) | HVL (mm Al) |  | 1.2 | 0.56 |  |  | HVL | MIN. HVL | MAX. HVL |  |  | HVL | MIN. HVL | MAX. HVL
60 | 400 | 0.05 | 20 | 0 | [14 empty] | 1.3 | 0.63 |  | kVp | (mm Al) | (mm Al) | (mm Al) |  | kVp | (mm Al) | (mm Al) | (mm Al)
60 | 400 | 0.05 | 20 | 0 | [14 empty] | 2.1 | 0.71 |  | 60 | 2.1855 | 1.9 | 2.4 |  | 50 | 1.5 | 1.5 | 1.7
80 | 400 | 0.05 | 20 | 0 | [14 empty] | 2.2 | 0.8 |  | 70 | 2.6109 | 2.3 | 2.9 |  | 60 | 1.5 | 1.5 | 1.7
80 | 400 | 0.05 | 20 | 0 | [14 empty] | 2.3 | 0.9 |  | 80 | 3.0364 | 2.6 | 3.4 |  | 70 | 1.5 | 1.5 | 1.7
80 | 400 | 0.05 | 20 | 0 | [14 empty] | 3.1 | 1 |  | 90 | 3.4618 | 3 | 3.8 |  | 80 | 2.3 | 2.3 | 2.7
80 | 400 | 0.05 | 20 | 0 | [14 empty] | 3.2 | 1.12 |  | 100 | 3.8873 | 3.4 | 4.3 |  | 90 | 2.5 | 2.5 | 3
80 | 400 | 0.05 | 20 | 0 | [14 empty] | 3.3 | 1.25 |  | 110 | 4.3127 | 3.8 | 4.8 |  | 100 | 2.7 | 2.7 | 3.2
80 | 400 | 0.05 | 20 | 3 | [14 empty] | 4.1 | 1.4 |  | 120 | 4.7382 | 4.2 | 5.2 |  | 110 | 3 | 3 | 3.5
80 | 400 | 0.05 | 20 | 3 | [14 empty] | 4.2 | 1.6 |  | 130 | 5.1636 | 4.6 | 5.7 |  | 120 | 3.2 | 3.2 | 3.75
80 | 400 | 0.05 | 20 | 3.5 | [14 empty] | 4.3 | 1.8
80 | 400 | 0.05 | 20 | 3.5 | [14 empty] | 5.1 | 2
80 | 400 | 0.05 | 20 | 2.5 | [14 empty] | 5.2 | 2.24
80 | 400 | 0.05 | 20 | 2.5 | [14 empty] | 5.3 | 2.5
80 | 500 | 0.05 | 25 | 0 | [14 empty] | 6.1 | 2.8
80 | 50 | 0.05 | 2.5 | 0 | [14 empty] | 6.2 | 3.15
80 | 800 | 0.05 | 40 | 0 | [14 empty] | 6.3 | 3.55
80 | 250 | 0.05 | 12.5 | 0 | [14 empty] | 7.1 | 4
100 | 400 | 0.05 | 20 | 0 | [14 empty] | 7.2 | 4.5
100 | 400 | 0.05 | 20 | 0 | [14 empty] | 7.3 | 5
120 | 400 | 0.05 | 20 | 0
120 | 400 | 0.05 | 20 | 0
140 | 400 | 0.05 | 20 | 0
50 | 100 | 0.1 | 10 | 0
70 | 100 | 0.1 | 10 | 0
90 | 100 | 0.1 | 10 | 0
90 | 100 | 0.1 | 10 | 0
90 | 100 | 0.1 | 10 | 0
90 | 100 | 0.1 | 10 | 0
90 | 100 | 0.1 | 10 | 0
80 | 100 | 0.1 | 10 | 0
80 | 110 | 0.1 | 11 | 0
80 | 250 | 0.1 | 25 | 0
80 | 50 | 0.1 | 5 | 0
80 | 160 | 0.1 | 16 | 0
110 | 100 | 0.1 | 10 | 0
130 | 100 | 0.1 | 10 | 0
80 | 200 | 0.01 | 2 | 0
80 | 200 | 0.02 | 4 | 0
80 | 200 | 0.04 | 8 | 0
80 | 200 | 0.1 | 20 | 0
80 | 200 | 0.25 | 50 | 0
80 | 200 | 0.4 | 80 | 0
[27 empty] | Exp |  | mAs | Measurement |  | =HVL data
Mode | kVp Set | mA Set | Time | mAs Set | Filter Set | Add Filt | Material | Distance | Units | Exp. CF. | Sensitivity | Chamber Type | Delay | Wave | Focal Spot | Technique Comments | Linearity | Accuracy | Beam Quality | Parameters | Reproduci... | Survey | Wave File | kVp Avg | kVp Eff | kVp Max | Time (s) | Exposure (mGy) |  | Utilization |  | =Small focus data
Radio | 60 | 400 | 0.05 | 20 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV) | Accu(kV)=Accuracy of kVp
Radio | 60 | 400 | 0.05 | 20 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 1 | 0 | 0 | 1 | [8 empty] | Repr, Accu (kV), Lin
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 1 | [8 empty] | Repr, HVL
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 1 | [8 empty] | Repr, HVL
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | Accu(T)=Accuracy of timer
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | Repr=Reproducibility (kVp, Expos, timer)
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | HVL=Half Value Layer
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | Lin=Linearity
Radio | 80 | 400 | 0.05 | 20 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL
Radio | 80 | 500 | 0.05 | 25 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 50 | 0.05 | 2.5 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 800 | 0.05 | 40 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 250 | 0.05 | 12.5 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 100 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 100 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 1 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 120 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 120 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 1 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 140 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 50 | 100 | 0.1 | 10 | 35-60 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 70 | 100 | 0.1 | 10 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 1 | [8 empty] | Repr, Accu (kV)
Radio | 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
Radio | 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
Radio | 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
Radio | 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
Radio | 80 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 110 | 0.1 | 11 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 250 | 0.1 | 25 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 50 | 0.1 | 5 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 80 | 160 | 0.1 | 16 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
Radio | 110 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 130 | 100 | 0.1 | 10 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
Radio | 80 | 200 | 0.01 | 2 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
Radio | 80 | 200 | 0.02 | 4 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
Radio | 80 | 200 | 0.04 | 8 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
Radio | 80 | 200 | 0.1 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
Radio | 80 | 200 | 0.25 | 50 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
Radio | 80 | 200 | 0.4 | 80 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
Leeds N3 |  | 65 | 70 | 75 | 80 | Nominal |  | TO.10 kV Adj | 65 | 70 | 75 | 80 |  | TO.10 | Diameter (mm) | sqrt(Area) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | 0.167 | 0.188 | 0.167 | 0.149 | 0.135 | 0.16 |  | 0.012 | 0.0119 | 0.011 | 0.0103 | 0.0097 |  | A | 11.1 | 9.8371 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
2 | 0.148 | 0.167 | 0.148 | 0.132 | 0.12 | 0.145 |  | 0.016 | 0.0168 | 0.0155 | 0.0145 | 0.0138 |  | B | 7.9 | 7.0012 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
3 | 0.128 | 0.145 | 0.128 | 0.114 | 0.104 | 0.123 |  | 0.023 | 0.0232 | 0.0215 | 0.0201 | 0.019 |  | C | 5.6 | 4.9629 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
4 | 0.109 | 0.124 | 0.109 | 0.097 | 0.088 | 0.108 |  | 0.032 | 0.0347 | 0.0322 | 0.0301 | 0.0285 |  | D | 4 | 3.5449 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
5 | 0.0876 | 0.0993 | 0.0876 | 0.078 | 0.0705 | 0.086 |  | 0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 |  | E | 2.8 | 2.4814 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
6 | 0.0749 | 0.0847 | 0.0749 | 0.067 | 0.0609 | 0.076 |  | 0.066 | 0.0766 | 0.0674 | 0.0599 | 0.0542 |  | F | 2 | 1.7725 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
7 | 0.0674 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | 0.066 |  | 0.086 | 0.0993 | 0.0876 | 0.078 | 0.0705 |  | G | 1.4 | 1.2407 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
8 | 0.0525 | 0.0593 | 0.0525 | 0.047 | 0.0427 | 0.055 |  | 0.123 | 0.145 | 0.128 | 0.114 | 0.104 |  | H | 1 | 0.88623 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
9 | 0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | 0.045 |  | 0.16 | 0.188 | 0.167 | 0.149 | 0.135 |  | J | 0.7 | 0.62036 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
10 | 0.0371 | 0.04 | 0.0371 | 0.0347 | 0.0329 | 0.039 |  | 0.23 | 0.267 | 0.238 | 0.214 | 0.194 |  | K | 0.5 | 0.44311 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
11 | 0.0322 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | 0.033 |  | 0.35 | 0.399 | 0.36 | 0.327 | 0.299 |  | L | 0.35 | 0.31018 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
12 | 0.0246 | 0.0269 | 0.0246 | 0.0233 | 0.022 | 0.027 |  | 0.5 | 0.618 | 0.573 | 0.533 | 0.497 |  | M | 0.25 | 0.22156 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
13 | 0.0215 | 0.0232 | 0.0215 | 0.0201 | 0.019 | 0.023 |  | 0.66 | 0.769 | 0.726 | 0.686 | 0.649
14 | 0.0172 | 0.0186 | 0.0172 | 0.0161 | 0.0153 | 0.018 |  | 0.93 | 0.969 | 0.954 | 0.936 | 0.915
15 | 0.0155 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | 0.016
16 | 0.013 | 0.014 | 0.013 | 0.0122 | 0.0115 | 0.013
17 | 0.011 | 0.0119 | 0.011 | 0.0103 | 0.0097 | 0.011
18 | 0.0086 | 0.0093 | 0.0086 | 0.0081 | 0.0077 | 0.009
19 | 0.0066 | 0.0071 | 0.0066 | 0.0062 | 0.0059 | 0.007
TO.10 Values | 0 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 1: 8.5; 2: 9
A | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 1: 0.01325; 2: 0.011
B | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 1: 0.01325; 2: 0.011
C | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 1: 0.01325; 2: 0.011
D | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 1: 0.0185; 2: 0.0155
E | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 1: 0.0185; 2: 0.0155
F | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 1: 0.0185; 2: 0.0155
G | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 1: 0.0386; 2: 0.0322
H | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 1: 0.0386; 2: 0.0322
J | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 1: 0.0386; 2: 0.0322
K | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 1: 0.0775; 2: 0.0674
L | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 1: 0.0775; 2: 0.0674
M | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 1: 0.0775; 2: 0.0674
TO.16 | Diameter (mm) | sqrt(Area) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 11.1 | 9.8371 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028 | 0.002 | 0.0015
B | 8 | 7.0898 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028 | 0.002 | 0.0015
C | 5.6 | 4.9629 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028 | 0.002 | 0.0015
D | 4 | 3.5449 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028 | 0.002
E | 2.8 | 2.4814 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028 | 0.002
F | 2 | 1.7725 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028 | 0.002
G | 1.4 | 1.2407 | 0.115 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028
H | 1 | 0.88623 | 0.115 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028
J | 0.7 | 0.62036 | 0.115 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101 | 0.0081 | 0.0062 | 0.004 | 0.0028
K | 0.5 | 0.44311 | 0.465 | 0.33 | 0.215 | 0.15 | 0.115 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145 | 0.0101
L | 0.35 | 0.31018 | 0.697 | 0.0465 | 0.33 | 0.215 | 0.15 | 0.115 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201 | 0.0145
M | 0.25 | 0.22156 | 0.945 | 0.697 | 0.465 | 0.33 | 0.215 | 0.15 | 0.115 | 0.0782 | 0.0601 | 0.04 | 0.0301 | 0.0201
METHOD TO SET THE LINE LENGTH FOR THE COMMENTS: (1) Set the maximum for line 1. (2) Vary the long word factor until the three lines are fairly evenly distributed.
[5 empty] | Column |  | Note: Maximum should be 407 for the parameters specified.
[5 empty] | AB | Rows |  | I wanted every body to know that now is the time for all good women to come to the aid of their country, and if you don't belive that, then know that the sly brown fox jumped quickly over the lazy dog. I wanted every body to know that now is the time for all good men to come to the aid of their country, and if you don't belive that, the know that the sly brown fox jumped quickly over the lazy dog. The end.
|  |  |  | Long word factor( variable): | 6 |  |  | Number of characters: | 409
|  |  |  | Max line1 (variable): | 137 | 1505 | 135 | I wanted every body to know that now is the time for all good women to come to the aid of their country, and if you don't belive that,
|  |  |  | Length for two lines (lower)-Calc: | 268 | 1506 | 133 | then know that the sly brown fox jumped quickly over the lazy dog. I wanted every body to know that now is the time for all good men
|  |  |  | 1st line split for two lines only-Calc: | 134 | 1507 | 141 | to come to the aid of their country, and if you don't belive that, the know that the sly brown fox jumped quickly over the lazy dog. The end.
|  |  |  | 1st line split (max)-Calc: | 135.5 | 1508 | 409
|  |  |  | Length for two lines (max)-Calc: | 272.5 | 1509
|  |  |  | Count for lines 1 & 2 when 3rd present-Calc: | 266.5 | 1510
Note on X2: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on AC2: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on I92: Eugene Mah: Nominal contrast values as a function of kVp
Note on O92: Eugene Mah: TO.10 nominal contrast values
Note on A113: Eugene Mah: TO.10 contrast values at the kVp specified in Fluoro!V319
Note on A127: Eugene Mah: TO.16 nominal contrast values at 70 kV and 1.5 mm Cu filtration

## DataPage
Sheet | RAD
SurveyID
LFS | 0
SFS | 0
Collimation table | 5 | 5 | 5 | 5
LFS Output
SFS Output
Gen Test | 60 | 400 | 0.05 | 20 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 60 | 400 | 0.05 | 20 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 20 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 500 | 0.05 | 25 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 50 | 0.05 | 2.5 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 800 | 0.05 | 40 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 250 | 0.05 | 12.5 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 400 | 0.05 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0
| 120 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 120 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0
| 140 | 400 | 0.05 | 20 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 50 | 100 | 0.1 | 10 | 35-60 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 70 | 100 | 0.1 | 10 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 110 | 0.1 | 11 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 250 | 0.1 | 25 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 50 | 0.1 | 5 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 160 | 0.1 | 16 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 110 | 100 | 0.1 | 10 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 130 | 100 | 0.1 | 10 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.01 | 2 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.02 | 4 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.04 | 8 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.1 | 20 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.25 | 50 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.4 | 80 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
HVL | #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!