## Summary
Medical Physicist CT Survey Report

This report summarizes the results of tests performed in accordance with the American College of
Radiology CT QC Manual.

Facility Name			Unit ID
Address 1			Manufacturer
Address 2			Model
City, State, ZIP			Serial Number
			Date of Manufacture
CTAP # (if applicable)
			Survey Date				30-Dec-99
Medical Physicist	Eugene Mah		Report Date
Signature

Medical Physicist Tests		Pass/Fail	Technologist QC Evaluation					Pass/Fail/NA

Review of CT Protocols			Water CT Number and SD (Daily)
Scout Prescription Accuracy			Artifact Evaluation (Daily)
Alignment Light Accuracy			Wet Laser QC (Weekly)
Image Thickness			Visual Checklist (Monthly)
Table Travel Accuracy			Dry Laser QC (Monthly)
Radiation Beam Width			Acquisition Display QC (Monthly)
Low-Contrast Performance
Spatial Resolution
CT Number Accuracy
Artifact Evaluation
Dosimetry
CT Number Uniformity
Acquisition Display Calibration

Comments

## Sheet1
1					Revision 1.0-20140806					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					CT System Compliance Inspection					CT System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1
8					System Information					Previous Date:					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Accession Number:					Survey ID:				Accession Number:					Survey ID:
14					Room Number:					Room Number:
15			CT Scanner					Site Number:
16				Manufacturer:					Manufacture Date:					CT Scanner					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18					Model:				Serial Number:					Accesssion Number:
19			X-Ray Tube					Focal Spot Sizes (mm)
20				Tube Designation/Use:					Large:					X-Ray Tube					Focal Spot Sizes (mm)					X-Ray Generator
21				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Manufacturer:
22			Insert					Micro:					Manufacture Date:				Small:					Model:
23				Manufacturer:					Insert					Micro:					Manufacture Date:
24				Model:					Filtration					Manufacturer:					Serial Number:
25				Serial Number:					Inherent:					Model:				Filtration
26			Housing					Added:					Serial Number:				Inherent:					X-Ray Tube 1
27				Manufacturer:					Housing					Added:					Tube Designation/Use:
28				Model:					Manufacturer:					Manufacture Date:
29				Serial Number:					Model:					Manufacturer:
30					Serial Number:					Model:
31					Serial Number:
32		Rule Number			Labels, Notices, Postings					Compliance					Manufacturer:
33		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Inspection Results					Model:
34		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
35				Pregnancy Warning sign is posted					TBD					Large:
36		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Labels, Notices, Postings					Small:
37				Operator manuals are available.					TBD				DHEC Registration sticker is present, clearly visible and legible					Micro:
38				Monthly radiation monitoring reports are posted.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Inherent:
39					Pregnancy Warning sign is posted					Added:
40					Radiation Safety					Radiation warning label posted on the generator control panel					Measurement Protocols
41				Radiation warning light at room entrance is functional with selected tube					TBD				Operator manuals are available.					Adult Abdomen
42				Lead aprons available.					TBD				Monthly radiation monitoring reports are posted.					kVp
43				A properly designed and installed apron rack is present.					TBD					mA
44				Documentation of annual protective apparel integrity inspection is available.					TBD					Radiation Safety					mAs
45				Patient restraint devices available.					TBD				Radiation warning light at room entrance is functional with selected tube					Eff mAs
46					Lead aprons available.					Time/rotation (s)
47				Reconstructed images are free of artifacts					TBD				A properly designed and installed apron rack is present.					Scan FOV (cm)
48				Acquisition station display performance is adequate					TBD				Documentation of annual protective apparel integrity inspection is available.					Display FOV (cm)
49				Daily QC performed by technologist					TBD				Patient restraint devices available.					Reconstruction Algorithm
50				Interlock systems function to prevent scanning while door is open					TBD					Axial/Helical
51		DHEC RHB 4.11.1.1		Tomographic plane is indicated by a light source					TBD				Reconstructed images are free of artifacts					Z-Axis collimation
52		DHEC RHB 4.11.1.3.1		Visual indication of x-ray production is present and functional					TBD				Acquisition station display performance is adequate					# of channels
53		DHEC RHB 4.11.1.6.1		Positioning with the alignment light is accurate to within 2mm					TBD				Daily QC performed by technologist					Table Increment/speed
54		DHEC RHB 4.11.1.6.3		Table travel is accurate to within 1mm					TBD				Interlock systems function to prevent scanning while door is open					Pitch
55		DHEC RHB 4.11.2.5.1		Technologist has a clear view of the patient during the scan					TBD				Tomographic plane is indicated by a light source					Recon scan width (mm)
56				Patient communication system is operational					TBD				Visual indication of x-ray production is present and functional					Recon scan interval (mm)
57					Positioning with the alignment light is accurate to within 2mm					Dose reduction techniques
58					Radiation Dose Measurements					Table travel is accurate to within 1mm					Indiated CTDIvol (mGy)
59				Routine head protocol CTDI is below ACR reference level of 75 mGy					NA				Technologist has a clear view of the patient during the scan					Adult Chest
60				Routine body protocol CTDI is below ACR reference level of 25 mGy					NA				Patient communication system is operational					kVp
61				Pediatric body protocol CTDI is below ACR reference level of 15 mGy					NA					mA
62				Pediatric head protocol CTDI is below ACR reference level of 35 mGy					NA					Radiation Dose Measurements					mAs
63					3	Routine head protocol CTDI is below ACR reference level of 75 mGy					Eff mAs
64					3	Routine body protocol CTDI is below ACR reference level of 25 mGy					Time/rotation (s)
65		Date:					Inspector:	Eugene Mah		3	Pediatric body protocol CTDI is below ACR reference level of 15 mGy					Scan FOV (cm)
66		Room Number					Survey ID:			3	Pediatric head protocol CTDI is below ACR reference level of 35 mGy					Display FOV (cm)
1					Medical University of South Carolina					Reconstruction Algorithm
2					Inspection Results					CT System Compliance Inspection					Axial/Helical
3					Measurement Protocols					Measurement Protocols					Z-Axis collimation
4					Adult			Pediatric					Adult			Pediatric					# of channels
5					Abdomen	Chest	Head	Abdomen	Head					Abdomen	Chest	Head	Abdomen	Head					Table Increment/speed
6				kVp					kVp					Pitch
7				mA					mA					Recon scan width (mm)
8				mAs					mAs					Recon scan interval (mm)
9				Eff mAs					Eff mAs					Dose reduction techniques
10				Time/rotation (s)					Time/rotation (s)					Indiated CTDIvol (mGy)
11				Scan FOV (cm)					Scan FOV (cm)					Adult Head
12				Display FOV (cm)					Display FOV (cm)					kVp
13				Reconstruction Algorithm					Reconstruction Algorithm					mA
14				Axial/Helical					Axial/Helical					mAs
15				Z-Axis collimation					Z-Axis collimation					Eff mAs
16				# of channels					# of channels					Time/rotation (s)
17				Table Increment/speed					Table Increment/speed					Scan FOV (cm)
18				Pitch					Pitch					Display FOV (cm)
19				Recon scan width (mm)					Recon scan width (mm)					Reconstruction Algorithm
20				Recon scan interval (mm)					Recon scan interval (mm)					Axial/Helical
21				Dose reduction techniques					Dose reduction techniques					Z-Axis collimation
22				Indiated CTDIvol (mGy)					Indiated CTDIvol (mGy)					# of channels
23					Adult			Pediatric					Table Increment/speed
24		CTDI Measurements					Abdomen	Chest	Head	Abdomen	Head					Pitch
25					Adult			Pediatric					kVp					Recon scan width (mm)
26					Abdomen	Chest	Head	Abdomen	Head	Criteria:					mA					Recon scan interval (mm)
27				Indicated CTDIvol (mGy):					CTDIvol is less than ACR Pass/Fail dose					mAs					Dose reduction techniques
28			CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					Reference	Pass/Fail			Eff mAs					Indiated CTDIvol (mGy)
29					(mGy)	(mGy)			Time/rotation (s)					Ped Abdomen
30					Adult Abdomen	25	30			Scan FOV (cm)					kVp
31				Average CTDIvol:					Adult Head:	75	80			Display FOV (cm)					mA
32				Phantom size indicated (cm):					Ped Abd (16cm)	15	20			Reconstruction Algorithm					mAs
33				Phantom size used (cm):					Ped Abd (32cm)	7.5	10			Axial/Helical					Eff mAs
34				Variation from previous year:					Pediatric Head:	35	40			Z-Axis collimation					Time/rotation (s)
35				Variation from indicated CTDI:					# of channels					Scan FOV (cm)
36				Acceptable:					Table Increment/speed					Display FOV (cm)
37				CTDIvol/mAs:					Pitch					Reconstruction Algorithm
38				Criteria:	CTDI is within 5% of previous year.					Recon scan width (mm)					Axial/Helical
39					Measured CTDI is within 20% of indicated CTDI					Recon scan interval (mm)					Z-Axis collimation
40					Dose reduction techniques					# of channels
41		Beam Profile					Indiated CTDIvol (mGy)					Table Increment/speed
42			Channels (N)					Pitch
43			Detector width (mm)					Recon scan width (mm)
44			Table velocity (mm/s)					Inspection Results					Recon scan interval (mm)
45			Beam width (mm)					CTDI Measurements					Dose reduction techniques
46			Measured beam width (mm)					Adult			Pediatric			Criteria:				Indiated CTDIvol (mGy)
47					Abdomen	Chest	Head	Abdomen	Head		CTDIvol is less than ACR Pass/Fail dose				Ped Head
48					Indicated CTDIvol (mGy)					Reference	Pass/Fail		kVp
49					CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					(mGy)	(mGy)		mA
50			Acceptable:					Adult Abdomen	25	30		mAs
51				Criteria:	Beam profile is within 3mm or 30% of specified value					Adult Head:	75	80		Eff mAs
52					Average CTDIvol:					Ped Abd (16cm)	15	20		Time/rotation (s)
53					Phantom size indicated (cm)					Ped Abd (32cm)	7.5	10		Scan FOV (cm)
54					Phantom size used (cm)					Pediatric Head:	35	40		Display FOV (cm)
55		Artifact Evaluation					Previous year:					Reconstruction Algorithm
56			Water phantom is free from rings, streaks, lines, cupping artifacts					Variation from prev year:					Axial/Helical
57					Variation from indicated:					Z-Axis collimation
58		Acquisition Console Monitor					Acceptable:					# of channels
59			SMPTE 5%/95% visible			Lmin (cd/m^2)					CTDIvol/100 mAs:					Table Increment/speed
60			Display free of artifacts			Lmax (cd/m^2)					Prev CTDIvol/100 mAs:					Pitch
61			Display uniformity		#DIV/0!	% Non-uniformity					Criteria:	CTDI is within 5% of previous year.					Recon scan width (mm)
62					Measured CTDI is within 20% of indicated CTDI					Recon scan interval (mm)
63					Beam Profile					Dose reduction techniques
64					Channels (N)					Indiated CTDIvol (mGy)
65		Date:					Inspector:	Eugene Mah			Detector width (mm)					Criteria:	Beam profile is within 3mm or 30% of specified value					CTDI Measurements
66		Room Number					Survey ID:				Table velocity (mm/s)					Adult Abdomen
1					Medical University of South Carolina			Beam width (mm)					Adult Chest
2					Inspection Results					CT System Compliance Inspection			Measured beam width (mm)					Adult Head
3		Alignment Accuracy					Ped Abdomen
4			Landmark position:		mm		Slice thickness:		mm					Ped Head
5					Average (mm):					CTDIvol/100 mAs
6					Foot	Head				Scout prescription accuracy:					Acceptable:					Adult Abdomen
7				Location of alignment slice:			mm			Central wire visible:					Adult Chest
8				All markers visible					Table movement accuracy:				Artifact Evaluation					Adult Head
9					Alignment light accuracy:					Water phantom is free from rings, streaks, lines, cupping artifacts					Ped Abdomen
10					Ped Head
11		CT Number Accuracy
12			kVp	80	100	110	120	130	140	150					Inspection Results					Comment Page 1
13			mAs					Alignment Accuracy
14			Polyethylene					Landmark position:		mm	Slice thickness		mm
15			Water
16			Water Std Dev					Foot	Head				Enter 1 for YES, 2 for NO
17			Acrylic					Location of alignment slice:			mm		Scout prescription accuracy:
18			Bone					All markers visible:					Central wire visible:
19			Air					Table movement accuracy:
20					Acceptable:					Alignment light accuracy:
21				Critiera:	Water HU is 0 +/- 5, air HU is within acceptable range at all kVp					CT Number Accuracy
22					CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp					kVp	80	100	110	120	130	140	150		Acceptable HU
23					mAs					Low	High
24		CT Slice Thickness Accuracy					Polyethylene					Polyethylene	-107	-84
25				Slice thickness (mm):					Water					Water	-7	7
26				Top (mm):					Water Std Dev					Acrylic	110	135
27				Bottom (mm):					Acrylic					Bone	850	970
28				Average (mm):					Bone					Air	-1005	-970
29				Water HU:					Air
30				Water HU Std Dev:					Acceptable:		Enter 1 for YES, 2 for NO
31				Acceptable:					Critiera:	Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
32					Criteria:	Slice thickness is within 1.5mm of the set thickness					CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
33					CT number for water is 0 +/- 5
34					CT Slice Thickness Accuracy
35		Low Contrast Performance					Slice thickness (mm):
36				Adult		Pediatric					Top (mm):
37				Abdomen	Head	Abdomen	Head					Bottom (mm):
38			Inside					6mm rods visible:					Average (mm):
39			Adjacent					Acceptable:					Water HU:
40			Adjacent StDev					Water HU Std Dev:
41			Contrast					Acceptable:
42			CNR					Criteria:	Slice thickness is within 1.5mm of the set thickness
43					Criteria:	All four 6mm rods are visible					CT number for water is 0 +/- 5
44					CNR > 1.0 for adult abdomen and head protocols
45					CNR > 0.4 for pediatric abdomen protocol					Low Contrast Performance
46					CNR > 1.0 for pediatric head protocol					Adult		Pediatric
47		Uniformity					Abdomen	Head	Abdomen	Head			Enter 1 for YES, 2 for NO, 3 for Not Applicable
48			Mean	St Dev	Diff (Center)	Diff (Mean)	Acceptable					Inside					6mm rods visible:		Acceptable:
49		Center:					Adjacent					Criteria:	All four 6mm rods are visible
50		Top:					Adj StDev					CNR > 1.0 for adult abdomen and head protocols
51		Bottom:					Contrast					CNR > 0.4 for pediatric abdomen protocol
52		Right:					CNR					CNR > 1.0 for pediatric head protocol
53		Left:
54			Criteria:	CT number of all ROIs is within +/- 5 HU of the central region					Uniformity
55			Uniformity image is free from artifacts					Mean	St Dev	Diff (Center)	Diff (Mean)	Acceptable
56					Center:
57		High Contrast Resolution					Top:
58				lp/cm	Acceptable					Bottom:
59			Abdomen				Critiera:	6 lp/cm for abdomen					Right:
60			Hi-res Chest					8 lp/cm for high resolution					Left:
61					Uniformity image is free from artifacts			Criteria:	CT number of all ROIs is within +/- 5 HU of the central region
62					High Contrast Resolution
63					lp/cm	Acceptable
64					Abdomen				Critiera:	6 lp/cm for abdomen
65		Date:					Inspector:	Eugene Mah		Hi-res Chest					8 lp/cm for high resolution
66		Room Number					Survey ID:
1					Medical University of South Carolina		Acquisition Console Monitor
2					Comments Page 1					CT System Compliance Inspection					Display uniformity
3					SMPTE 5%/95% visible			Top left:
4					Display free of artifacts			Top right:
5					Display uniformity			Center:
6					Lmin (cd/m^2)			Bottom left:
7					Lmax (cd/m^2)			Bottom right:
8					Critiera:	Lmin <= 1.2 cd/m^2, Lmax >= 90 cd/m^2			Average:	#DIV/0!
9					Display uniformity is +/-15%			SD:	#DIV/0!
10					% Non-uniformity:	#DIV/0!
11
12					Comments Page 1
13					Previous Year Comments
14					Comments:
15					Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
16					Additional Comments:
17					Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
18					Additional Comments:
19					Do not exceed 390 chars		0			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
20					Additional Comments:
21					Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
22					Additional Comments:
23					Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
24					Additional Comments:
25					Do not exceed 390 chars		0
26					Additional Comments:
27					Do not exceed 390 chars		0
28					Additional Comments:
29					Do not exceed 390 chars		0
30					Additional Comments:
31					Do not exceed 390 chars		0
32					Additional Comments:
33					Do not exceed 390 chars		0
34					Additional Comments:
35					Do not exceed 390 chars		0
36					Additional Comments:
37					Do not exceed 390 chars		0
38					Additional Comments:
39					Do not exceed 390 chars		0
40					Additional Comments:
41					Do not exceed 390 chars		0
42					Additional Comments:
43					Do not exceed 390 chars		0
44					Additional Comments:
45					Do not exceed 390 chars		0
46					Additional Comments:
47					Do not exceed 390 chars		0
48					Additional Comments:
49					Do not exceed 390 chars		0
50					Additional Comments:
51					Do not exceed 390 chars		0
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:

## Sheet2
CT Beam Profile – Siemens
NxT	Definition Flash			NxT	Definition AS			NxT	Sensation 64		NxT	Sensation 16		NxT	Emotion 16
	Z-sharp	Nominal (mm)	Tolerance)		Z-sharp	Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)
2x1	Off	3.5	1.0	2x1	Off	4.7	1.0	30x0.6	22.5	1.5	16x1.5	30.5	1.0	4x0.6	4.9	1.0
1x5	Off	6.0	1.0	1x5	Off	6.6	1.0	24x1.2	33.5	1.5	16x0.75	18.5	1.0	12x0.6	12.6	1.0
10x0.6	On	10.5	1.5	10x0.6	On	11.2	1.0	12x1.2	19.5	1.5	12x1.5	24.5	1.0	16x0.6	14.5	1.0
8x0.6	On	11.5	1.5	8x0.6	Off	8.5	1.0	1x10	11.5	1.5	12x0.75	15.5	1.0	2x5	12.6	1.0
1x10	Off	11.1	1.5	1x10	Off	11.6	1.0	1x5	6.5	1.5	2x5	12.5	1.0	12x1.2	18.5	1.0
20x0.6	On	18.8	1.5	8x0.6	On	11.5	1.0	2x1	4.0	1.0	2x1	3.5	1.0	2x8	18.5	1.0
12x1.2	Off	19.4	1.5	6x1.2	Off	11.8	1.0	6x0.6	7.0	1.0	2x0.6	3.0	1.0	16x1.2	24.2	1.0
32x0.6	Off	24.4	2.0	12x1.2	Off	19.8	1.0
32x0.6	On	27.4	2.0	60x0.6	Off	41.0	1.5
32x1.2	Off	44.0	2.0	64x0.6	On	45.5	1.5
64x0.6	On	45.5	2.0	32x1.2	Off	43.5	1.5
				30x0.6	Off	23.5	1.0
				32x0.6	On	28.0	1.0
				20x0.6	Off	18.0	1.0
				16x1.2	Off	25.0	1.0
				20x0.6	On	20.0	1.0