## Sheet1
1 | [11 empty] | Revision 1.3-20190605 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [22 empty] | All:
5 | [5 empty] | Bone Mineral Density Compliance Inspection | [10 empty] | Bone Mineral Density Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [5 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility:
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department:
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division:
13 |  |  | Survey ID: | [10 empty] | Survey ID: | [8 empty] | Survey ID:
14 | [13 empty] | Room Number: | [8 empty] | Room Number:
15 |  | X-Ray Generator | [20 empty] | Site Number:
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [9 empty] | Authorized Use:
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation:
18 |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number:
19 | [13 empty] | Max kVp: |  |  |  | Max mA:
20 |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes (mm) | [14 empty] | X-Ray Generator
21 |  |  | Tube Designation/Use: |  |  |  |  | Large: |  |  |  |  | X-Ray Tube 1 |  |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Manufacturer:
22 |  |  | Manufacture Date: |  |  |  |  | Small: | [5 empty] | Tube Designation/Use: |  |  |  | Large: |  |  |  |  | Model:
23 |  | Insert | [5 empty] | Micro: | [5 empty] | Manufacture Date: |  |  |  | Small: |  |  |  |  | Max kVp:
24 |  |  | Manufacturer: | [9 empty] | Insert |  |  |  |  | Micro: |  |  |  |  | Manufacture Date:
25 |  |  | Model: |  |  |  |  | Filtration | [5 empty] | Manufacturer: | [8 empty] | Serial Number:
26 |  |  | Serial Number: |  |  |  |  | Inherent: | [5 empty] | Model: |  |  |  | Filtration |  |  |  |  | Max mA:
27 |  | Housing | [5 empty] | Added: | [5 empty] | Serial Number: |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
28 |  |  | Manufacturer: | [9 empty] | Housing |  |  |  |  | Added: |  |  |  |  | Tube Designation/Use:
29 |  |  | Model: | [10 empty] | Manufacturer: | [8 empty] | Manufacture Date:
30 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | Manufacturer:
31 | [13 empty] | Serial Number: | [8 empty] | Model:
32 | [22 empty] | Serial Number:
33 | [5 empty] | Inspection Results |  |  |  |  | Bone Mineral Density Compliance Inspection | [11 empty] | Manufacturer:
34 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance | [6 empty] | Inspection Results | [5 empty] | Model:
35 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Serial Number:
36 | DHEC RHB 10.2.1 |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | TBD | [6 empty] | Labels, Notices, Postings | [5 empty] | Large:
37 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD |  |  | DHEC Registration sticker is present, clearly visible and legible | [9 empty] | Small:
38 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [9 empty] | Micro:
39 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | Radiation warning label posted on the generator control panel | [9 empty] | Inherent:
40 | [12 empty] | Operator manuals are available. | [9 empty] | Added:
41 |  |  |  | Radiation Safety | [8 empty] | Monthly radiation monitoring reports are posted.
42 |  |  | Operator has full visibility of the exposure factors and patient | [6 empty] | TBD | [12 empty] | Equipment: |  | Change | Piranha CB2-17090320
43 |  |  | Operator position is at least 1m away from patient and unit | [6 empty] | TBD | [6 empty] | Radiation Safety | [5 empty] | Detector: |  | Change | R100B
44 | [12 empty] | Operator has full visibility of the exposure factors and patient
45 |  |  |  | Physical Inspection | [8 empty] | Operator position is at least 1m away from patient and unit
46 |  |  | System clean and free of debris | [6 empty] | TBD
47 |  |  | All cables free from excessive wear or restraints | [6 empty] | TBD | [6 empty] | Physical Inspection
48 |  |  | System free of missing or damaged components | [6 empty] | TBD |  |  | System clean and free of debris
49 |  |  | All other movements and locks functional | [6 empty] | TBD |  |  | All cables free from excessive wear or restraints
50 |  |  | Unit installed as shown on shielding plan | [6 empty] | NA |  |  | System free of missing or damaged components
51 | [12 empty] | All other movements and locks functional
52 | [11 empty] | 3 | Unit installed as shown on shielding plan
53
54 | [11 empty] | Scan Techniques | [6 empty] | Measurement Locations
55 | [15 empty] | Field size
56 | [11 empty] | Mode | kV | mA | Time (s) | Trans (cm) | Long (cm) |  | Location | Distance (cm)
57 | [11 empty] | High Def | 140/100 | 2.5 | 163 | 11.43 | 20.32 |  | Tableside
58 | [11 empty] | Array | 140/100 | 2.5 | 82 | 11.43 | 20.32 |  | Operator
59 | Date: | [8 empty] | Inspector: | Eugene Mah | Fast Array | 140/100 | 2.5 | 41 | 11.43 | 20.32 |  | Door
60 | Room Number | [8 empty] | Survey ID: |  | Turbo | 140/100 | 2.5 | 15 | 11.43 | 20.32
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Inspection Results |  |  |  |  | Bone Mineral Density Compliance Inspection | Background
3 | Scan Techniques | [10 empty] | Time (s): | 82 | Reading (nGy): |  |  |  |  | Avg Bkg: (nGy):
4 | [5 empty] | Field Size |  |  | Measurement Locations
5 |  | Mode | kV | mA | Time (s) | Trans (cm) | Long (cm) |  | Location | Distance (cm) |  | Scatter Measurements | [5 empty] | Average
6 |  | High Def | 140/100 | 2.5 | 163 | 11.43 | 20.32 |  | Tableside |  |  | High Def | Location | Distance (cm) | nGy/scan |  |  | nGy/scan | uGy/h
7 |  | Array | 140/100 | 2.5 | 82 | 11.43 | 20.32 |  | Operator |  |  |  | Tableside
8 |  | Fast Array | 140/100 | 2.5 | 41 | 11.43 | 20.32 |  | Door |  |  |  | Operator
9 |  | Turbo | 140/100 | 2.5 | 15 | 11.43 | 20.32 | [5 empty] | Door
10
11
12 | [11 empty] | Array | Location | Distance (cm) | nGy/scan |  |  | nGy/scan | uGy/h
13 | [12 empty] | Tableside
14 | [12 empty] | Operator
15 | Scatter Measurements | [5 empty] | Entrance exposure Measurements | [5 empty] | Door
16
17 | High Def | Location | Distance (cm) | nGy/scan | uGy/h |  | Mode | Detector (mGy) | Patient Entrance
18 |  | Tableside | [6 empty] | (mGy) | (mGy/s) |  | Fast Array | Location | Distance (cm) | nGy/scan |  |  | nGy/scan | uGy/h
19 |  | Operator |  |  |  |  | High Def | [5 empty] | Tableside
20 |  | Door |  |  |  |  | Array | [5 empty] | Operator
21 | [6 empty] | Fast Array | [5 empty] | Door
22 | [6 empty] | Turbo
23 | Array | Location | Distance (cm) | nGy/scan | uGy/h
24 |  | Tableside |  |  |  |  | Test equipment: | Piranha CB2-17090320 |  | Detector: | R100B | Turbo | Location | Distance (cm) | nGy/scan |  |  | nGy/scan | uGy/h
25 |  | Operator |  |  |  |  | Calibration date: |  |  | Calibration date: |  |  | Tableside
26 |  | Door |  |  |  |  | Calibration due: |  |  | Calibration due: |  |  | Operator
27 | [12 empty] | Door
28
29 | Fast Array | Location | Distance (cm) | nGy/scan | uGy/h
30 |  | Tableside | [10 empty] | Test equipment: |  |  | Calibration date: |  |  | Calibration due:
31 |  | Operator
32 |  | Door
33 | [11 empty] | Exposure measurements
34 | [11 empty] | 107 | Source – Detector distance (cm) |  |  |  | Test equipment used: | Piranha CB2-17090320 |  | Detector: | R100B
35 | Turbo | Location | Distance (cm) | nGy/scan | uGy/h | [6 empty] | 42.4 | Source – Patient distance (cm) |  |  |  |  | Piranha CB2-17090320 |  |  | R100B
36 |  | Tableside | [14 empty] | Calibration date: |  |  | Calibration date:
37 |  | Operator | [9 empty] | Detector entrance dose measurements |  |  |  |  | Calibration due: |  |  | Calibration due:
38 |  | Door | [13 empty] | Average |  | Pt Entrance
39 | [11 empty] | Mode | mGy/scan |  |  | (mGy) | SD | (mGy) | (mGy/s)
40 | [11 empty] | High Def
41 |  | Test equipment: | [9 empty] | Array
42 |  | Calibration date: | [9 empty] | Fast Array
43 |  | Calibration due: | [9 empty] | Turbo
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59 | Date: | [8 empty] | Inspector: | Eugene Mah
60 | Room Number | [8 empty] | Survey ID: