## QC Test Summary-Siemens
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Siemens
Site Name | [8 empty] | Report Date
Address | [8 empty] | Survey Date | 0
Medical Physicist's Name |  |  | Eugene Mah | [5 empty] | Signature
X-Ray Unit Manufacturer |  |  | Siemens | [5 empty] | Model | Mammomat Revelation
Date of Installation | [8 empty] | Room ID
QC Manual Version # |  |  | VC10 |  |  |  | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  | Model |  | Location |  | QC Manual Version #
|  | Review Workstation* | Agfa |  | Impax |  |  |  | v4.2 09-2015
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System (www.accessdata.fda.gov/cdrh_docs/presentations/pghs/Polic_Guidance_Help_System.htm).
Survey Type |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [9 empty] | Annual Survey
Features |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Image Quality
Largest 5 fibers, 4 speck groups and 4 masses visible*
(*largest 4 fibers, 3 speck groups and 3 masses acceptable if spatial resolution and CNR pass)
| Phantom image scores: |  |  |  | Fibers | 0 | Specks | 0 | Masses | 0
2. Artifact Detection
3. Printer Check (if applicable)
4. SNR, CNR and AEC Repeatability
| Measured values: |  |  | SNR |  | CNR
CV for mAs and entrance air kerma <5%
Max deviation of mean pixel values and SNR within ±15% of mean for measurements
5. Radiation Dose
| Average glandular dose for average breast is <3 mGy (300 mrad) | [8 empty] | #DIV/0! | mGy
6. Spatial Resolution
7. AEC Test
8. Detector Uniformity
9. Mechanical Tests
10. Acquisition Workstation Monitor Check
11. Site Audit/Evaluation of Technologist QC Program
12. Collimation, Dead Space & Compression Paddle Position
13. HVL and Radiation Output
14. Tube Voltage Measurement & Repeatability
15. Average Glandular Dose (DBT)
16. Geometric Accuracy in X and Y Direction and Z-Resolution (DBT)
17. Radiation Field (DBT)
18. System Imaging Quality (DBT)
≥ 4 fibers, ≥ 3 speck groups and ≥ 3 masses must be visible
| Phantom image scores: |  |  |  | Fibers | 0 | Specks | 0 | Masses | 0
19. Artifact Detection (DBT)
20. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on N22: Click in boxes to use drop-down lists
Note on H25: Click in boxes to use drop-down lists
Note on H48: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Siemens
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Siemens, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY |  | PASS/FAIL
1. | Phantom Image Quality |  | Novation & Fusion-Daily; Inspiration-Weekly
2. | Artifact Detection | Weekly
3. | SNR and CNR Measurements | Weekly
4. | Detector Calibration |  | Novation-Weekly; Inspiration & Fusion-Quarterly
5. | Repeat/Reject Analysis | Quarterly
6. | Compression Force | Semi-annually
7. | System Imaging Quality (DBT) | Daily when DBT performed
8. | Printer Check (if applicable) | Daily, when images printed |  | NA
9. | Review Workstation QC-Overall (NA if only hardcopy read) | See FDA guidance
10. | Mobile Unit Quality Control (if applicable) | After every move |  | NA
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for Siemens FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## MQSA Requirements
MEDICAL PHYSICIST'S CHECKLIST MQSA REQUIREMENTS FOR MAMMOGRAPHY EQUIPMENT
Facility Name: | 0
Unit Manufacturer: |  |  | Model:
Serial number: |  |  | Year Mfr:
Medical Physicist: | Eugene Mah |  | Room ID:
Signature: |  |  | Survey Date: | 0
Feature | FDA Rule Section | Requirement | Applies to | Meets FDA Requirements? (if NA, please explain)
Motion of tube-image receptor assembly | 3(i) | The assembly shall be capable of being fixed in any position where it is designed to operate. Once fixed in any such position, it shall not undergo unintended motion. | S-F & FFDM
| 3(ii) | This mechanism shall not fail in the event of power interruption. | S-F & FFDM
Image receptor sizes | 4(i) | Systems using screen-film image receptors shall provide, at a minimum, for operation with image receptors of 18 x 24 cm and 24 x 30 cm. | S-F
| 4(ii) | Systems using screen-film image receptors shall be equipped with moving grids matched to all image receptor sizes provided. | S-F
| 4(iii) | Systems used for magnification procedures shall be capable of operation with the grid removed from between the source and image receptor. | S-F & FFDM
Beam limitation and light fields | 5(i) | All systems shall have beam limiting devices that allow the useful beam to extend to or beyond the chest wall edge of the image receptor. | S-F & FFDM
| 5(ii) | For any mammography system with a light beam that passes through the X-ray beam-limiting device, the light shall provide an average illumination of not less than 160 lux (15 ft-candles) at 100 cm or the maximum source-image receptor distance (SID), whichever is less. | S-F & FFDM (except Fischer)
Magnification | 6(i) | Systems used to perform noninterventional problem-solving procedures shall have radiographic magnification capability available for use by the operator. | S-F & FFDM
| 6(ii) | Systems used for magnification procedures shall provide, at a minimum, at least one magnification value within the range of 1.4 to 2.0. | S-F & FFDM
Focal spot selection | 7(i) | When more than one focal spot is provided, the system shall indicate, prior to exposure, which focal spot is selected. | S-F & FFDM
| 7(ii) | When more than one target material is provided, the system shall indicate, prior to exposure, the preselected target material. | S-F & FFDM
| 7(iii) | When the target material and/or focal spot is selected by a system algorithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure. | S-F & FFDM
Application of compression | 8(i)(A) | Each system shall provide an initial power-driven compression activated by hands-free controls operable from both sides of the patient. | S-F & FFDM
| 8(i)(B) | Each system shall provide fine adjustment compression controls operable from both sides of the patient. | S-F & FFDM
Compression paddle | 8(ii)(A) | Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system. | S-F & FFDM
| 8(ii)(B) | The compression paddle shall be flat and parallel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when compression is applied. | S-F & FFDM (except Fischer)
| 8(ii)(C) | Equipment intended by the manufacturer’s design to not be flat and parallel to the breast support table during compression shall meet the manufacturer’s design specifications and maintenance requirements. | S-F & FFDM
| 8(ii)(D) | The chest wall edge of the compression paddle shall be straight and parallel to the edge of the image receptor. | S-F & FFDM
| 8(ii)(E) | The chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image. | S-F & FFDM
Technique factor selection and display | 9(i) | Manual selection of mAs or at least one of its component parts (mA and/or time) shall be available. | S-F & FFDM
| 9(ii) | The technique factors (kVp and either mA and seconds or mAs) to be used during an exposure shall be indicated before the exposure begins, except when AEC is used, in which case the technique factors that are set prior to the exposure shall be indicated. | S-F & FFDM
| 9(iii) | Following AEC mode use, the system shall indicate the actual kVp and mAs (or mA and time) used during the exposure. | S-F & FFDM
Automatic exposure control | 10(i) | Each screen-film system shall provide an AEC mode that is operable in all combinations of equipment configuration provided, e.g., grid, non-grid; magnification, nonmagnification; and various target-filter combinations. | S-F
| 10(ii) | The positioning or selection of the detector shall permit flexibility in the placement of the detector under the target tissue. The size and the available positions of the detector shall be clearly indicated at the X-ray input surface of the breast compression paddle. The selected position of the detector shall be clearly indicated. | S-F
| 10(iii) | The system shall provide means for the operator to vary the selected optical density from the normal (zero) setting. | S-F
X-ray film* | 11 | The facility shall use X-ray film for mammography that has been designated by the film manufacturer as appropriate for mammography. | S-F
Intensifying screens* | 12 | The facility shall use intensifying screens for mammography that have been designated by the screen manufacturer as appropriate for mammography and shall use film that is matched to the screen’s spectral output as specified by the manufacturer. | S-F
Film processing solutions* | 13 | For processing mammography films, the facility shall use chemical solutions that are capable of developing the films used by the facility in a manner equivalent to the minimum requirements specified by the film manufacturer. | S-F
Lighting* | 14 | The facility shall make special lights for film illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the interpreting physicians. | S-F & FFDM (for hardcopy comparison)
Film masking devices* | 15 | Facilities shall ensure that film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are available to all interpreting physicians interpreting for the facility. | S-F & FFDM (for hardcopy comparison)
* NA is acceptable for new units at existing facilities if these were previously evaluated and have not changed
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA

## Sheet1
1 | [11 empty] | Revision 1.3-20200416 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina | [5 empty] | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
4 | [22 empty] | All: | Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5 | [5 empty] | Mammography System Compliance Inspection | [10 empty] | Mammography System Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10 |  | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [5 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah | kV Set | mAs Set | Add Filt | Target | Filter | kVp | ms | Dose (mGy) | Dose rate (mGy/s) | HVL (mm Al) | Corr kV
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility: |  |  |  | 24 | 50 | 0
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department: |  |  |  | 24 | 50 | 0
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division: |  |  |  | 25 | 50 | 0
13 |  |  | Survey ID: |  |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: |  |  |  |  | Survey ID: |  |  |  | 25 | 50 | 0
14 | [13 empty] | Room Number: |  |  |  | Accession Number: |  |  |  |  | Room Number: |  |  |  | 26 | 50 | 0
15 |  | X-Ray Generator | [20 empty] | Site Number: |  |  |  | 28 | 20 | 0
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [9 empty] | Authorized Use: |  |  |  | 28 | 50 | 0
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation: |  |  |  | 28 | 50 | 0
18 |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number: |  |  |  |  | Number of X-ray Tubes: |  |  |  | 28 | 50 | 0
19 | [13 empty] | Max kVp: |  |  |  | Max mA: |  |  |  |  | Accesssion Number: |  |  |  | 28 | 50 | 0
20 |  | X-Ray Tube 1 | [20 empty] | X-Ray Generator |  |  |  | 28 | 100 | 0
21 |  |  | Tube Designation/Use: |  |  |  |  | Target: |  |  |  |  | X-Ray Tube 1 |  |  |  |  | Target: |  |  |  |  | Manufacturer: |  |  |  | 28 | 320 | 0
22 |  |  | Manufacture Date: | [10 empty] | Tube Designation/Use: | [8 empty] | Model: |  |  |  | 30 | 50 | 0
23 |  | Insert | [5 empty] | Filter: | [5 empty] | Manufacture Date: | [8 empty] | Max kVp: |  |  |  | 32 | 50 | 0
24 |  |  | Manufacturer: | [9 empty] | Insert |  |  |  |  | Filter: |  |  |  |  | Manufacture Date: |  |  |  | 32 | 50 | 0
25 |  |  | Model: | [10 empty] | Manufacturer: | [8 empty] | Serial Number: |  |  |  | 34 | 50 | 0
26 |  |  | Serial Number: |  |  |  | Focal Spot Sizes (mm) | [6 empty] | Model: | [8 empty] | Max mA: |  |  |  | 34 | 50 | 0
27 |  | Housing | [5 empty] | Large: | [5 empty] | Serial Number: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | X-Ray Tube 1 |  |  |  | 28 | 50 | 0
28 |  |  | Manufacturer: |  |  |  |  | Small: |  |  |  |  | Housing |  |  |  |  | Large: |  |  |  |  | Tube Designation/Use: |  |  |  | 28 | 50 | 0
29 |  |  | Model: |  |  |  | Filtration | Inherent: | [5 empty] | Manufacturer: |  |  |  | Small: |  |  |  |  | Manufacture Date: |  |  |  | 30 | 50 | 0
30 |  |  | Serial Number: |  |  |  |  | Added: | [5 empty] | Model: | [8 empty] | Manufacturer: |  |  |  | 30 | 50 | 0
31 | [13 empty] | Serial Number: |  |  |  | Filtration |  |  |  |  | Model: |  |  |  | 32 | 50 | 0
32 | [17 empty] | Inherent: |  |  |  |  | Serial Number: |  |  |  | 32 | 50 | 0
33 | [5 empty] | Technique Chart | [5 empty] | 1 | Is this a tomosynthesis unit? |  |  |  |  | Added: |  |  |  |  | Manufacturer: |  |  |  | 34 | 50 | 0
34 | [11 empty] | 1 | Stereotactic breast biopsy unit or add-on present? | [9 empty] | Model: |  |  |  | 34 | 50 | 0
35 | Compressed | Fatty Breast |  |  | 50% Fatty - |  |  | Dense Breast |  |  |  |  | Is this a new installation? | [9 empty] | Serial Number: |  |  |  | 36 | 50 | 0
36 | Breast |  |  |  | 50% Dense Breast | [17 empty] | Large: |  |  |  | 38 | 50 | 0
37 | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | [6 empty] | Inspection Results | [5 empty] | Small: |  |  |  | 28 | 50 | 0
38 | <3 cm | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Target 1: |  |  |  | 28 | 50 | 0
39 | 3 to 5 cm | [11 empty] | DHEC Registration sticker is present, clearly visible and legible | [9 empty] | Target 2: |  |  |  | 28 | 50 | 0
40 | 5 to 7 cm | [11 empty] | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [9 empty] | Filter 1: |  |  |  | 28 | 50 | 0
41 | >7 cm | [11 empty] | Radiation warning label posted on the generator control panel | [9 empty] | Filter 2: |  |  |  | 30 | 50 | 0
42 | [12 empty] | Operator manuals are available. | [9 empty] | Inherent: |  |  |  | 30 | 50 | 0
43 | [5 empty] | Inspection Results | [6 empty] | Monthly radiation monitoring reports are posted. | [9 empty] | Added |  |  |  | 32 | 50 | 0
44 | Rule Number | [8 empty] | Compliance |  |  | Unit installed as shown on shielding plan | [9 empty] | Technique Chart |  |  |  | 32 | 50 | 0
45 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Mechanical Tests | [5 empty] | Fatty Breast |  |  |  | 34 | 50 | 0
46 | DHEC RHB 10.2.1 |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  |  | Printer passes manufacturer QC tests | [9 empty] | Target/Filter: |  |  |  | 34 | 50 | 0
47 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD |  |  | Repeat analysis performed | [9 empty] | kVp: |  |  |  | 38 | 50 | 0
48 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | Control panel lights | [9 empty] | Density: |  |  |  | 38 | 50 | 0
49 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | System stops above floor | [9 empty] | Target/Filter:
50 |  |  | Unit installed as shown on shielding plan | [6 empty] | TBD |  |  | Motorized movements | [9 empty] | kVp:
51 | [5 empty] | Mammography Unit Assembly Evaluation | [6 empty] | Lifts from lowest to highest position | [9 empty] | Density:
52 |  |  | Free-standing unit is mechanically stable | [6 empty] | TBD |  |  | Foot switches functional | [9 empty] | Target/Filter:
53 |  |  | All moving parts move smoothly, without obstructions to motion | [6 empty] | TBD |  |  | Movements blocked on error | [9 empty] | kVp:
54 |  |  | All locks and detents work properly | [6 empty] | TBD |  |  | Emergency stop button functional | [9 empty] | Density:
55 |  |  | Image receptor holder assembly is free from vibrations | [6 empty] | TBD |  |  | Grid motion | [9 empty] | Target/Filter:
56 |  |  | Image receptor slides smoothly into holder assembly | [6 empty] | TBD |  |  | No sharp edges or cracks on table or compression paddles | [9 empty] | kVp:
57 |  |  | Image receptor is held securely by assembly in any orientation | [6 empty] | TBD |  |  | Radiation shield present | [9 empty] | Density:
58 |  |  | Compressed breast thickness scale accurate to 0.5 cm, reproducible to 2 mm | [6 empty] | TBD |  |  | Manual compression functional | [9 empty] | 50/50 Breast
59 |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [6 empty] | TBD |  |  | Autodecompression functional | [9 empty] | Target/Filter:
60 |  |  | Operator technique control charts are posted | [6 empty] | TBD |  |  | Decompression button on control box functional | [9 empty] | kVp:
61 |  |  | Operator protected during exposure by adequate radiation shielding | [6 empty] | TBD |  |  | Emergency decompression | [9 empty] | Density:
62 |  |  | All indicator lights working properly | [6 empty] | TBD |  |  | Measured phantom thickness is 35-43 mm | [9 empty] | Target/Filter:
63 |  |  | Autodecompression can be overridden to maintain compression | [6 empty] | TBD | [6 empty] | Mammography Unit Assembly Evaluation | [5 empty] | kVp:
64 |  |  | Manual emergency compression release in the event of a power failure | [6 empty] | TBD |  |  | Free-standing unit is mechanically stable | [9 empty] | Density:
65 | [12 empty] | All moving parts move smoothly, without obstructions to motion | [9 empty] | Target/Filter:
66 | [12 empty] | All locks and detents work properly | [9 empty] | kVp:
67 | [12 empty] | Image receptor holder assembly is free from vibrations | [9 empty] | Density:
68 | [12 empty] | Image receptor slides smoothly into holder assembly | [9 empty] | Target/Filter:
69 | [12 empty] | Image receptor is held securely by assembly in any orientation | [9 empty] | kVp:
70 | [12 empty] | Compressed breast thickness scale accurate to +0.5 cm, reproducible to +2 mm | [9 empty] | Density:
71 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [9 empty] | Dense Breast
72 | Room Number | [8 empty] | Survey ID: |  |  | Operator technique control charts are posted | [9 empty] | Target/Filter:
1 | [10 empty] | Medical University of South Carolina |  | Operator protected during exposure by adequate radiation shielding | [9 empty] | kVp:
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | All indicator lights working properly | [9 empty] | Density:
3 | Rule Number |  |  |  |  | MQSA Equipment Standards | [6 empty] | Auto decompression can be overridden to maintain compression (and status displayed) | [9 empty] | Target/Filter:
4 | 900.12(b)(3i) |  | Tube-receptor can be fixed into any operating position and does not move | [6 empty] | TBD |  |  | Manual emergency compression release can be activated in the event of a power failure | [9 empty] | kVp:
5 | 900.12(b)(3ii) |  | Tube-receptor locking mechanism does not fail during power interruption | [6 empty] | TBD | [12 empty] | Density:
6 | 900.12(b)(4i) |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [6 empty] | TBD | [6 empty] | MQSA Equipment Standards | [5 empty] | Target/Filter:
7 | 900.12(b)(4ii) |  | Moving grids are present for all image receptor sizessizes | [6 empty] | TBD |  |  | Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position | [9 empty] | kVp:
8 | 900.12(b)(4iii) |  | Magnification can be done without grid | [6 empty] | TBD |  |  | Tube-receptor assembly locking mechanism does not fail during power interruption | [9 empty] | Density:
9 | 900.12(b)(5i) |  | Beam limiting device allows useful beam to extend beyond chest wall edge | [6 empty] | TBD |  |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [9 empty] | Target/Filter:
10 | 900.12(b)(5ii) |  | Average illumination of the light field is at least 160 lux | [6 empty] | YES |  |  | Moving grids are present for all image receptor sizes | [9 empty] | kVp:
11 | 900.12(b)(6i) |  | Magnification mode available for non-interventional problem solving | [6 empty] | TBD |  |  | Magnification can be done without grid | [9 empty] | Density:
12 | 900.12(b)(6ii) |  | System provides magnification factor between 1.4 – 2.0 | [6 empty] | TBD |  |  | Beam limiting device allows useful beam to extend beyond chest wall edge of receptor
13 | 900.12(b)(7i) |  | Prior indication of selected focal spot is provided | [6 empty] | TBD |  | 1 | Average illumination of the light field is at least 160 lux | [9 empty] | MGD: |  |  | #DIV/0!
14 | 900.12(b)(7i) |  | Prior indication of selected target is provided | [6 empty] | TBD |  |  | Magnification mode is available for non-interventional problem solving procedures | [9 empty] | Combo 2D MGD:
15 | 900.12(b)(7iii) |  | Post exposure display of target/focal spot is provided | [6 empty] | TBD |  |  | System provides magnification factor between 1.4 – 2.0 | [9 empty] | Combo Tomo MGD:
16 | 900.12(b)(8)(iB) |  | Initial power driven compression can be activated hands-free from both sides | [6 empty] | TBD |  |  | Prior indication of selected focal spot is provided | [9 empty] | / mGy/mAs:
17 |  |  | Fine adjustment of compression can be activated from both sides | [6 empty] | TBD |  |  | Prior indication of selected target is provided | [9 empty] | / mGy/s:
18 | 900.12(b)(8)(iiA) |  | Compression paddles match the sizes of all image receptors | [6 empty] | TBD |  |  | Post exposure display of target/focal spot is provided | [9 empty] | / mGy/mAs:
19 | 900.12(b)(8)(iiB) |  | Compression paddles do not flex by more than 1 cm | [6 empty] | TBD |  |  | Initial power driven compression can be activated hands-free from both sides | [9 empty] | / mGy/s:
20 | 900.12(b)(8)(iiC) |  | Compression paddles are flat and parallel to the breast support table | [6 empty] | TBD |  |  | Fine adjustment of compression can be activated from both sides | [9 empty] | Image Quality (Printer)
21 | 900.12(b)(8)(iiD) |  | Chest wall edge of the compression paddle is straight and parallel to receptor | [6 empty] | TBD |  |  | Compression paddles match the sizes of all image receptors | [9 empty] | kV:
22 | 900.12(b)(8)(iiE) |  | Chest wall edge may be bent upwards and does not appear in the image | [6 empty] | TBD |  |  | Compression paddles do not flex by more than 1 cm | [9 empty] | mA:
23 | 900.12(b)(9i) |  | Manual selection of mAs or mA and/or time is available | [6 empty] | TBD |  |  | Compression paddles are flat and parallel to the breast support table | [9 empty] | Bkg Density:
24 | 900.12(b)(9ii) |  | Prior indication of selected technique factors is provided | [6 empty] | TBD |  |  | Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor | [9 empty] | Inside disk:
25 | 900.12(b)(9iii) |  | Post exposure display of technique factors is provided for AEC exposures | [6 empty] | TBD |  |  | Chest wall edge may be bent upwards for patient comfort and does not appear in the image | [9 empty] | Outside disk:
26 | 900.12(b)(10i) |  | AEC mode is operable in all combinations of equipment configuration | [6 empty] | TBD |  |  | Manual selection of mAs or mA and/or time is available | [9 empty] | Difference:
27 | 900.12(b)(10ii) |  | AEC sensor positioning is flexible | [6 empty] | NA |  |  | Prior indication of selected technique factors is provided | [9 empty] | Fibers:
28 |  |  | AEC sensor position and size is indicated on the compression paddle | [6 empty] | NA |  |  | Post exposure display of technique factors is provided for AEC exposures | [9 empty] | Specks:
29 |  |  | Selected position of the AEC sensor is clearly indicated | [6 empty] | NA |  |  | AEC mode is operable in all combinations of equipment configuration | [9 empty] | Masses:
30 | 900.12(b)(10iii) |  | AEC density control setting is provided | [6 empty] | NA |  | 3 | AEC sensor positioning is flexible | [9 empty] | Image Quality (Acq Unit)
31 | 900.12(b)(11) |  | X-ray film used is appropriate for mammography | [6 empty] | NA |  | 3 | AEC sensor position and size is indicated on the compression paddle | [9 empty] | Target/Filter:
32 | 900.12(b)(12) |  | Intensifying screens used are appropriate for mammography and matched to film | [6 empty] | NA |  | 3 | Selected position of the AEC sensor is clearly indicated | [9 empty] | kV:
33 | 900.12(b)(13) |  | Film processing chemicals are appropriate for the film being used | [6 empty] | NA |  | 3 | AEC density control setting is provided | [9 empty] | mA:
34 | 900.12(b)(14) |  | Lights for film illumination i.e. hot lights are available | [6 empty] | TBD |  | 3 | X-ray film used is appropriate for mammography | [9 empty] | Fibers:
35 | 900.12(b)(15) |  | Film masking devices are available to all interpreting physicians | [6 empty] | TBD |  | 3 | Intensifying screens used are appropriate for mammography and matched to the film used | [9 empty] | Specks:
36 | [11 empty] | 3 | Film processing chemicals are appropriate for the film being used | [9 empty] | Masses:
37 |  |  | Printer passes manufacturer QC tests | [6 empty] | TBD |  |  | Lights for film illumination i.e. hot lights are available | [9 empty] | Target/Filter:
38 |  |  | Repeat analysis performed | [6 empty] | TBD |  |  | Film masking devices are available to all interpreting physicians | [9 empty] | kV:
39 |  |  | Control panel lights | [6 empty] | TBD | [6 empty] | Acquisition Monitor | [5 empty] | mA:
40 |  |  | System stops above floor | [6 empty] | TBD |  |  | Gray scale contrast: 5%/95% squares visible | [9 empty] | Fibers:
41 |  |  | Motorized movements | [6 empty] | TBD |  |  | Spatial resolution: All lines in all bar patterns can be resolved | [9 empty] | Specks:
42 |  |  | Lifts from lowest to highest position | [6 empty] | TBD |  |  | Visual inspection: No visible streaking | [9 empty] | Masses:
43 |  |  | Foot switches functional | [6 empty] | TBD |  |  | Visual inspection: No fluttering in image | [9 empty] | Target/Filter:
44 |  |  | Movements blocked on error | [6 empty] | TBD |  |  | Visual inspection: No shadows in image | [9 empty] | kV:
45 |  |  | Emergency stop button functional | [6 empty] | TBD | [12 empty] | mA:
46 |  |  | Grid motion | [6 empty] | TBD | [12 empty] | Fibers:
47 |  |  | No sharp edges or cracks on table or compression paddles | [6 empty] | TBD | [6 empty] | Mammography Phototimer Technique Chart | [5 empty] | Specks:
48 |  |  | Radiation shield present | [6 empty] | TBD | [12 empty] | Masses:
49 |  |  | Manual compression functional | [6 empty] | TBD |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  | SNR:
50 |  |  | Autodecompression functional | [6 empty] | TBD |  | Breast | [10 empty] | CNR:
51 |  |  | Decompression button on control box functional | [6 empty] | TBD |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  | Tomo/3D
52 |  |  | Emergency decompression | [6 empty] | TBD |  | <3 cm | [10 empty] | MGD 20mm:
53 |  |  | Measured phantom thickness is 39-45 mm | [6 empty] | TBD |  | 3 to 5 cm | [10 empty] | MGD 40mm:
54 | [11 empty] | 5 to 7 cm | [10 empty] | MGD 60mm:
55 | [5 empty] | Acquisition Monitor | [5 empty] | >7 cm | [10 empty] | Geometry x P1:
56 |  |  | Gray scale contrast: 5%/95% squares visible | [6 empty] | TBD | [12 empty] | Geometry y P1:
57 |  |  | Spatial resolution: All lines in all bar patterns can be resolved | [6 empty] | TBD |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  | Geometry x P2:
58 |  |  | Visual inspection: No visible streaking | [6 empty] | TBD |  | Breast | [10 empty] | Geometry y P2:
59 |  |  | Visual inspection: No fluttering in image | [6 empty] | TBD |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  | Geometry x P3:
60 |  |  | Visual inspection: No shadows in image | [6 empty] | TBD |  | <3 cm | [10 empty] | Geometry y P3:
61 | [11 empty] | 3 to 5 cm | [10 empty] | ASF(a) P1:
62 | [5 empty] | Tomosynthesis | [5 empty] | 5 to 7 cm | [10 empty] | ASF(b) P1:
63 |  |  | AGD for 20, 40, 60 mm PMMA thickness is below maximum allowed MGD | [6 empty] | TBD |  | >7 cm | [10 empty] | ASF(a) P2:
64 |  |  | AGD in 2D + Tomoscan mode is below 3 mGy | [6 empty] | TBD | [12 empty] | ASF(b) P2:
65 |  |  | X/Y geometric accuracy is within +/- 2% of actual phantom size | [6 empty] | TBD | [6 empty] | Inspection Results | [5 empty] | ASF(a) P3:
66 |  |  | Z resolution (ASF(a) and ASF(b)) is acceptable | [6 empty] | TBD |  | Compression Force Indicator |  |  | Indicated | Measured | [6 empty] | ASF(b) P3:
67 |  |  | Phantom image quality is acceptable | [6 empty] | TBD |  |  |  | Max force (N): | [8 empty] | 0 deg
68 |  |  | Collimator or compression paddle are not visible in the +/- 25 degree images | [6 empty] | TBD |  |  |  | Force maintained for 1 minute: |  | Enter 1 for YES, 2 for NO, 3 for NA | [6 empty] | Fibers:
69 |  |  | Image is free of clinically relevant artifacts (dust, defective pixels, damage) | [6 empty] | TBD | [12 empty] | Specks:
70 | [12 empty] | Criteria: | Maximum force is between 12 - 20 kg (120-200 N) | [8 empty] | Masses:
71 | Date: | [8 empty] | Inspector: | Eugene Mah |  |  | Change in compression force after 1 min is within +/- 20 N | [8 empty] | 90 deg
72 | Room Number | [8 empty] | Survey ID: | [12 empty] | Fibers:
1 | [10 empty] | Medical University of South Carolina | Compression Thickness Indicator | [10 empty] | Specks:
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | BR-12 (cm) | Indicated | Difference | [7 empty] | Masses:
3 | Compression Force Indicator | [11 empty] | 1
4 |  |  | Max force (N): |  |  | Acceptable: | [6 empty] | 2
5 |  |  | Force maintained for 1 minute: | [9 empty] | 4 |  |  | Mean:
6 |  | Criteria: | Maximum force is between 12 - 20 kg (120-200 N) | [9 empty] | 4 |  |  | StDev:
7 |  |  | Change in compression force after 1 min is within +/- 20 N | [9 empty] | 4
8 | [12 empty] | 4 |  |  |  | Criteria: | Difference between indicated and actual thickness is <= 5 mm
9 | Compression Thickness Indicator | [11 empty] | 6 |  |  |  |  | Reproducibility is <= 2 mm
10 | BR-12 (cm) | Indicated | Difference |  |  |  | Paddle flex is less than 1 cm: | YES |  |  |  |  | 8
11 | 1 | [5 empty] | Overall unit assembly: | YES |  |  |  | Enter 1 for YES, 2 for NO, 3 for NA
12 | 2 | [5 empty] | Compression thickness indicator: | TBD |  |  |  | 1 | Paddle flex is less than 1 cm |  |  |  | Compression thickness indicator:
13 | 4 | [5 empty] | Mean: |  |  |  |  | 1 | Overall unit assembly
14 | 4 | [5 empty] | StDev:
15 | 4 | [10 empty] | Artifact Evaluation
16 | 4 |  |  | Criteria: | Difference between indicated and actual thickness is <= 5 mm | [7 empty] | Target:
17 | 6 |  |  |  | Reproducibility is <= 2 mm | [7 empty] | Filter:
18 | 8 | [11 empty] | 2D/3D:
19 | [12 empty] | kV:
20 | Artifact Evaluation | [11 empty] | mAs:
21 | Target: | [11 empty] | EI:
22 | Filter: | [11 empty] | Acceptable:
23 | 2D/3D: | [11 empty] | Criteria: | No clinically relevant artifacts detected
24 | Focal spot: | [11 empty] | Comments:
25 | kV: | [11 empty] | Do not exceed 390 chars
26 | mAs: | [11 empty] | New:
27 | EI:
28 | Acceptable: | [10 empty] | Light Field Brightness |  |  | 1 | Enter 1 for YES, 2 for NO, 3 for NA
29 | Criteria: | No clinically relevant artifacts detected | [10 empty] | Measurements (lux) |  |  |  |  | Average Illumination
30 |  | Comments: | [10 empty] | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux)
31 | [19 empty] | Acceptable:
32 | [12 empty] | Criteria: | Average light field illumination >= 160 lux
33
34 | Light Field Brightness | [10 empty] | Light Field/X-ray Field Alignment
35 |  | Measurements (lux) |  |  |  |  | Average Illumination | [5 empty] | Target/Filter: |  |  |  | Target/Filter:
36 |  | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux) | [5 empty] | kV: |  |  |  | kV:
37 | [8 empty] | Acceptable: |  |  |  | mAs: |  |  |  | mAs:
38 |  | Criteria: | Average light field illumination >= 160 lux. | [10 empty] | Deviation (mm) |  |  |  | Deviation (mm)
39 | [13 empty] | 24x30 | 18x24 Low | 18x24 High |  | 24x30 | 18x24 Low | 18x24 High
40 | Light Field/X-ray Field Alignment |  |  |  |  | X-ray Field/Image Receptor Alignment | [6 empty] | Left:
41 | Target/Filter: |  |  |  |  | Target/Filter: |  |  | Target/Filter: |  |  |  | Right:
42 |  | Deviation (mm) |  |  |  | Deviation (mm) |  |  | Deviation (mm) |  |  |  | Anterior:
43 |  | 24x30 | 18x24 Low | 18x24 High |  | 24x30 | 18x24 Low | 18x24 High | 24x30 | 18x24 Low | 18x24 High |  | Chest Wall:
44 | Left: |  |  |  | Left: | [7 empty] | Left/Right:
45 | Right: |  |  |  | Right: | [7 empty] | Ant/Chest:
46 | Anterior: |  |  |  | Anterior: | [7 empty] | Criteria: | Sum of left and right edge deviation is less than 2% SID (13 mm)
47 | Chest Wall: |  |  |  | Chest Wall: | [8 empty] | Sum of anterior and chest wall edge deviation is less than 2% SID (13mm)
48 | Left/Right: |  |  |  | Paddle:
49 | Ant/Chest: |  |  |  | Missed Tissue: | [6 empty] | X-ray Field/Image Receptor Alignment
50 | Target/Filter: | [5 empty] | Light field/X-ray field alignment acceptable | [6 empty] | Deviation (mm) |  |  |  | Deviation (mm)
51 | Left: | [5 empty] | X-ray field/image receptor alignment acceptable | [6 empty] | 24x30 | 18x24 Low | 18x24 High |  | 24x30 | 18x24 Low | 18x24 High
52 | Right: | [5 empty] | Chest wall missed tissue/Compression paddle acceptable | [5 empty] | Left:
53 | Anterior: |  |  |  | Criteria: | Sum of left and right edge deviation is less than 13 mm | [6 empty] | Right:
54 | Chest Wall: |  |  |  |  | Sum of anterior and chest wall edge deviation is less than 13mm | [6 empty] | Anterior:
55 | Left/Right: |  |  |  |  | Compression paddle deviation is <= 6.5mm | [6 empty] | Chest Wall:
56 | Ant/Chest: |  |  |  |  | Chest wall missed tissue is <= 5 mm | [6 empty] | Compression Paddle:
57 | [5 empty] | X-ray field does not extend beyond the detector by more than 2% SID (13 mm) at any one side | [6 empty] | Missed Tissue:
58 | [12 empty] | Criteria: | X-ray field does not extend beyond the detector by more than 2% SID (13 mm) at any one side
59 | System Resolution | [12 empty] | Chest wall edge of the compression paddle does not extend beyond the detector by more than 6.5 mm
60 | Target: | [12 empty] | Chest wall missed tissue is <= 5 mm
61 | kV: |  |  |  | Acceptable:
62 | mAs: | [10 empty] | System Resolution
63 | lp/mm: | [11 empty] | Target/Filter:
64 | Criteria: | Limiting system resolution must be 6 lp/mm or higher | [10 empty] | kV:
65 | [12 empty] | mAs:
66 | [12 empty] | lp/mm:
67 | [12 empty] | Acceptable:
68 | [12 empty] | Criteria: | Limiting system resolution must be 6 lp/mm or higher
69
70 | [11 empty] | AEC Thickness Tracking
71 | Date: | [8 empty] | Inspector: | Eugene Mah | AEC Mode: |  |  | AEC Position: |  |  | Target/Filter: | W/Rh
72 | Room Number | [8 empty] | Survey ID: |  |  | Thickness
1 | [10 empty] | Medical University of South Carolina |  | (cm) | AEC Region | kV | mAs | Mean Pixel | Deviation
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | 2 | [6 empty] | Average:
3 | AEC Thickness Tracking | [11 empty] | 4 | [6 empty] | SD:
4 | AEC Mode: |  |  | AEC Position: |  |  | Target/Filter: | W/Rh |  |  |  |  | 6 | [6 empty] | Coeff of Var:
5 | Thickness | [11 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [5 empty] | Acceptable:
6 | (cm) | AEC Region | kV | mAs | Mean Pixel | Deviation | [7 empty] | Coefficient of variation <= 5%
7 | 2 | [6 empty] | Average:
8 | 4 | [6 empty] | SD: |  |  |  |  | AEC security cutoff is activated
9 | 6 | [6 empty] | Coeff of Var: |  |  |  |  | Visible signal indicates AEC security cutoff and subsequent exposures blocked
10 | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [5 empty] | Acceptable:
11 |  | Coefficient of variation <= 5% | [9 empty] | SNR, CNR, AEC Reproducibility
12 | [11 empty] | AEC Mode: |  |  | Target/Filter: | W/Rh | kV: | 28
13 | TBD | AEC security cutoff is activated
14 | TBD | Visible signal indicates AEC security cutoff and subsequent exposures blocked | [16 empty] | Entrance
15 | [12 empty] | mAs | Mass mean | Bkg mean | Bkg SD | SNR | CNR | Dose (mGy) |  | Mean Dev | SNR Dev
16 | SNR, CNR, AEC Reproducibility
17 | AEC Mode: |  |  | Target/Filter: | W/Rh | kV: | 28
18
19 | [7 empty] | Entrance
20 |  | mAs | Mass mean | Bkg mean | Bkg SD | SNR | CNR | Dose (mGy) | Mean Dev | SNR Dev
21 | [11 empty] | Average: | [7 empty] | Max:
22 | [11 empty] | SD:
23 | [11 empty] | Coeff of Var:
24
25 | [12 empty] | Criteria: | Coefficient of variation in mAs, entrance dose is <= 5%
26 | Average: | [12 empty] | Maximum deviation of Mass mean/SNR is <= 15%
27 | SD:
28 | Coeff of Var: | [10 empty] | Detector Uniformity
29 | [12 empty] | Target/Filter | / | / | / |  | Deviation from Mean
30 | Criteria: | Coefficient of variation in mAs, entrance dose is <= 5% | [10 empty] | kV |  |  |  |  | / | / | /
31 |  | Maximum deviation of Mass mean/SNR is <= 15% | [10 empty] | mAs
32 | [12 empty] | ROI 1
33 | Detector Uniformity | [11 empty] | ROI 2
34 | Target/Filter | / | / | / |  | Deviation from Mean | [6 empty] | ROI 3
35 | kV |  |  |  |  | / | / | / |  |  |  |  | ROI 4
36 | mAs | [11 empty] | ROI 5
37 | ROI 1 | [11 empty] | Average
38 | ROI 2 | [11 empty] | Max dev
39 | ROI 3
40 | ROI 4 | [11 empty] | Criteria: | Maximum deviation from average is <= 7%
41 | ROI 5
42 | Average | [10 empty] | Mean Glandular Dose - 2D
43 | Max dev | [11 empty] | Phantom: | 800004-1204146 |  | kV: | 28
44 | [12 empty] | Target/Filter: | W/Rh |  | mA: | 100
45 | Criteria: | Maximum deviation from average is <= 7% | [13 empty] | Reported |  | Air KERMA
46 | [13 empty] | mAs | Bkg mean | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #N/A
47 | Mean Glandular Dose - 2D | [12 empty] | 0 | 0 |  | 0 | #DIV/0! |  | ESE (mGy): | #DIV/0!
48 | Phantom: | 800004-1204146 |  | kV: | 28 | [8 empty] | 0 | 0 |  | 0 | #DIV/0! |  | DGN (mrad/R): | #N/A
49 | Target/Filter: | W/Rh |  | mA: | 100 | [8 empty] | 0 | 0 |  | 0 | #DIV/0! |  | MGD (mGy): | #DIV/0!
50 |  |  |  | Reported |  | Air KERMA | [7 empty] | 0 | 0 |  | 0 | #DIV/0! |  | Prev MGD (mGy):
51 |  | mAs | Bkg mean | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #N/A |  |  |  | Average: | 0 | 0 |  | 0 | #DIV/0! |  | % Change: | #DIV/0!
52 |  | 0 | 0 |  | 0 | #DIV/0! |  | ESE (mGy): | #DIV/0! |  |  |  | Coeff of Var: | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
53 |  | 0 | 0 |  | 0 | #DIV/0! |  | DGN (mrad/R): | #N/A |  |  |  | Criteria: | MGD must be less than 3 mGy and should be less than 2 mGy | [5 empty] | Deviation from Reported MGD: | #DIV/0!
54 |  | 0 | 0 |  | 0 | #DIV/0! |  | MGD (mGy): | #DIV/0! |  |  |  |  | Deviation from reported MGD is less than 15% | [5 empty] | mAs at 3 mGy: | #DIV/0!
55 |  | 0 | 0 |  | 0 | #DIV/0! |  | Prev MGD (mGy):
56 | Average: | 0 | 0 |  | 0 | #DIV/0! |  | % Change: | #DIV/0! |  |  | Mean Glandular Dose - Tomo
57 | Coeff of Var: | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | [6 empty] | PMMA | kV | mAs | MGD (mGy) | Prev MGD | Max MGD | Acceptable
58 | Criteria: | MGD must be less than 3 mGy and should be less than 2 mGy | [5 empty] | Deviation from Reported MGD: | #DIV/0! |  |  |  | 20 | 26 |  |  |  | 1.1
59 |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | #DIV/0! |  |  |  | 40 | 28 |  |  |  | 2.2
60 | [12 empty] | 60 | 30 |  |  |  | 4.7
61 | Mean Glandular Dose - Tomo | [6 empty] | Mean Glandular Dose - 2D + Tomo |  |  |  |  | Criteria: | Reported MGD is less than the max MGD
62 | PMMA | kV | mAs | MGD (mGy) | Max MGD | Acceptable |  |  |  |  | Measured
63 | 20 | 26 |  |  | 1.1 |  |  |  | kV | mAs | AGD (mGy) | Mean Glandular Dose - 2D + Tomo
64 | 40 | 28 |  |  | 2.2 |  |  | 2D | 28 |  |  |  | Combo 2D |  |  | Reported |  | Air KERMA
65 | 60 | 30 |  |  | 4.7 |  |  | Tomoscan | 28 |  |  |  |  | mAs | Bkg mean | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #N/A
66 | Criteria: | Reported MGD is less than the max MGD | [7 empty] | 2D + Tomoscan: | [9 empty] | ESE (mGy):
67 | [9 empty] | Acceptable: | [9 empty] | DGN (mrad/R): | #N/A
68 | [19 empty] | MGD (mGy):
69 | [19 empty] | Prev MGD (mGy):
70 | [12 empty] | Average: | [6 empty] | % Change:
71 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Coeff of Var:
72 | Room Number | [8 empty] | Survey ID: |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
1 | [10 empty] | Medical University of South Carolina |  |  | Deviation from reported MGD is less than 15% | [5 empty] | mAs at 3 mGy:
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection
3 | Meter: | Piranha | SN: |  |  | Calibration Date: | [6 empty] | Combo Tomo |  |  | Reported |  | Air KERMA
4 | [5 empty] | Calibration Due: | [7 empty] | mAs | Bkg mean | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #N/A
5 | kVp Accuracy and Output | [18 empty] | ESE (mGy):
6 |  |  |  | Measured Values | [15 empty] | DGN (mrad/R): | #N/A
7 | Target/Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s | kV Error | [10 empty] | MGD (mGy):
8 | / | 24 | 50 | [16 empty] | Prev MGD (mGy):
9 |  | 25 | 50 | [9 empty] | Average: | [6 empty] | % Change:
10 |  | 26 | 50 | [9 empty] | Coeff of Var:
11 |  | 28 | 50 | [9 empty] | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
12 |  | 30 | 50 | [10 empty] | Deviation from reported MGD is less than 15% | [5 empty] | mAs at 3 mGy:
13 |  | 32 | 50
14 |  | 34 | 50 | [12 empty] | Reported | Measured
15 | / | 28 | 50 | [10 empty] | kV | mAs | AGD (mGy) | AGD (mGy)
16 |  | 30 | 50 | [9 empty] | 2D | 28 |  |  |  | Acceptable:
17 |  | 32 | 50 | [9 empty] | Tomoscan | 28
18 |  | 34 | 50 | [11 empty] | 2D + Tomoscan: |  |  | mGy
19 |  | 36 | 50 | [9 empty] | Criteria: | MGD is less than 3 mGy
20 | / | 28 | 50
21 |  | 30 | 50 | [8 empty] | Geometric Accuracy - X/Y
22 |  | 32 | 50 | [9 empty] | Phantom dimensions | x (mm): | 101.3
23 |  | 34 | 50 | [10 empty] | y (mm): | 108 | Previous
24 |  | 38 | 50 | [9 empty] | x (mm) | y (mm) | Acceptable | x (mm) | y (mm)
25 | [11 empty] | Position 1
26 | [11 empty] | Position 2
27 | [11 empty] | Position 3
28 | [12 empty] | Criteria: | x/y measurements within +/- 2% of actual phantom dimensions
29
30 | [11 empty] | Geometric Accuracy - Z resolution
31 | [11 empty] | Slice height | z=-2: | -2 | z=-1: | -1 | z=0: |  | z=+1: | 1 | z=+2: | 2
32 | [11 empty] | Position 1 | ROI max | Bkg mean | ROI max | Bkg mean | ROI max | Bkg mean | ROI max | Bkg mean | ROI max | Bkg mean
33 | [11 empty] | 1
34 | [11 empty] | 2
35 | [11 empty] | 3
36 | [11 empty] | 4
37 | [11 empty] | 5
38 | [11 empty] | 6
39 | [11 empty] | Position 2 | z=-2: | -2 | z=-1: | -1 | z=0: |  | z=+1: | 1 | z=+2: | 2
40 | [11 empty] | 1
41 | [11 empty] | 2
42 | [11 empty] | 3
43 | [11 empty] | 4
44 | [11 empty] | 5
45 | [11 empty] | 6
46 | [11 empty] | Position 3 | z=-2: | -2 | z=-1: | -1 | z=0: |  | z=+1: | 1 | z=+2: | 2
47 | [11 empty] | 1
48 | [11 empty] | 2
49 | [11 empty] | 3
50 | [11 empty] | 4
51 | [11 empty] | 5
52 | [11 empty] | 6
53
54 | kVp and Output Reproducibility | [11 empty] | Position 1 |  |  | Position 2 |  |  | Position 3
55 |  | Target/Filter: | / | kV: | 28 |  | Target/Filter: | / | kV: | 28 |  |  | s(zi) | s(bgi) | ASF(i) | s(zi) | s(bgi) | ASF(i) | s(zi) | s(bgi) | ASF(i)
56 |  |  |  | mAs: | 50 |  |  |  | mAs: | 50 |  | z=-2
57 |  | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s | z=-1
58 | [11 empty] | z=0
59 | [11 empty] | z=+1
60 | [11 empty] | z=+2
61 | [12 empty] | ASF(a) |  |  | ASF(a) |  |  | ASF(a)
62 | Mean: | [11 empty] | ASF(b) |  |  | ASF(b) |  |  | ASF(b)
63 | Std Dev: | [10 empty] | Previous | ASF(a) |  |  | ASF(a) |  |  | ASF(a)
64 | Coeff of Var: | [11 empty] | ASF(b) |  |  | ASF(b) |  |  | ASF(b)
65 |  |  |  |  | Output > 2.7 mGy/s: |  |  |  |  | Output > 2.7 mGy/s: |  |  | Criteria: | ASF(a) is < 0.9 for all positions. ASF(b) is < 0.6 for all positions
66
67 |  | Criteria: | Coefficient of variation <= 2% | [8 empty] | Phantom Image Quality - 3D
68 |  |  | Minimum acceptable output is 2.7 mGy/s at 28 kVp for W/Rh | [9 empty] | 0 deg | 90 deg |  | Previous | 0 deg | 90 deg
69 | [11 empty] | Fibers: |  |  |  | Fibers:
70 | [11 empty] | Specks: |  |  |  | Specks:
71 | Date: | [8 empty] | Inspector: | Eugene Mah | Masses: |  |  |  | Masses:
72 | Room Number | [8 empty] | Survey ID: |  | Acceptable:
1 | [10 empty] | Medical University of South Carolina |  | Criteria: | Must be able to see 4 fibers, 3 speck groups, 3 masses
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Test 3: Radiation field - Collimator or compression paddle are not visible in the +/- 25 degree images
3 | Exposure Linearity | [11 empty] | Test 4: Tomo Phantom Image Quality - Minmum number of objects visible
4 | Target: | / | Filter: | 28 | [8 empty] | Test 5: Artifact Detection - No clinically relevant artifacts found
5
6 | Set mAs | kVp | mGy | mGy/mAs | mGy/s |  |  | Linearity acceptable |  |  |  | kVp Accuracy and Output
7 | 20 | [5 empty] | NA | Unit can maintain output for 3 sec at 28 kVp |  |  |  | Meter: | Piranha |  | Calibration Date:
8 | 50 | [5 empty] | NA | Output > 2.7 mGy/s averaged over 3 sec |  |  |  | Serial #: |  |  | Calibration Due:
9 | 100 | [14 empty] | Measured Values
10 | 320 | [11 empty] | Target/Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s | kV Error
11 |  |  | Coeff of Linearity: | [9 empty] | / | 24 | 50
12 |  | Criteria: | Coefficient of linearity is <= 10% | [10 empty] | 25 | 50
13 | Half Value Layer | [12 empty] | 26 | 50
14 | Target/Filter | / | / | / | [9 empty] | 28 | 50
15 | kVp | 28 | 28 | 28 | [9 empty] | 30 | 50
16 | HVL (mm Al) | [12 empty] | 32 | 50
17 | 0 mm diff | [12 empty] | 34 | 50
18 | Acceptable: | [11 empty] | / | 28 | 50
19 |  | Criteria: | HVL at 28 kVp is greater than 0.28 mm Al | [10 empty] | 30 | 50
20 |  |  | Exposure deviation for 0 mm Al measurements is less than 2% | [10 empty] | 32 | 50
21 | Signal to Noise Ratio | [12 empty] | 34 | 50
22 |  | kV: | 28 |  |  | Current | Previous | Baseline | Change | Acceptable |  |  |  | 36 | 50
23 |  | mA: | 100 |  | SNR: | [8 empty] | 38 | 50
24 |  | Target: |  |  | CNR: | [7 empty] | / | 28 | 50
25 |  | Filter: | [11 empty] | 30 | 50
26 |  | Criteria: | SNR must be >= 40. CNR must be >= 2. | [10 empty] | 32 | 50
27 |  |  | SNR/CNR must not differ from baseline by more than +/- 15% | [10 empty] | 34 | 50
28 | [13 empty] | 38 | 50
29 | Image Quality Evaluation | [11 empty] | Criteria: | kV accuracy is <= 5%
30 |  | Printer |  |  |  | Acq Workstation |  |  |  | Review Workstation
31 |  | kV: |  |  | Mode: | 2D |  | 3D |  | Left | Right | kV, Exposure, and Output Reproducibility
32 |  | mA: |  |  | Target/Filter |  |  |  |  | Workstation: |  |  | Target/Filter: | / | kV: | 28 |  | Target/Filter: | / | kV: | 28
33 |  | Background Density: |  |  | kV: |  |  |  | Fibers: | [5 empty] | mAs: | 50 |  |  |  | mAs: | 50
34 |  | Inside disk: |  |  | mA: |  |  |  | Specks: |  |  |  | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
35 |  | Outside disk: |  |  | Fibers: |  |  |  | Masses:
36 |  | Difference: |  |  | Specks: |  |  |  |  | Workstation:
37 |  | Fibers: |  |  | Masses: |  |  |  | Fibers:
38 |  | Specks: | [6 empty] | Specks:
39 |  | Masses: | [6 empty] | Masses: |  |  | Mean:
40 |  | Criteria: | Must see at least 5 fibers, 4 specks, 4 masses | [8 empty] | Std Dev:
41 | [11 empty] | Coeff of Var:
42 | [11 empty] | Previous:
43 | Geometric Accuracy - X/Y | [11 empty] | Criteria: | Coefficient of variation <= 2%
44 |  | Phantom dimensions | x (mm): | 101.3 | [9 empty] | Minimum acceptable output is 2.7 mGy/s at 28 kVp for W/Rh
45 |  |  | y (mm): | 108
46 |  | x (mm) | y (mm) | Acceptable | [7 empty] | Exposure Linearity
47 | Position 1 | [11 empty] | Target/Filter: | / | kV: | 28
48 | Position 2 | [11 empty] | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
49 | Position 3 | [11 empty] | 20
50 |  | Criteria: | x/y measurements within +/- 2% of actual phantom dimensions | [9 empty] | 50
51 | [12 empty] | 100
52 | Geometric Accuracy - Z resolution | [11 empty] | 320
53 |  | z=0 (mm) | ASF(a) | ASF(b) | Acceptable | [10 empty] | Coeff of Linearity:
54 | Position 1 | [11 empty] | Criteria: | Coefficient of linearity is <= 10%
55 | Position 2 | [10 empty] | 3 | Unit Can Maintain Output for 3 sec at 28 kVp
56 | Position 3 | [10 empty] | 3 | Radiation output is > 2.7 mGy/s averaged over 3 s
57 |  | Criteria: | ASF(a) is < 0.9 for all positions. ASF(b) is < 0.6 for all positions | [8 empty] | Half Value Layer |  | 1 | Compression paddle in place? (1 for Yes, 2 for No)
58 | [12 empty] | Target/Filter | / | / | /
59 | Phantom Image Quality - 3D | [11 empty] | kV | 28 | 28 | 28
60 |  | 0 deg | 90 deg | [9 empty] | Al (mm) | Air Kerma (mGy)
61 | Fibers: | [11 empty] | 0
62 | Specks: | [11 empty] | 0
63 | Masses: | [11 empty] | HVL
64 | Acceptable: | [11 empty] | 0 mm
65 |  | Criteria: | Must be able to see 4 fibers, 3 speck groups, 3 masses | [9 empty] | Acceptable:
66 | [12 empty] | Criteria: | HVL at 28 kVp is greater than 0.28 mm Al
67 |  | Test 3: Radiation field - Collimator or compression paddle are not visible in the +/- 25 degree images | [11 empty] | Exposure deviation for 0 mm Al measurements is less than 2%
68 |  | Test 4: Tomo Phantom Image Quality - Minmum number of objects visible | [9 empty] | Image Quality Evaluation
69 |  | Test 5: Artifact Detection - No clinically relevant artifacts found | [9 empty] | Printer | Model: |  | SN: |  |  | QC Manual Version:
70
71 | Date: | [8 empty] | Inspector: | Eugene Mah |  |  |  | Previous |  |  |  | Previous
72 | Room Number | [8 empty] | Survey ID: |  |  | kV: |  |  |  | Fibers:
1 | [10 empty] | Medical University of South Carolina |  | mA: |  |  |  | Specks:
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Background Density: |  |  |  | Masses:
3 | [12 empty] | Inside disk: |  |  |  | Acceptable:
4 | Comments: | [11 empty] | Outside disk: |  |  |  | Criteria: | Must be able to see 5 fibers, 4 speck groups, 4 masses
5 | [12 empty] | Difference:
6
7 | [11 empty] | Workstation
8 | [13 empty] | Workstation: |  |  |  | Workstation:
9 | [13 empty] | Left | Right |  |  | Left | Right
10 | [12 empty] | Fibers: |  |  |  | Fibers:
11 | [12 empty] | Specks: |  |  |  | Specks:
12 | [12 empty] | Masses: |  |  |  | Masses:
13
14 | [11 empty] | Acquisition Unit |  |  | Phantom serial number: | 800004-1204146
15 | [12 empty] | Mode: | 2D |  | 3D
16 | [12 empty] | Target/Filter: |  |  |  | Previous | Target/Filter
17 | [12 empty] | kV: |  |  |  |  | kV
18 | [12 empty] | mAs: |  |  |  |  | mAs
19 | [12 empty] | Fibers: |  |  |  |  | Fibers
20 | [12 empty] | Specks: |  |  |  |  | Specks
21 | [12 empty] | Masses: |  |  |  |  | Masses
22
23 | [12 empty] | Criteria: | 2D - Must see at least 5 fibers, 4 specks, 4 masses
24 | [13 empty] | 3D - Must see at least 4 fibers, 3 specks, 3 masses
25 | [11 empty] | Signal to Noise
26 | [16 empty] | Mean | SD
27 | [12 empty] | kV: | 28 |  | Mass:
28 | [12 empty] | mA: | 100 |  | Background:
29 | [12 empty] | Target:
30 | [12 empty] | Filter: |  |  |  | Current | Previous | Change | Baseline | Acceptable
31 | [15 empty] | SNR:
32 | [15 empty] | CNR:
33 | [12 empty] | Criteria: | SNR must be >= 40. CNR must be >= 2.
34 | [13 empty] | SNR/CNR must not differ from baseline by more than +/- 15%
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56 | [16 empty] | COMMENTS
57 | [12 empty] | Note: Type all comments in the shaded cells.
58 | [15 empty] | Previous Year Comments:
59 | [12 empty] | Comments:
60 | [12 empty] | Do NOT exceed 390 |  | 0 |  | 0
61 | [12 empty] | Additional comments:
62 | [12 empty] | Do NOT exceed 390 |  | 0 |  | 0
63 | [12 empty] | Additional comments:
64 | [12 empty] | Do NOT exceed 390 |  | 0 |  | 0
65 | [12 empty] | Additional comments:
66 | [12 empty] | Do NOT exceed 390 |  | 0 |  | 0
67 | [12 empty] | Additional comments:
68 | [12 empty] | Do NOT exceed 390 |  | 0 |  | 0
69 | [12 empty] | Additional comments:
70 | [12 empty] | Do NOT exceed 390 |  | 0 |  | 0
71 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Additional comments:
72 | Room Number | [8 empty] | Survey ID: |  |  | Do NOT exceed 390 |  | 0 |  | 0
Additional comments:
[13 empty] | Do NOT exceed 390 |  | 0 |  | 0

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
| kVp | [9 empty] | kVp | [9 empty] | kVp
HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
0.23 | 116 | [8 empty] | 0.23 | [9 empty] | 0.23
0.24 | 121 | 124 | [7 empty] | 0.24 | [9 empty] | 0.24
0.25 | 126 | 129 | 131 | [6 empty] | 0.25 | [9 empty] | 0.25
0.26 | 130 | 133 | 135 | 138 | [5 empty] | 0.26 | [9 empty] | 0.26
0.27 | 135 | 138 | 140 | 142 | 143 |  |  |  |  | 0.27 | [9 empty] | 0.27
0.28 | 140 | 142 | 144 | 146 | 147 | 149 |  |  |  | 0.28 |  |  | 149 | 151 | 154 |  |  |  |  | 0.28 |  |  | 150 | 155 | 159
0.29 | 144 | 146 | 148 | 150 | 151 | 153 | 154 |  |  | 0.29 |  |  | 154 | 156 | 158 | 159 |  |  |  | 0.29 |  |  | 155 | 160 | 164 | 168
0.3 | 149 | 151 | 153 | 155 | 156 | 157 | 158 | 159 |  | 0.3 |  |  | 158 | 160 | 162 | 162 | 163 |  |  | 0.3 |  |  | 160 | 164 | 168 | 172 | 176
0.31 | 154 | 156 | 157 | 159 | 160 | 161 | 162 | 163 | 164 | 0.31 |  |  | 163 | 164 | 166 | 166 | 167 | 167 |  | 0.31 |  |  | 165 | 168 | 172 | 174 | 180 | 182
0.32 | 158 | 160 | 162 | 163 | 164 | 166 | 167 | 168 | 168 | 0.32 |  |  | 167 | 169 | 171 | 171 | 171 | 172 | 172 | 0.32 |  |  | 169 | 173 | 177 | 181 | 184 | 186 | 188
0.33 | 163 | 165 | 166 | 168 | 169 | 170 | 171 | 173 | 173 | 0.33 |  |  | 171 | 173 | 175 | 176 | 176 | 176 | 176 | 0.33 |  |  | 174 | 178 | 181 | 185 | 188 | 190 | 192
0.34 | 168 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 0.34 |  |  | 176 | 178 | 179 | 179 | 180 | 180 | 180 | 0.34 |  |  | 179 | 183 | 186 | 190 | 193 | 195 | 196
0.35 |  | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 0.35 |  |  | 180 | 181 | 183 | 183 | 184 | 185 | 185 | 0.35 |  |  | 184 | 187 | 190 | 194 | 197 | 199 | 201
0.36 |  |  | 179 | 181 | 182 | 183 | 184 | 185 | 185 | 0.36 |  |  | 185 | 186 | 187 | 187 | 188 | 188 | 189 | 0.36 |  |  | 189 | 192 | 195 | 198 | 201 | 204 | 205
0.37 |  |  |  | 185 | 186 | 187 | 188 | 189 | 190 | 0.37 |  |  | 189 | 190 | 191 | 191 | 192 | 193 | 193 | 0.37 |  |  | 193 | 196 | 199 | 202 | 205 | 207 | 209
0.38 |  |  |  |  | 190 | 191 | 192 | 193 | 194 | 0.38 |  |  | 193 | 194 | 196 | 196 | 197 | 197 | 197 | 0.38 |  |  | 198 | 201 | 204 | 207 | 209 | 211 | 213
0.39 |  | 174 178: 196; 175 179: 197; 176 180: 198; 177 181: 198; 0.34 0.35: 0.39 |  |  | 198 | 199 | 200 | 200 | 201 | 201 | 202 | 0.39 |  |  | 203 | 206 | 208 | 211 | 214 | 216 | 217
0.4 |  | 175 179: 201; 176 180: 202; 177 181: 203; 0.34 0.35: 0.4 |  |  | 202 | 203 | 204 | 204 | 205 | 205 | 206 | 0.4 |  |  | 208 | 211 | 213 | 216 | 218 | 220 | 221
0.41 |  | 176 180: 206; 177 181: 207; 0.34 0.35: 0.41 |  |  | 206 | 207 | 208 | 208 | 209 | 209 | 210 | 0.41 |  |  | 213 | 215 | 217 | 220 | 222 | 224 | 225
0.42 |  | 177 181: 211; 0.34 0.35: 0.42 |  |  | 211 | 211 | 212 | 212 | 213 | 213 | 214 | 0.42 |  |  | 218 | 220 | 222 | 224 | 226 | 228 | 229
Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast | [12 empty] | Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
| kVp | [12 empty] | kVp
HVL | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | HVL | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39
0.3 | 152 | 157 | 163 | 166 | 170 | 173 | 175 | 177 | 179 | 182 | 184 | 187 | 0.4 | 222 | 226 | 229 | 231 | 234 | 236 | 239 | 241 | 244 | 246 | 248 | 250 | 252
0.325 | 163 | 169 | 174 | 177 | 181 | 183 | 186 | 188 | 190 | 192 | 195 | 197 | 0.425 | 233 | 236 | 239 | 242 | 244 | 246 | 248 | 251 | 253 | 256 | 258 | 260 | 262
0.35 | 175 | 180 | 185 | 188 | 191 | 194 | 196 | 198 | 200 | 202 | 205 | 207 | 0.45 | 244 | 247 | 249 | 252 | 254 | 256 | 258 | 260 | 263 | 265 | 267 | 269 | 271
0.375 | 186 | 191 | 196 | 199 | 202 | 205 | 207 | 209 | 211 | 213 | 215 | 218 | 0.475 | 254 | 257 | 260 | 262 | 264 | 266 | 268 | 270 | 273 | 275 | 277 | 279 | 281
0.4 | 198 | 203 | 207 | 210 | 213 | 215 | 217 | 219 | 221 | 223 | 226 | 228 | 0.5 | 265 | 267 | 270 | 272 | 274 | 276 | 278 | 280 | 282 | 284 | 286 | 288 | 290
0.425 | 209 | 214 | 218 | 221 | 224 | 226 | 228 | 230 | 232 | 234 | 236 | 238 | 0.525 | 275 | 278 | 280 | 282 | 284 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 300
0.45 | 221 | 226 | 230 | 232 | 235 | 237 | 238 | 240 | 242 | 244 | 246 | 248 | 0.55 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 299 | 301 | 303 | 305 | 307 | 309
0.475 | 233 | 237 | 241 | 243 | 245 | 247 | 249 | 251 | 253 | 254 | 256 | 258 | 0.575 | 296 | 298 | 300 | 302 | 304 | 305 | 307 | 309 | 311 | 313 | 315 | 317 | 318
0.5 | 244 | 248 | 252 | 254 | 256 | 258 | 260 | 261 | 263 | 265 | 267 | 269 | 0.6 | 306 | 308 | 310 | 312 | 313 | 315 | 317 | 319 | 320 | 322 | 324 | 326 | 328
0.525 | 256 | 260 | 263 | 265 | 267 | 269 | 270 | 272 | 273 | 275 | 277 | 279 | 0.625 | 316 | 318 | 320 | 322 | 323 | 325 | 326 | 328 | 330 | 332 | 333 | 335 | 337
0.55 | 267 | 271 | 274 | 276 | 278 | 279 | 281 | 282 | 284 | 285 | 287 | 289 | 0.65 | 326 | 328 | 330 | 331 | 333 | 334 | 336 | 338 | 339 | 341 | 343 | 344 | 346
0.575 | 279 | 282 | 285 | 287 | 288 | 290 | 291 | 292 | 294 | 296 | 297 | 299 | 0.675 | 336 | 338 | 339 | 341 | 342 | 344 | 345 | 347 | 349 | 350 | 352 | 354 | 355
0.6 | 290 | 293 | 296 | 297 | 299 | 300 | 301 | 303 | 304 | 306 | 308 | 310 | 0.7 | 346 | 348 | 349 | 350 | 352 | 353 | 355 | 356 | 358 | 359 | 361 | 363 | 364
0.625 | 301 | 304 | 306 | 308 | 309 | 310 | 312 | 313 | 315 | 316 | 318 | 320 | 0.725 | 356 | 357 | 358 | 360 | 361 | 362 | 364 | 365 | 367 | 368 | 370 | 372 | 373
0.65 | 312 | 314 | 317 | 318 | 320 | 321 | 322 | 323 | 325 | 326 | 328 | 330 | 0.75 | 365 | 367 | 368 | 369 | 370 | 372 | 373 | 375 | 376 | 378 | 379 | 381 | 382
0.675 | 322 | 325 | 327 | 328 | 330 | 331 | 333 | 333 | 335 | 336 | 338 | 340 | 0.775 | 374 | 376 | 377 | 378 | 379 | 381 | 382 | 383 | 385 | 386 | 388 | 390 | 391
0.7 | 333 | 335 | 337 | 339 | 340 | 341 | 342 | 343 | 345 | 346 | 348 | 350 | 0.8 | 384 | 385 | 386 | 387 | 388 | 390 | 391 | 392 | 394 | 395 | 397 | 398 | 400
0.725 | 342 | 345 | 347 | 348 | 349 | 351 | 352 | 353 | 354 | 356 | 358 | 360 | 0.825 | 393 | 394 | 395 | 396 | 397 | 399 | 400 | 401 | 403 | 404 | 406 | 407 | 408
0.75 | 352 | 355 | 357 | 358 | 359 | 360 | 361 | 363 | 364 | 366 | 368 | 369 | 0.85 | 402 | 403 | 404 | 405 | 406 | 407 | 409 | 410 | 411 | 413 | 414 | 416 | 417
0.775 | 361 | 365 | 367 | 368 | 369 | 370 | 371 | 372 | 374 | 375 | 377 | 379 | 0.875 | 410 | 411 | 412 | 413 | 415 | 416 | 417 | 418 | 420 | 421 | 423 | 424 | 425
0.8 | 369 | 374 | 376 | 377 | 378 | 379 | 380 | 382 | 383 | 385 | 387 | 389 | 0.9 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 427 | 428 | 429 | 431 | 432 | 434
Siemens Inspiration
| PMMA thickness (mm) | 170 174: Equiv Breast thickness (mm); 171 175: HVL (mm Al)
|  |  | 0.25 | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 | Slope | Intercept | RSQ
g-factor | 45 | 53 | 0.13 | 0.155 | 0.177 | 0.198 | 0.22 | 0.245 | 0.272 | 0.295 | 0.4681 | 0.01256 | 0.99882
c-factor | 45 | 53 |  | 1.109 | 1.105 | 1.102 | 1.099 | 1.096 | 1.091 | 1.088 | -0.069286 | 1.1297 | 0.99503
g-factor | 50 | 60 | 0.112 | 0.135 | 0.154 | 0.172 | 0.192 | 0.214 | 0.236 | 0.261 | 0.41619 | 0.007619 | 0.998
c-factor | 50 | 60 |  | 1.164 | 1.16 | 1.151 | 1.15 | 1.144 | 1.139 | 1.134 | -0.099286 | 1.1935 | 0.98456
| Mo/Mo | 170 174: Mo/Rh; 171 175: W/Rh
s-factor | 1 | 170 174: 1.017; 171 175: 1.042
kV Accuracy/Output |  | 170 174: Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV
Target/Filter | kV | mGy/mAs | Target/Filter | kV | mGy/mAs | Target/Filter | kV | mGy/mAs
/ | 24 | 171 175: /; 172 176: 28; 174 178: /; 175 179: 28
| 25 | 172 176: 30; 175 179: 30
| 26 | 172 176: 32; 175 179: 32
| 28 | 172 176: 34; 175 179: 34
| 30 | 172 176: 36; 175 179: 38
| 32 | 172 176: 38
34
/ | bar: 170 174 to 171 175
Slope: | #DIV/0! | 170 174: #DIV/0!; 171 175: #DIV/0!
Intercept: | #DIV/0! | 170 174: #DIV/0!; 171 175: #DIV/0!
r^2: | #DIV/0! | 170 174: #DIV/0!; 171 175: #DIV/0!
Half Value Layer |  | 170 174: Calculate HVL as a function of kV
Target/Filter | kV | HVL | Target/Filter | kV | HVL | Target/Filter | kV | HVL |  | Target | Filter | C
/ | #DIV/0! | #DIV/0! | / | #DIV/0! | #DIV/0! | / | #DIV/0! | #DIV/0! |  | Mo | Mo | 0.12
| #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | Mo | Rh | 0.19
| 0 | 0 |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | Rh | Rh | 0.22
| #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | Critiera: | Min HVL > kV/100 (no paddle)
| 0 | 170 174: 0; 172 176: 0; 173 177: 0; 175 179: #DIV/0!; 176 180: #DIV/0! | Min HVL > kV/100 + 0.03 (with paddle)
| #DIV/0! | 170 174: #DIV/0!; 172 176: 0; 173 177: 0 | Max HVL < kV/100 + C
| #DIV/0! | 170 174: #DIV/0!
/ | bar: 170 174 to 171 175
Slope: | #DIV/0! | 170 174: #DIV/0!; 171 175: #DIV/0!
Intercept: | #DIV/0! | 170 174: #DIV/0!; 171 175: #DIV/0!
r^2: | #DIV/0! | 170 174: #DIV/0!; 171 175: #DIV/0!
DGN | #N/A
Mo/Mo | Mo/Rh | 170 174: Rh/Rh; 172 176: W/Ag; 173 177: W/Rh
#N/A | #N/A | 170 174: #N/A; 171 175: #N/A; 172 176: #N/A; 173 177: #N/A
Radcal kV Correction for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Rh | < | 27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| >= | 27.586 | -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Ag | < | 30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| >= | 30.1 | -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
Radcal kV Correction for Hologic Dimension
| kV | 171 175: kV^0; 172 176: kV^1; 173 177: kV^2; 174 178: kV^3; 175 179: kV^4
W/Rh | < | 170 174: 26.9; 171 175: 138.8867; 172 176: -10.7264; 173 177: 0.26216; 174 178: -0.00082
| >= | 26.9 | -5009.7752 | 605.732 | -27.3019 | 0.54671 | -0.0040986
W/Ag | kV | 171 175: kV^0; 172 176: kV^1; 173 177: kV^2; 174 178: kV^3; 175 179: kV^4
| < | 170 174: 28.7; 171 175: 296.3419; 172 176: -31.6292; 173 177: 1.1803; 174 178: -0.01417
| >= | 170 174: 28.7; 171 175: 4.8345; 172 176: 0.91924
W/Al |  | 171 175: kV^0; 172 176: kV^1; 173 177: kV^2; 174 178: kV^3; 175 179: kV^4
| < | 170 174: 28.7; 171 175: 49.3111; 172 176: -2.9302; 173 177: 0.07379
|  | 28.7 < kV < 30.1 | -24.875 | 1.8032
| >= | 170 174: 30.1; 171 175: -4.8346; 172 176: 1.1571
Advanced Settings
|  | Grid | Auto Decomp | OPCOMP | AEC Segm | Dose Level
| Phantom Image Quality | 170 174: In; 171 175: Off; 172 176: On; 173 177: Off; 174 178: Normal
| Artifact Detection | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| SNR/CNR | 170 174: In; 171 175: Off; 172 176: On; 173 177: Off; 174 178: Normal
| SNR/CNR Repeatability | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| Spatial Resolution | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| AEC | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| Detector Uniformity | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| Collimation | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| MGD | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| HVL | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
| kV | 170 174: In; 171 175: Off; 172 176: Off; 173 177: Off; 174 178: Normal
Note on L56: Eugene Mah: Linear fit for this one is not good
Note on A91: Radcal kV correction for W anodes uses a 5th order polynomial divided into two domains.
Note on A103: Radcal kV correction for W anodes uses a 5th order polynomial divided into two domains.

## DataPage
Sheet | MAMMO_SIE
SurveyID

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
Mammomat Fusion
Mammomat Inspiration
Mammomat Inspiration Tomosynthesis
Mammomat Novation
Mammomat Revelation

## Corrected kV
GENERAL
Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
Read the true corrected kV in the adjacent cell to the right.
If that cell is light green, the corrected kV is accurate.
If the cell is red, the entered kV is out of range and the correction should not be used.
If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.
Tungsten-Rhodium (W / Rh)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| Radcal kV | True kV
| 25.03 | 21.8081
Tungsten-Silver (W / Ag)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
| Radcal kV | True kV
| 37.1 | 38.9675
27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0