## Sheet2
Medical University of South Carolina
Charleston, South Carolina
[5 empty] | Radiation Shielding | [5 empty] | Conversion Assistance
[11 empty] | Feet | Meter | Meter | Feet
[5 empty] | Location Information | [6 empty] | 0 |  | 0
[11 empty] | inch | mm | mm | inch
Location | [7 empty] | Workload Information |  |  |  | 0 |  | 0
| Date: | [6 empty] | # Patients/week:
| Facility: | [6 empty] | Workload/patient (mA-min/pt):
| Department: | [6 empty] | Views/patient:
| Area/Division | [6 empty] | Weekly workload (mA-min):
| Type of Room: | [6 empty] | Primary field size (cm^2):
Performed by:
Unshielded primary air kerma per patient (Kp)
Unshielded secondary air kerma per patient (Ksec)
Fit parameters
|  | Attenuator | a | b | g | ag | b/a
Lead
Concrete
Gypsum
Steel
Glass
Wood
|  | All distances in meters | [6 empty] | P | Transmission
| Barrier | D(pri) | D(sec) | D(sca) | D(leak) | U | T | Area | (mSv/wk) | B(x)
Note on K29: Eugene Mah: Dose limits per NCRP 116 recommendations.
Note on D30: Eugene Mah: Distance from primary beam (m)
Note on H30: Eugene Mah: Use factor
Note on I30: Eugene Mah: Occupancy factor
Note on J30: Eugene Mah: Type of area U - Uncontrolled C - Controlled C if blank

## Sheet1
[6 empty] | P |  |  | N | 100 | pt
Barrier | D(sec) | D(sca) | D(leak) | T | Area | (mSv/wk) | B |  | Ksec | 0.036 | mGy/pt
Wall A1 | 3.1 | 3.1 | 3.1 | 0.2 | U | 0.02 | 0.044491 |  | W | 1333 | mA-min
Wall A2 | 3.8 | 3.8 | 3.8 | 0.2 | U | 0.02 | 0.066852 |  | Views/pt | 6
Door A | 3.1 | 3.1 | 3.1 | 0.2 | U | 0.02 | 0.044491
Wall B | 2.4 | 2.4 | 2.4 | 1 | U | 0.02 | 0.0053333 |  | Attenuator | a | b | g | ag | b/a
Wall C | 1.8 | 1.8 | 1.8 | 0.1 | U | 0.02 | 0.03 |  | Lead | 28.41 | 62.94 | 0.3201 | 9.094 | 2.2154
Wall D | 4.3 | 4.3 | 4.3 | 1 | U | 0.02 | 0.01712 |  | Concrete | 0.1541 | 0.5974 | 0.309 | 0.047617 | 3.8767
Floor | 3 | 3 | 3 | 1 | U | 0.02 | 0.0083333 |  | Gypsum | 0.06557 | 0.3031 | 0.2593 | 0.017002 | 4.6225
Ceiling | 3 | 3 | 3 | 1 | U | 0.02 | 0.0083333 |  | Steel | 5.402 | 16.81 | 0.5086 | 2.7475 | 3.1118
Control | 3.2 | 3.2 | 3.2 | 1 | C | 0.1 | 0.047407 |  | Glass | 0.1514 | 0.6343 | 0.2725 | 0.041257 | 4.1896
[9 empty] | Wood | 0.01186 | 0.01312 | 0.1603 | 0.0019012 | 1.1062
Barrier | Lead | Concrete | Gypsum | Steel | Glass | Wood
Wall A1 | 0.046854 | 6.0116 | 11.7329 | 0.2414 | 5.55 | 140.8924
Wall A2 | 0.039203 | 4.9865 | 9.7762 | 0.1973 | 4.6213 | 120.6244 |  | 31.75
Door A | 0.046854 | 6.0116 | 11.7329 | 0.2414 | 5.55 | 140.8924
Wall B | 0.09395 | 12.6712 | 24.3595 | 0.52581 | 11.5352 | 255.0048
Wall C | 0.0547 | 7.0798 | 13.7637 | 0.28762 | 6.514 | 161.0205
Wall D | 0.066599 | 8.7322 | 16.8961 | 0.35904 | 8.0005 | 190.5269
Floor | 0.083084 | 11.0837 | 21.3487 | 0.45938 | 10.1109 | 229.8431
Ceiling | 0.083084 | 11.0837 | 21.3487 | 0.45938 | 10.1109 | 229.8431
Control | 0.045631 | 5.8465 | 11.4184 | 0.23428 | 5.4007 | 137.6962