## Sheet1
[7 empty] | Revision 1.0-20230427 | [10 empty] | Print Area
|  |  | Medical University of South Carolina | [8 empty] | Medical University of South Carolina
|  |  | Charleston, South Carolina | [8 empty] | Charleston, South Carolina
[18 empty] | All: | Page1,Page2,HVLPage,KAP_AKR,ExpChartFluoroPage,ExpChartDAPage,ImgQualityPage,Com1Page,Com2Page,OutputGraphFluoroPage,OutputGraphDAPage,LeedsTO10Page
|  |  | O-Arm System Compliance Inspection | [8 empty] | O-Arm System Compliance Inspection
[8 empty] | Date: | [6 empty] | Inspector: |  |  | Measurement Parameter | Last Year |  | This Year
[8 empty] | Previous Date: | [9 empty] | Print Area |  |  |  | P2
| Location | [10 empty] | System Information | [5 empty] | Date |  |  |  | D6
| Facility: |  |  | Site Number: |  |  |  | Location |  | Input Changes Only |  |  |  |  | Input Changes Only |  |  | Inspector |  |  |  | L6
| Department: |  |  | Authorized Use: |  |  |  | Facility: |  |  |  |  | Site Number: |  |  |  |  | Facility: |  |  |  | E10
| Area/Division: |  |  | Date of Installation: |  |  |  | Department: |  |  |  |  | Authorized Use: |  |  |  |  | Department: |  |  |  | E11
| Survey ID: |  |  | Accession Number: |  |  |  | Area/Division: |  |  |  |  | Date of Installation: |  |  |  |  | Area/Division: |  |  |  | E12
| Room Number: | [6 empty] | Survey ID: |  |  |  |  | Accession Number: |  |  |  |  | Survey ID: |  |  |  | E13
| X-Ray Generator | [6 empty] | Room Number: | [9 empty] | Room Number: |  |  |  | E14
| Manufacturer: |  |  | Manufacture Date: |  |  |  | X-Ray Generator | [9 empty] | Site Number: |  |  |  | J10
| Model: |  |  | Serial Number: |  |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | Authorized Use: |  |  |  | J11
| Max kVp: |  |  | Max mA: |  |  |  | Model: |  |  |  |  | Serial Number: |  |  |  |  | Date of Installation: |  |  |  | J12
|  |  |  | Software Version: |  |  |  | Max kVp: |  |  |  |  | Max mA: |  |  |  |  | Accesssion Number: |  |  |  | J13
| X-Ray Tube | [11 empty] | Software Version: |  |  |  |  | X-Ray Generator
| Tube Designation/Use: |  |  | Focal Spot Sizes (mm) |  |  |  | X-Ray Tube | [9 empty] | Manufacturer: |  |  |  | E16
| Manufacture Date: |  |  | Large: |  |  |  | Tube Designation/Use: |  |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Model: |  |  |  | E17
| Insert |  |  | Small: |  |  |  | Manufacture Date: |  |  |  |  | Large: |  |  |  |  | Max kVp: |  |  |  | E18
| Manufacturer: |  |  | Micro: |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Manufacture Date: |  |  |  | J16
| Model: | [6 empty] | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Serial Number: |  |  |  | J17
| Serial Number: |  |  | Filtration |  |  |  | Model: | [9 empty] | Max mA: |  |  |  | J18
| Housing |  |  | Inherent: |  |  |  | Serial Number: |  |  |  |  | Filtration |  |  |  |  | Software Version: |  |  |  | J19
| Manufacturer: |  |  | Added: |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
| Model: | [6 empty] | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Tube Designation/Use: |  |  |  | E21
| Serial Number: | [6 empty] | Model: | [9 empty] | Manufacture Date: |  |  |  | E22
[8 empty] | Serial Number: | [9 empty] | Manufacturer: |  |  |  | E24
[18 empty] | Model: |  |  |  | E25
Labels, Notices, Postings | [5 empty] | Compliance | [11 empty] | Serial Number: |  |  |  | E26
DHEC Registration sticker is present, clearly visible and legible | [5 empty] | TBD | [11 empty] | Manufacturer: |  |  |  | E28
DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [5 empty] | TBD | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Model: |  |  |  | E29
Radiation warning label posted on the generator control panel | [5 empty] | TBD |  |  | Is this a new installation? |  |  |  |  | Tube Number: |  |  |  | Serial Number: |  |  |  | E30
Tube loading charts for each focal spot are posted | [5 empty] | TBD | [8 empty] | **Input only one of the following |  |  | Large: |  |  |  | J22
Tube loading charts are correct for installed tube(s). | [5 empty] | TBD | Labels, Notices, Postings | [7 empty] | numbers or names |  |  | Small: |  |  |  | J23
Operator manuals are available. | [5 empty] | TBD |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | 1 | 1 |  | Micro: |  |  |  | J24
Monthly radiation monitoring reports are posted. | [5 empty] | TBD |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | 2 | 2 |  | Inherent: |  |  |  | J27
[8 empty] | Radiation warning label posted on the generator control panel | [6 empty] | 3 | 3 |  | Added: |  |  |  | J28
[8 empty] | Tube loading charts for each focal spot are posted | [6 empty] | AP | AP |  | New install: |  |  |  | C35
Date: |  |  |  |  | Inspector: |  |  | Tube loading charts are correct for installed tube(s). | [6 empty] | Front | Front |  | Tube Number: |  |  |  | K35
Room: |  |  |  |  | Survey ID: |  |  | Operator manuals are available. | [6 empty] | Frontal | Frontal |  | Tube Designation: |  | Change | 1 | K38
[6 empty] | Medical University of South Carolina |  | Monthly radiation monitoring reports are posted. | [6 empty] | Lat | Lat |  |  |  | Change | 2 | K39
[6 empty] | O-Arm System Compliance Inspection | [8 empty] | Lateral | Lateral |  |  |  | Change | 3 | K40
Fluoroscopy Systems | [6 empty] | Fluoroscopy Systems | [12 empty] | Change | AP | K41
Interlock systems function to prevent fluoro while II is parked | [5 empty] | TBD |  | Interlock systems function to prevent fluoro while II is parked | [11 empty] | Change | Front | K42
The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [5 empty] | TBD |  | The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [11 empty] | Change | Frontal | K43
Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels) | [5 empty] | TBD |  | Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels) | [11 empty] | Change | Lat | K44
The useful beam is intercepted by the image assembly at any SID | [5 empty] | TBD |  | The useful beam is intercepted by the image assembly at any SID | [11 empty] | Change | Lateral | K45
Fluoroscopy patient dose charts are posted | [5 empty] | TBD |  | Fluoroscopy patient dose charts are posted | [9 empty] | Receptor Size
Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [5 empty] | TBD |  | Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [9 empty] | Normal:
Max Entrance Exposure Rate <= 176 mGy/min (high level) | [5 empty] | TBD |  | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [9 empty] | Mag 1:
High contrast resolution within MUSC specifications | [5 empty] | TBD |  | High contrast resolution within MUSC specifications | [9 empty] | Mag 2:
Low contrast resolution within MUSC specifications | [5 empty] | TBD |  | Low contrast resolution within MUSC specifications | [9 empty] | Pt Dose Measured Fluoro
Transmission through II < 2 uGy/hr for each mGy/min at patient entrance | [5 empty] | TBD |  | Transmission through II < 2 uGy/hr for each mGy/min at patient entrance | [9 empty] | Manual/Auto:
Beam limiting device available (functioning) to restrict the x-ray field | [5 empty] | TBD |  | Beam limiting device available (functioning) to restrict the x-ray field | [9 empty] | Lucite/Al:
The useful beam is centered on the image receptor | [5 empty] | TBD |  | The useful beam is centered on the image receptor | [9 empty] | # dose rates:
Collimator edges are visible at full open | [5 empty] | TBD |  | Collimator edges are visible at full open | [9 empty] | High level present:
Elapsed timer (5 min maximum) functional | [5 empty] | TBD |  | Elapsed timer (5 min maximum) functional | [9 empty] | cm/in:
Fluoroscopic kV and mA meter present and functional | [5 empty] | TBD |  | Fluoroscopic kV and mA meter present and functional | [9 empty] | Auto/mA:
Radiation Safety | [17 empty] | Test equipment:
Continuous pressure on exposure switch required to make an exposure | [5 empty] | TBD | Radiation Safety | [10 empty] | Calibration date:
Operator has full visibility of the exposure factors and patient | [5 empty] | TBD |  | Continuous pressure on exposure switch required to make an exposure with selected tube | [9 empty] | Calibration due:
Lead aprons available | [5 empty] | TBD |  | Operator has full visibility of the exposure factors and patient | [9 empty] | Technique 1:
Lead gloves available | [5 empty] | TBD |  | Lead aprons available | [9 empty] | Technique 2:
Properly designed and installed apron rack is present | [5 empty] | TBD |  | Lead gloves available | [9 empty] | Technique 3:
Documentation of annual protective apparel integrity inspection is available | [5 empty] | TBD |  | Properly designed and installed apron rack is present | [9 empty] | Maximum Dose Rates Fluoro
Patient restraint devices available | [5 empty] | TBD |  | Documentation of annual protective apparel integrity inspection is available | [9 empty] | kV: |  |  | 0
Gonadal shielding is present and functional | [5 empty] | TBD |  | Patient restraint devices available | [9 empty] | mA: |  |  | 0
Physical Inspection | [7 empty] | Gonadal shielding is present and functional | [9 empty] | mGy/min: |  |  | 0
System clean and free of debris | [5 empty] | TBD | [11 empty] | kV: |  |  | 0
All cables free from excessive wear or restraints | [5 empty] | TBD | Physical Inspection | [10 empty] | mA: |  |  | 0
System free of missing or damaged components | [5 empty] | TBD |  | System clean and free of debris | [9 empty] | mGy/min: |  |  | 0
Tube support movements and locks are acceptable | [5 empty] | TBD |  | All cables free from excessive wear or restraints | [9 empty] | kV: |  |  | 0
All other movements and locks functional | [5 empty] | TBD |  | System free of missing or damaged components | [9 empty] | mA: |  |  | 0
Unit installed as shown on shielding plan | [5 empty] | TBD |  | Tube support movements and locks are acceptable | [9 empty] | mGy/min: |  |  | 0
Automatic Brightness Control / Display Monitor | [7 empty] | All other movements and locks functional | [9 empty] | HVL
Automatic Brightness Control maintains image brightness for varying attenuation | [5 empty] | TBD |  | Unit installed as shown on shielding plan | [9 empty] | HVL Measured kV: |  | Change | 80
Display monitor is free of artifacts, shadows, bright spots, and burn-in | [5 empty] | TBD | [11 empty] | HVL Measured: |  | Change | TBD
Display monitor is free of distortion | [5 empty] | TBD | Automatic Brightness Control / Display Monitor | [10 empty] | Fluoro Scatter/Transmission
Display monitor is free of excessive persistence (after-image) | [5 empty] | TBD |  | Automatic Brightness Control maintains image brightness for varying attenuation | [9 empty] | Transmission: |  |  | 0
Display monitor is free of static or rolling noise bands | [5 empty] | TBD |  | Display monitor is free of artifacts, shadows, bright spots, and burn-in | [9 empty] | Unshielded Scatter (eyes):
The last image hold (LIH) display is free of jitter or abnormal appearance | [5 empty] | TBD |  | Display monitor is free of distortion | [9 empty] | Unshielded Scatter (waist):
Date: |  |  |  |  | Inspector: |  |  | Display monitor is free of excessive persistence (after-image) | [9 empty] | Shielded Scatter (eyes):
Room: |  |  |  |  | Survey ID: |  |  | Display monitor is free of static or rolling noise bands | [9 empty] | Shielded Scatter (waist):
[6 empty] | Medical University of South Carolina |  | The last image hold (LIH) display is free of jitter or abnormal appearance | [9 empty] | Receptor Entr Exp Fluoro
[6 empty] | O-Arm System Compliance Inspection | [11 empty] | Normal:
Patient Entrance Exposure Rate (Fluoroscopy) | [17 empty] | Mag 1:
Image | Acrylic | [5 empty] | Entrance Exposure Rate | [10 empty] | Mag 2:
Mode | Attenuator | [7 empty] | Type 1 for manual, leave blank for auto | [6 empty] | Auto/mA mode |  | Normal:
(cm) | (cm) | kVp | mA | mGy/min |  |  |  |  | Type 1 if using Lucite, leave blank if using Al | [6 empty] | Test equipment |  | Mag 1:
cm | 10 | 0 | 0 | 0 |  |  |  |  | How many dose rate settings are present (exclusing High Level)? | [6 empty] | Calibration date |  | Mag 2:
| 20 | 0 | 0 | 0 |  |  |  |  | Is High Level provided? (1=Yes, 2=No) | [6 empty] | Calibration due |  | Normal:
| 30 | 0 | 0 | 0 |  |  | cm | cm | Unit for image receptor size | [8 empty] | Mag 1:
cm | 10 | 0 | 0 | 0 |  |  | Blank | [10 empty] | Mag 2:
| 20 | 0 | 0 | 0 |  |  | Patient Entrance Exposure Rate (Fluoroscopy) | [10 empty] | Leeds Phantom
| 30 | 0 | 0 | 0 |  |  | Image | Acrylic | [9 empty] | kVp: |  | Change | 70
cm | 10 | 0 | 0 | 0 |  |  | Mode | Attenuator | [9 empty] | Huttner Normal:
| 20 | 0 | 0 | 0 |  |  | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | Huttner Mag 1:
| 30 | 0 | 0 | 0 |  |  | Normal | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Huttner Mag 2:
| Max | 0 | 0 | 0 |  |  |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N3 Normal:
[8 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N3 Mag 1:
[7 empty] | Mag 1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N3 Mag 2
[8 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Field Shape:
[8 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SID:
[7 empty] | Mag 2 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Front:
[8 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Back:
[8 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Left:
[8 empty] | Max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Right:
[8 empty] | Last Year: | [9 empty] | TO.10.A
[8 empty] | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES | TO.10.B
[8 empty] | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | TO.10.C
[8 empty] | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point) | [9 empty] | TO.10.D
[8 empty] | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [9 empty] | TO.10.E
[8 empty] | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point) | [9 empty] | TO.10.F
[8 empty] | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [8 empty] | TO.10.G
[9 empty] | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [8 empty] | TO.10.H
[8 empty] | Comments: | [9 empty] | TO.10.J
[8 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | TO.10.K
[8 empty] | New: | [9 empty] | TO.10.L
TO.10.M
[7 empty] | Half Value Layer | [10 empty] | TO.10.A
[7 empty] | 2 | Use DHEC standards? (Type 1 to base acceptance on DHEC criteria) | [9 empty] | TO.10.B
TO.10.C
[7 empty] | kV | mA | mm Al | mGy/min |  | Data Processing | [5 empty] | TO.10.D
[7 empty] | 80 |  | 0 |  |  | mm Al | Avg Exp | DATA(1/2) | Data Al | HVL |  | TO.10.E
Date: |  |  |  |  | Inspector: |  |  |  | 0 |  |  | 0 |  |  |  | TBD |  | TO.10.F
Room: |  |  |  |  | Survey ID: | [12 empty] | TO.10.G
[6 empty] | Medical University of South Carolina | [11 empty] | TO.10.H
[6 empty] | O-Arm System Compliance Inspection | [8 empty] | Data Exp | LN(EXP) |  | TO.10.J
[12 empty] | TBD | #NAME? |  |  |  |  | TO.10.K
[12 empty] | Acceptable (mm Al) | [5 empty] | TO.10.L
[12 empty] | Low | 2.6 |  |  |  |  | TO.10.M
[12 empty] | Desired | 3.1 |  | HVL Acceptable: | TBD |  | TO.10.A
[12 empty] | High | 3.5 |  |  |  |  | TO.10.B
[15 empty] | Previous HVL @ kV: |  |  | TO.10.C
TO.10.D
[7 empty] | Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate | [10 empty] | TO.10.E
[8 empty] | Source-PDC distance (cm) |  |  |  |  | Source-Isocenter |  |  | Test equipment |  | TO.10.F
[8 empty] | Source-Ref Pt distance (cm) |  |  |  |  | Table-Isocenter |  |  | Calibration date |  | TO.10.G
[8 empty] | Reference Point correction factor |  |  |  |  | RefPt-Isocenter |  |  | Calibration due |  | TO.10.H
TO.10.J
[9 empty] | Start KAP | End KAP | Measured | Meas KAP | Indicated |  |  |  |  | TO.10.K
[7 empty] | kV | mA | µGy·m² |  | KAP | @ Ref Pt | KAP | Deviation |  |  |  | TO.10.L
TO.10.M
Comments
Pt Entr Exp Fluoro:
Receptor Entr Exp Fluoro:
Scatter/Transmission:
[9 empty] | Start AK | End AK | Measured | Meas CAK | Indicated |  |  |  |  | Leeds Phantom:
[7 empty] | kV | mA | mGy |  | CAK | @ Ref Pt | CAK | Deviation |  |  |  | TO.10:
Beam Limitation:
CBCT
CTDI:
[9 empty] | Measured | Meas AKR | Indicated | [6 empty] | kV:
[7 empty] | kV | mA | AKR | @ Ref Pt | AKR | Deviation | [5 empty] | mA:
f(0):
CBCT comments:
Image Quality
Phantom serial number:
CT Number
kV:
[7 empty] | Criteria: | Measured KAP/AK/AKR is within 40% of the displayed values | [9 empty] | mAs:
Polyethylene:
[7 empty] | Scatter and Transmission | [10 empty] | Water:
[11 empty] | Pt Entrance | [6 empty] | Water SD:
[9 empty] | kVp | mA | mGy/min | uGy/hr | Last Year |  | Acceptable |  |  | Acrylic:
Date: |  |  |  |  | Inspector: |  |  | Transmission* @ 10cm |  |  |  | 0 |  |  | TBD |  |  | Bone:
Room: |  |  |  |  | Survey ID: | [5 empty] | Unshielded | Shielded | Unshielded | Shielded | Unshielded | Shielded |  | Air:
[6 empty] | Medical University of South Carolina |  |  |  |  | uGy/min | uGy/min | uGy/hr | uGy/hr | uGy/hr | uGy/hr |  | Low Contrast
[6 empty] | O-Arm System Compliance Inspection | Scatter – Fluoro | Eye level | [9 empty] | Contrast:
[8 empty] | Waist level | [9 empty] | CNR:
[8 empty] | *Measured with 8” Lucite in beam | [9 empty] | Spatial Resolution:
Largest field size mode, highest dose rate mode
[8 empty] | Criteria: | Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure | [8 empty] | Comment Page:
Comments:
[8 empty] | Do not exceed 390 chars |  | 0
New:
Image Receptor Entrance Exposure Rate – Fluoroscopy
[7 empty] | Dose | Receptor |  |  |  |  | Previous Year
[7 empty] | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | nGy/frame
Measure using pancake probe placed against
image receptor input window
Use 1.5mm Cu filtration
Comments:
[8 empty] | Do not exceed 390 chars |  | 0
New:
kV Accuracy and Reproducibility
[8 empty] | Test Equipment Used: | Piranha |  |  | Calibration Date:
[8 empty] | Serial Number: | CB2-17090320 |  |  | Calibration Due:
[7 empty] | kV Accuracy |  |  | Exposure | Exp rate
[7 empty] | kV set | mA Set | kV | (mGy) | (mGy/min) | HVL | Tot Filt | kV % error
[7 empty] | 60 | 5
[7 empty] | 80 | 5
[7 empty] | 100 | 5
[7 empty] | 120 | 5
[8 empty] | Criteria: | kVp accurate to within 5% of indicated. |  |  |  | Acceptable:
kV/exposure reproducibility
[9 empty] | Exposure | Exp rate |  |  |  | Air Kerma Rate Linearity
Date: |  |  |  |  | Inspector: |  | kV Set: | kV | (mGy) | (mGy/min) | HVL | Tot Filt |  | kV Set: | mA | mGy/min
Room: |  |  |  |  | Survey ID: |  | 80 | [6 empty] | 80 | 1
[6 empty] | Medical University of South Carolina | [8 empty] | 10
[6 empty] | O-Arm System Compliance Inspection | [8 empty] | 7.5
5
Coeff of Linearity:
[7 empty] | Mean: | [7 empty] | Acceptable:
[7 empty] | Std Dev: | [7 empty] | Criteria: | Coefficient of linearity <= 10%.
Coeff of Var:
Acceptable:
[8 empty] | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%.
Leeds Phantom
[7 empty] | Receptor Size | MS1 | MS3 | MS4 | Huttner | N3 | kVp: | 70 | 70
[14 empty] | MS1-4 |  | Huttner: Enter col.row
[14 empty] | 1 | Clearly visible | N3: Enter # of visible disks
[14 empty] | 2 | Visible
[10 empty] | Acceptable: |  |  |  | 3 | Barely visible
[14 empty] | 4 | Not visible
Comments:
[8 empty] | Do not exceed 390 chars |  | 0
New:
[7 empty] | TO.10 |  |  |  |  | Field Size
[7 empty] | Row | [6 empty] | Type 1 for circular field, 2 for square field
A
B
[7 empty] | C |  |  |  |  | Horizontal:
[7 empty] | D |  |  |  |  | Vertical:
[7 empty] | E |  |  |  |  | Diagonal:
[7 empty] | F |  |  |  |  | Acceptable: | NA | NA | NA
[7 empty] | G | [5 empty] | Criteria: | Measured field size is at least 85% of nominal field size
H
J
K
L
M
Comments:
[8 empty] | Do not exceed 390 chars |  | 0
New:
Useful Beam Limitation
Date: |  |  |  |  | Inspector: |  |  |  | SID (cm)
Room: |  |  |  |  | Survey ID: |  |  | Measured | Difference | Sum | Last Year |  |  | T | B | L | R
[6 empty] | Medical University of South Carolina | Top: | [5 empty] | Visible:
[6 empty] | O-Arm System Compliance Inspection | Bottom: | [5 empty] | X-ray:
[7 empty] | Left: |  |  |  |  | Acceptable | SID:
[7 empty] | Right: | [5 empty] | SOD:
[7 empty] | Criteria: | Difference between x-ray field and visible image area is less than 2% SID |  |  |  |  | SID/SOD:
Sum of length and width differences is less than 4% SID
Comments:
[8 empty] | Do not exceed 390 chars |  | 0
New:
CBCT
CTDI:
kV:
mA:
Planar Average Dose Profile
[9 empty] | Center | Top | Bottom | Left | Right
Average:
SD:
f(0):
Previous f(0):
Variation from previous:
Variation from indicated:
Comments:
[8 empty] | Do not exceed 390 chars |  | 0
New:
[7 empty] | Image Quality |  |  |  | Phantom serial number:
[7 empty] | CT Number |  |  |  | Low Contrast
[7 empty] | kV: |  |  |  | Inside:
[7 empty] | mAs: |  |  |  | Adjacent:
[7 empty] | Polyethylene: |  |  |  | Adj SD:
[7 empty] | Water: |  |  |  | Contrast:
[7 empty] | Water Std Dev: |  |  |  | CNR:
Acrylic:
Bone:
Air:
Date: |  |  |  |  | Inspector:
Room: |  |  |  |  | Survey ID: |  | Uniformity
[8 empty] | Mean | SD | Diff (center) | Diff (Mean) | Acceptable
Center:
Top:
Bottom:
Left:
Right:
High Contrast Resolution
[9 empty] | lp/cm |  | Acceptable:
Comments Page 1
Previous Year Comments
Comments:
[8 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[8 empty] | Do not exceed 390 chars |  | 0

## Tables
Leeds MS
1 | Clearly visible
2 | Visible
3 | Barely visible
4 | Not visible
Huttner Grating |  |  | Leeds N3 |  | 65 | 70 | 75 | 80 | Nominal
1.1 | 0.5 |  | 1 | 0.167 | 0.188 | 0.167 | 0.149 | 0.135 | 0.16
1.2 | 0.56 |  | 2 | 0.148 | 0.167 | 0.148 | 0.132 | 0.12 | 0.145
1.3 | 0.63 |  | 3 | 0.128 | 0.145 | 0.128 | 0.114 | 0.104 | 0.123
2.1 | 0.71 |  | 4 | 0.109 | 0.124 | 0.109 | 0.097 | 0.088 | 0.108
2.2 | 0.8 |  | 5 | 0.0876 | 0.0993 | 0.0876 | 0.078 | 0.0705 | 0.086
2.3 | 0.9 |  | 6 | 0.0749 | 0.0847 | 0.0749 | 0.067 | 0.0609 | 0.076
3.1 | 1 |  | 7 | 0.0674 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | 0.066
3.2 | 1.12 |  | 8 | 0.0525 | 0.0593 | 0.0525 | 0.047 | 0.0427 | 0.055
3.3 | 1.25 |  | 9 | 0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | 0.045
4.1 | 1.4 |  | 10 | 0.0371 | 0.04 | 0.0371 | 0.0347 | 0.0329 | 0.039
4.2 | 1.6 |  | 11 | 0.0322 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | 0.033
4.3 | 1.8 |  | 12 | 0.0246 | 0.0269 | 0.0246 | 0.0233 | 0.022 | 0.027
5.1 | 2 |  | 13 | 0.0215 | 0.0232 | 0.0215 | 0.0201 | 0.019 | 0.023
5.2 | 2.24 |  | 14 | 0.0172 | 0.0186 | 0.0172 | 0.0161 | 0.0153 | 0.018
5.3 | 2.5 |  | 15 | 0.0155 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | 0.016
6.1 | 2.8 |  | 16 | 0.013 | 0.014 | 0.013 | 0.0122 | 0.0115 | 0.013
6.2 | 3.15 |  | 17 | 0.011 | 0.0119 | 0.011 | 0.0103 | 0.0097 | 0.011
6.3 | 3.55 |  | 18 | 0.0086 | 0.0093 | 0.0086 | 0.0081 | 0.0077 | 0.009
7.1 | 4 |  | 19 | 0.0066 | 0.0071 | 0.0066 | 0.0062 | 0.0059 | 0.007
7.2 | 4.5
7.3 | 5
TO.10 kV Adj | 65 | 70 | 75 | 80
0.012 | 0.0119 | 0.011 | 0.0103 | 0.0097 |  |  |  |  | .
0.016 | 0.0168 | 0.0155 | 0.0145 | 0.0138
0.023 | 0.0232 | 0.0215 | 0.0201 | 0.019
0.032 | 0.0347 | 0.0322 | 0.0301 | 0.0285
0.045 | 0.0512 | 0.045 | 0.0399 | 0.036
0.066 | 0.0766 | 0.0674 | 0.0599 | 0.0542
0.086 | 0.0993 | 0.0876 | 0.078 | 0.0705
0.123 | 0.145 | 0.128 | 0.114 | 0.104
0.16 | 0.188 | 0.167 | 0.149 | 0.135
0.23 | 0.267 | 0.238 | 0.214 | 0.194
0.35 | 0.399 | 0.36 | 0.327 | 0.299
0.5 | 0.618 | 0.573 | 0.533 | 0.497
0.66 | 0.769 | 0.726 | 0.686 | 0.649
0.93 | 0.969 | 0.954 | 0.936 | 0.915
TO.10 | Diameter (mm) | sqrt(Area) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 11.1 | 9.8371 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
B | 7.9 | 7.0012 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
C | 5.6 | 4.9629 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
D | 4 | 3.5449 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
E | 2.8 | 2.4814 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
F | 2 | 1.7725 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
G | 1.4 | 1.2407 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
H | 1 | 0.88623 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
J | 0.7 | 0.62036 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
K | 0.5 | 0.44311 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
L | 0.35 | 0.31018 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
M | 0.25 | 0.22156 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
TO.10 Values | 0 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9
A | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
B | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
C | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
D | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
E | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
F | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
G | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
H | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
J | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
K | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
L | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
M | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
Half Value Layer
| DHEC |  | 0.915 2: MUSC; 0.915 6: HVL Graph
| Min HVL |  | 0.915 2: HVL; 0.915 3: Min HVL; 0.915 4: Max HVL; 0.915 6: mm Al; 0.915 7: Avg Exp
kVp | (mm Al) |  | 0.936 1: kVp; 0.915 2: (mm Al); 0.915 3: (mm Al); 0.915 4: (mm Al)
30 | 0.3 |  | 0.936 1: 60; 0.915 2: 2.2; 0.915 3: 1.8; 0.915 4: 2.4
40 | 0.4 |  | 0.936 1: 70; 0.915 2: 2.6; 0.915 3: 2.2; 0.915 4: 2.9
50 | 0.5 |  | 0.936 1: 80; 0.915 2: 3; 0.915 3: 2.6; 0.915 4: 3.5
51 | 1.2 |  | 0.936 1: 90; 0.915 2: 3.5; 0.915 3: 3.1; 0.915 4: 4
60 | 1.3 |  | 0.936 1: 100; 0.915 2: 3.9; 0.915 3: 3.5; 0.915 4: 4.5
70 | 1.5 |  | 0.936 1: 110; 0.915 2: 4.3; 0.915 3: 3.9; 0.915 4: 5.1
71 | 2.1 |  | 0.936 1: 120; 0.915 2: 5.7; 0.915 3: 4.4; 0.915 4: 5.6
80 | 2.3 |  | 0.936 1: 130; 0.915 2: 5.2; 0.915 3: 4.8; 0.915 4: 6.1
90 | 2.5
100 | 2.7
110 | 3
120 | 3.2
130 | 3.5
140 | 3.8
150 | 4.1
Note on A30: Eugene Mah: Nominal contrast values as a function of kVp
Note on A46: Eugene Mah: TO.10 nominal contrast values
Note on A60: Eugene Mah: TO.10 contrast values at the kVp specified in Sheet1!I229

## Data Entry
Patient Entrance Exposure Rate (Fluoroscopy)* | [10 empty] | Measurements | [6 empty] | Measurements
Image | Acrylic
Mode | Attenuator | [9 empty] | Set kV (kV) | Set mA (mA) | Tube voltage (kV) | Exposure (mGy) | Exposure rate (mGy/min) | HVL (mm Al) | Total filtr. (mm Al) | Set kV (kV) | Set mA (mA) | Tube voltage (kV) | Exposure (mGy) | Exposure rate (mGy/min) | HVL (mm Al) | Total filtr. (mm Al) | Set frames/s (FPS) | Frames/s (FPS) | Exposure/frame (µGy/frame) | Frame exp. rate (mGy/s) | Pulse width (ms)
(cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | 60 | 5 | [5 empty] | 60 | 5
Normal | 10 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 20 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 30 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
Mag 1 | 10 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 20 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 30 | [9 empty] | 100 | 5 | [5 empty] | 100 | 5
Mag 2 | 10 | [9 empty] | 120 | 5 | [5 empty] | 120 | 5
| 20 | [9 empty] | 80 | 1 | [5 empty] | 80 | 5
| 30 | [9 empty] | 80 | 10 | [5 empty] | 80 | 5
| Max | [9 empty] | 80 | 7.5 | [5 empty] | 80 | 5