## Sheet1
1					Revision 1.1-20180517					Print Area
2					Medical University of South Carolina					Medical University of South Carolina
3					Charleston, South Carolina					Charleston, South Carolina
4					All:
5					Bone Mineral Density Compliance Inspection					Bone Mineral Density Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Survey ID:					Survey ID:
14					Room Number:					Room Number:
15			X-Ray Generator					Site Number:
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18				Max kVp:					Max mA:					Model:				Serial Number:
19					Max kVp:				Max mA:
20			X-Ray Tube 1					Focal Spot Sizes (mm)					X-Ray Generator
21				Tube Designation/Use:					Large:					X-Ray Tube 1					Focal Spot Sizes (mm)					Manufacturer:
22				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Model:
23			Insert					Micro:					Manufacture Date:				Small:					Max kVp:
24				Manufacturer:					Insert					Micro:					Manufacture Date:
25				Model:					Filtration					Manufacturer:					Serial Number:
26				Serial Number:					Inherent:					Model:				Filtration					Max mA:
27			Housing					Added:					Serial Number:				Inherent:					X-Ray Tube 1
28				Manufacturer:					Housing					Added:					Tube Designation/Use:
29				Model:					Manufacturer:					Manufacture Date:
30				Serial Number:					Model:					Manufacturer:
31					Serial Number:					Model:
32					Serial Number:
33					Inspection Results					Bone Mineral Density Compliance Inspection					Manufacturer:
34		Rule Number			Labels, Notices, Postings					Compliance					Inspection Results					Model:
35		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
36		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD					Labels, Notices, Postings					Large:
37		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD				DHEC Registration sticker is present, clearly visible and legible					Small:
38				Operator manuals are available.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Micro:
39				Monthly radiation monitoring reports are posted.					TBD				Radiation warning label posted on the generator control panel					Inherent:
40					Operator manuals are available.					Added:
41					Radiation Safety					Monthly radiation monitoring reports are posted.
42				Operator has full visibility of the exposure factors and patient					TBD					Equipment:		Change	Piranha CB2-17090320
43				Operator position is at least 1m away from patient and unit					TBD					Radiation Safety					Detector:		Change	Piranha R100B
44					Operator has full visibility of the exposure factors and patient
45					Physical Inspection					Operator position is at least 1m away from patient and unit
46				System clean and free of debris					TBD
47				All cables free from excessive wear or restraints					TBD					Physical Inspection
48				System free of missing or damaged components					TBD				System clean and free of debris
49				All other movements and locks functional					TBD				All cables free from excessive wear or restraints
50					System free of missing or damaged components
51					All other movements and locks functional
52
53
54					Scan Techniques					Measurement Locations
55					Field size
56					Mode	kV	mA	Time (s)	Trans (cm)	Long (cm)		Location	Distance (cm)
57					High Def	140/100	2.5	163	11.43	20.32		Tableside
58					Array	140/100	2.5	82	11.43	20.32		Operator
59					Fast Array	140/100	2.5	41	11.43	20.32		Door
60					Turbo	140/100	2.5	15	11.43	20.32
61
62					Background
63					Time (s):	82	Reading (nGy):					Avg Bkg: (nGy):
64
65		Date:					Inspector:	Eugene Mah		Scatter Measurements					Average
66		Room Number					Survey ID:			High Def	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
1					Medical University of South Carolina			Tableside
2					Inspection Results					Bone Mineral Density Compliance Inspection			Operator
3		Scan Techniques					Door
4					Field Size			Measurement Locations
5			Mode	kV	mA	Time (s)	Trans (cm)	Long (cm)		Location	Distance (cm)
6			High Def	140/100	2.5	163	11.43	20.32		Tableside				Array	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
7			Array	140/100	2.5	82	11.43	20.32		Operator					Tableside
8			Fast Array	140/100	2.5	41	11.43	20.32		Door					Operator
9			Turbo	140/100	2.5	15	11.43	20.32					Door
10
11
12					Fast Array	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
13					Tableside
14					Operator
15		Scatter Measurements					Exposure Measurements					Door
16
17		High Def	Location	Distance (cm)	nGy/scan	uGy/h		Mode	Detector (mGy)	Pt Entrance (mGy)
18			Tableside					Turbo	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
19			Operator					High Def					Tableside
20			Door					Array					Operator
21					Fast Array					Door
22					Turbo
23		Array	Location	Distance (cm)	nGy/scan	uGy/h
24			Tableside
25			Operator					Exposure measurements
26			Door					107	Source – Detector distance (cm)				Test equipment used:	Piranha CB2-17090320		Detector:	Piranha R100B
27					42.4	Source – Patient distance (cm)					Piranha CB2-17090320			Piranha R100B
28
29		Fast Array	Location	Distance (cm)	nGy/scan	uGy/h					Detector entrance dose measurements
30			Tableside					Average		Pt Entrance
31			Operator					Mode	mGy/scan			(mGy)	SD	(mGy)
32			Door					High Def
33					Array
34					Fast Array
35		Turbo	Location	Distance (cm)	nGy/scan	uGy/h					Turbo
36			Tableside
37			Operator
38			Door
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID: