## Sheet1
Revision 1.0-20230427 | [10 empty] | Print Area
Medical University of South Carolina
Charleston, South Carolina
[11 empty] | All: | Page1,Page2,HVLPage,KAP_AKR,ExpChartFluoroPage,ExpChartDAPage,ImgQualityPage,Com1Page,Com2Page,OutputGraphFluoroPage,OutputGraphDAPage,LeedsTO10Page
O-Arm System Compliance Inspection
| Date: | [6 empty] | Inspector: |  |  | Measurement Parameter | Last Year | This Year
| Previous Date: | [9 empty] | Print Area
[5 empty] | System Information | [5 empty] | Date
| Location |  | Input Changes Only |  |  |  |  | Input Changes Only |  |  | Inspector
| Facility: |  |  |  |  | Site Number: |  |  |  |  | Facility:
| Department: |  |  |  |  | Authorized Use: |  |  |  |  | Department:
| Area/Division: |  |  |  |  | Date of Installation: |  |  |  |  | Area/Division:
| Survey ID: |  |  |  |  | Accession Number: |  |  |  |  | Survey ID:
| Room Number: | [9 empty] | Room Number:
| X-Ray Generator | [9 empty] | Site Number:
| Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | Authorized Use:
| Model: |  |  |  |  | Serial Number: |  |  |  |  | Date of Installation:
| Max kVp: |  |  |  |  | Max mA: |  |  |  |  | Accesssion Number:
[6 empty] | Software Version: |  |  |  |  | X-Ray Generator
| X-Ray Tube | [9 empty] | Manufacturer:
| Tube Designation/Use: |  |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Model:
| Manufacture Date: |  |  |  |  | Large: |  |  |  |  | Max kVp:
| Insert |  |  |  |  | Small: |  |  |  |  | Manufacture Date:
| Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Serial Number:
| Model: | [9 empty] | Max mA:
| Serial Number: |  |  |  |  | Filtration |  |  |  |  | Software Version:
| Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
| Manufacturer: |  |  |  |  | Added: |  |  |  |  | Tube Designation/Use:
| Model: | [9 empty] | Manufacture Date:
| Serial Number: | [9 empty] | Manufacturer:
Model:
Serial Number:
Manufacturer:
Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Model:
|  | Is this a new installation? |  |  |  |  | Tube Number: |  |  |  | Serial Number:
[8 empty] | **Input only one of the following |  |  | Large:
Labels, Notices, Postings | [7 empty] | numbers or names |  |  | Small:
| DHEC Registration sticker is present, clearly visible and legible | [6 empty] | 1 | 1 |  | Micro:
| DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | 2 | 2 |  | Inherent:
| Radiation warning label posted on the generator control panel | [6 empty] | 3 | 3 |  | Added:
| Tube loading charts for each focal spot are posted | [6 empty] | AP | AP |  | New install:
| Tube loading charts are correct for installed tube(s). | [6 empty] | Front | Front |  | Tube Number:
| Operator manuals are available. | [6 empty] | Frontal | Frontal |  | Tube Designation: |  | 1
| Monthly radiation monitoring reports are posted. | [6 empty] | Lat | Lat |  |  |  | 2
[8 empty] | Lateral | Lateral |  |  |  | 3
Fluoroscopy Systems | [12 empty] | AP
| Interlock systems function to prevent fluoro while II is parked | [11 empty] | Front
| The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [11 empty] | Frontal
| Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels) | [11 empty] | Lat
| The useful beam is intercepted by the image assembly at any SID | [11 empty] | Lateral
| Fluoroscopy patient dose charts are posted | [9 empty] | Receptor Size
| Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [9 empty] | Normal:
| Max Entrance Exposure Rate <= 176 mGy/min (high level) | [9 empty] | Mag 1:
| High contrast resolution within MUSC specifications | [9 empty] | Mag 2:
| Low contrast resolution within MUSC specifications | [9 empty] | Pt Dose Measured Fluoro
| Transmission through II < 2 uGy/hr for each mGy/min at patient entrance | [9 empty] | Manual/Auto:
| Beam limiting device available (functioning) to restrict the x-ray field | [9 empty] | Lucite/Al:
| The useful beam is centered on the image receptor | [9 empty] | # dose rates:
| Collimator edges are visible at full open | [9 empty] | High level present:
| Elapsed timer (5 min maximum) functional | [9 empty] | cm/in:
| Fluoroscopic kV and mA meter present and functional | [9 empty] | Auto/mA:
Test equipment:
Radiation Safety | [10 empty] | Calibration date:
| Continuous pressure on exposure switch required to make an exposure with selected tube | [9 empty] | Calibration due:
| Operator has full visibility of the exposure factors and patient | [9 empty] | Technique 1:
| Lead aprons available | [9 empty] | Technique 2:
| Lead gloves available | [9 empty] | Technique 3:
| Properly designed and installed apron rack is present | [9 empty] | Maximum Dose Rates Fluoro
| Documentation of annual protective apparel integrity inspection is available | [9 empty] | kV: |  | 0
| Patient restraint devices available | [9 empty] | mA: |  | 0
| Gonadal shielding is present and functional | [9 empty] | mGy/min: |  | 0
[11 empty] | kV: |  | 0
Physical Inspection | [10 empty] | mA: |  | 0
| System clean and free of debris | [9 empty] | mGy/min: |  | 0
| All cables free from excessive wear or restraints | [9 empty] | kV: |  | 0
| System free of missing or damaged components | [9 empty] | mA: |  | 0
| Tube support movements and locks are acceptable | [9 empty] | mGy/min: |  | 0
| All other movements and locks functional | [9 empty] | HVL
| Unit installed as shown on shielding plan | [9 empty] | HVL Measured kV: |  | 80
[11 empty] | HVL Measured: |  | TBD
Automatic Brightness Control / Display Monitor | [10 empty] | Fluoro Scatter/Transmission
| Automatic Brightness Control maintains image brightness for varying attenuation | [9 empty] | Transmission: |  | 0
| Display monitor is free of artifacts, shadows, bright spots, and burn-in | [9 empty] | Unshielded Scatter (eyes):
| Display monitor is free of distortion | [9 empty] | Unshielded Scatter (waist):
| Display monitor is free of excessive persistence (after-image) | [9 empty] | Shielded Scatter (eyes):
| Display monitor is free of static or rolling noise bands | [9 empty] | Shielded Scatter (waist):
| The last image hold (LIH) display is free of jitter or abnormal appearance | [9 empty] | Receptor Entr Exp Fluoro
Normal:
Mag 1:
Entrance Exposure Rate | [10 empty] | Mag 2:
|  | Type 1 for manual, leave blank for auto | [6 empty] | Auto/mA mode |  | Normal:
|  | Type 1 if using Lucite, leave blank if using Al | [6 empty] | Test equipment |  | Mag 1:
|  | How many dose rate settings are present (exclusing High Level)? | [6 empty] | Calibration date |  | Mag 2:
|  | Is High Level provided? (1=Yes, 2=No) | [6 empty] | Calibration due |  | Normal:
cm | cm | Unit for image receptor size | [8 empty] | Mag 1:
Blank | [10 empty] | Mag 2:
Patient Entrance Exposure Rate (Fluoroscopy)* | [10 empty] | Leeds Phantom
Image | Acrylic | [9 empty] | kVp: |  | 70
Mode | Attenuator | [9 empty] | Huttner Normal:
(cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | Huttner Mag 1:
Normal | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Huttner Mag 2:
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N3 Normal:
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N3 Mag 1:
Mag 1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N3 Mag 2
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Field Shape:
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SID:
Mag 2 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Front:
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Back:
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Left:
| Max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Right:
| Last Year: | [9 empty] | TO.10.A
| *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES | TO.10.B
| Undertable fluoroscopy units: 1 cm above table top | [9 empty] | TO.10.C
| Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point) | [9 empty] | TO.10.D
| C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [9 empty] | TO.10.E
| Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point) | [9 empty] | TO.10.F
| Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [8 empty] | TO.10.G
|  | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [8 empty] | TO.10.H
| Comments: | [9 empty] | TO.10.J
| Do not exceed 390 chars |  | 0 | [7 empty] | TO.10.K
| New: | [9 empty] | TO.10.L
TO.10.M
Half Value Layer | [10 empty] | TO.10.A
2 | Use DHEC standards? (Type 1 to base acceptance on DHEC criteria) | [9 empty] | TO.10.B
TO.10.C
kV | mA | mm Al | mGy/min |  | Data Processing | [5 empty] | TO.10.D
80 |  | 0 |  |  | mm Al | Avg Exp | DATA(1/2) | Data Al | HVL |  | TO.10.E
|  | 0 |  |  | 0 |  |  |  | TBD |  | TO.10.F
TO.10.G
TO.10.H
[8 empty] | Data Exp | LN(EXP) |  | TO.10.J
[5 empty] | TBD | TBD |  |  |  |  | TO.10.K
[5 empty] | Acceptable (mm Al) | [5 empty] | TO.10.L
[5 empty] | Low | 2.6 |  |  |  |  | TO.10.M
[5 empty] | Desired | 3.1 |  | HVL Acceptable: | TBD |  | TO.10.A
[5 empty] | High | 3.5 |  |  |  |  | TO.10.B
[8 empty] | Previous HVL @ kV: |  |  | TO.10.C
TO.10.D
Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate | [10 empty] | TO.10.E
| Source-PDC distance (cm) |  |  |  |  | Source-Isocenter |  |  | Test equipment |  | TO.10.F
| Source-Ref Pt distance (cm) |  |  |  |  | Table-Isocenter |  |  | Calibration date |  | TO.10.G
| Reference Point correction factor |  |  |  |  | RefPt-Isocenter |  |  | Calibration due |  | TO.10.H
TO.10.J
|  | Start KAP | End KAP | Measured | Meas KAP | Indicated |  |  |  |  | TO.10.K
kV | mA | µGy·m² |  | KAP | @ Ref Pt | KAP | Deviation |  |  |  | TO.10.L
TO.10.M
Comments
Pt Entr Exp Fluoro:
Receptor Entr Exp Fluoro:
Scatter/Transmission:
|  | Start AK | End AK | Measured | Meas CAK | Indicated |  |  |  |  | Leeds Phantom:
kV | mA | mGy |  | CAK | @ Ref Pt | CAK | Deviation |  |  |  | TO.10:
Beam Limitation:
CT Measurement Protocols
Protocol 1:
Protocol name:
|  | Measured | Meas AKR | Indicated | [6 empty] | kVp:
kV | mA | AKR | @ Ref Pt | AKR | Deviation | [5 empty] | mA:
mAs:
Time/rotation (s):
Scan FOV (cm):
Display FOV (cm):
Z-Axis collimation:
# of channels:
Criteria: | Measured KAP/AK/AKR is within 40% of the displayed values | [9 empty] | Recon scan width (mm):
Recon scan interval (mm):
Scatter and Transmission | [10 empty] | Indiated CTDIvol (mGy):
|  |  |  | Pt Entrance | [6 empty] | Protocol 2:
|  | kVp | mA | mGy/min | uGy/hr | Last Year |  | Acceptable |  |  | Protocol name:
| Transmission* @ 10cm |  |  |  | 0 |  |  | TBD |  |  | kVp:
|  |  |  | Unshielded | Shielded | Unshielded | Shielded | Unshielded | Shielded |  | mA:
|  |  |  | uGy/min | uGy/min | uGy/hr | uGy/hr | uGy/hr | uGy/hr |  | mAs:
Scatter – Fluoro | Eye level | [9 empty] | Time/rotation (s):
| Waist level | [9 empty] | Scan FOV (cm):
| *Measured with 8” Lucite in beam | [9 empty] | Display FOV (cm):
| Largest field size mode, highest dose rate mode | [9 empty] | Z-Axis collimation:
| Criteria: | Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure | [8 empty] | # of channels:
| Comments: | [9 empty] | Recon scan width (mm):
| Do not exceed 390 chars |  | 0 | [7 empty] | Recon scan interval (mm):
| New: | [9 empty] | Indiated CTDIvol (mGy):
Protocol 3:
Image Receptor Entrance Exposure Rate – Fluoroscopy | [10 empty] | Protocol name:
Dose | Receptor |  |  |  |  | Previous Year |  |  |  |  | kVp:
Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | nGy/frame |  |  |  |  | mA:
[8 empty] | Measure using pancake probe placed against |  |  | mAs:
[8 empty] | image receptor input window |  |  | Time/rotation (s):
[8 empty] | Use 1.5mm Cu filtration |  |  | Scan FOV (cm):
Display FOV (cm):
Z-Axis collimation:
# of channels:
Recon scan width (mm):
Recon scan interval (mm):
Indiated CTDIvol (mGy):
CTDI Measurements
| Comments: | [9 empty] | Protocol 1:
| Do not exceed 390 chars |  | 0 | [7 empty] | Protocol 2:
| New: | [9 empty] | Protocol 3:
kV Accuracy and Reproducibility
| Test Equipment Used: | Piranha |  |  | Calibration Date:
| Serial Number: | CB2-17090320 |  |  | Calibration Due:
kV Accuracy |  |  | Exposure | Exp rate
kV set | mA Set | kV | (mGy) | (mGy/min) | HVL | Tot Filt | kV % error
60 | 5
80 | 5
100 | 5
120 | 5
| Criteria: | kVp accurate to within 5% of indicated. |  |  |  | Acceptable:
kV/exposure reproducibility
|  | Exposure | Exp rate |  |  |  | Air Kerma Rate Linearity
kV Set: | kV | (mGy) | (mGy/min) | HVL | Tot Filt |  | kV Set: | mA | mGy/min
80 | [6 empty] | 80 | 1
10
7.5
5
Coeff of Linearity:
Mean: | [7 empty] | Acceptable:
Std Dev: | [7 empty] | Criteria: | Coefficient of linearity <= 10%.
Coeff of Var:
Acceptable:
| Criteria: | kVp coefficient of variation (std dev./mean) <= 5%.
Leeds Phantom
Receptor Size | MS1 | MS3 | MS4 | Huttner | N3 | kVp: | 70 | 70
[7 empty] | MS1-4 |  | Huttner: Enter col.row
[7 empty] | 1 | Clearly visible | N3: Enter # of visible disks
[7 empty] | 2 | Visible
|  |  | Acceptable: |  |  |  | 3 | Barely visible
[7 empty] | 4 | Not visible
Comments:
| Do not exceed 390 chars |  | 0
New:
TO.10 |  |  |  |  | Field Size
Row | [6 empty] | Type 1 for circular field, 2 for square field
A
B
C |  |  |  |  | Horizontal:
D |  |  |  |  | Vertical:
E |  |  |  |  | Diagonal:
F |  |  |  |  | Acceptable: | NA | NA | NA
G | [5 empty] | Criteria: | Measured field size is at least 85% of nominal field size
H
J
K
L
M
Comments:
| Do not exceed 390 chars |  | 0
New:
Useful Beam Limitation
SID (cm)
| Measured | Difference | Sum | Last Year |  |  | T | B | L | R
Top: | [5 empty] | Visible:
Bottom: | [5 empty] | X-ray:
Left: |  |  |  |  | Acceptable | SID:
Right: | [5 empty] | SOD:
Criteria: | Difference between x-ray field and visible image area is less than 2% SID |  |  |  |  | SID/SOD:
Sum of length and width differences is less than 4% SID
Comments:
| Do not exceed 390 chars |  | 0
New:
|  | Protocol 1 | Protocol 2 | Protocol 3
Protocol name:
kVp:
mA:
mAs:
Time/rotation (s):
Scan FOV (cm):
Display FOV (cm):
Z-Axis collimation:
# of channels:
Recon scan width (mm):
Recon scan interval (mm):
Indiated CTDI (mGy):
CTDI Measurements
| Beam width (mm): | [5 empty] | Meter: | Piranha |  | Calibration Date:
| Chamber length (mm): | [5 empty] | Serial #: |  |  | Calibration Due:
Chamber correction factor:
| Protocol: | [7 empty] | Reference | Pass/Fail
| Indicated CTDIvol (mGy): | [7 empty] | (mGy) | (mGy)
| Center (mGy) | [6 empty] | Adult Abdomen | 25 | 30
[8 empty] | Adult Head: | 75 | 80
[8 empty] | Ped Abd (16cm) | 15 | 20
| Center CTDI (mGy): | [6 empty] | Ped Abd (32cm) | 7.5 | 10
| Periphery (mGy) | [6 empty] | Pediatric Head: | 35 | 40
Periphery CTDI (mGy):
CTDIw (mGy):
Phantom size indicated (cm)
Phantom size used (cm)
Previous year:
Variation from prev year:
Variation from indicated:
Acceptable:
CTDIw/100 mAs:
Prev CTDIw/100 mAs:

## Sheet2
CBCT CTDI
| Average |  | f(0) | #DIV/0!
Center (mGy) | #DIV/0! | f(center)
Top | #DIV/0! | f(top)
Left | #DIV/0! | f(left)
Right | #DIV/0! | f(right)
Bottom | #DIV/0! | f(bottom)
Top-left | #DIV/0! | f(TL)
Top-right | #DIV/0! | f(TR)
Bottom-left | #DIV/0! | f(BL)
Bottom-right | #DIV/0! | f(BR)

## Tables
Leeds MS
1 | Clearly visible
2 | Visible
3 | Barely visible
4 | Not visible
Huttner Grating |  |  | Leeds N3 |  | 65 | 70 | 75 | 80 | Nominal
1.1 | 0.5 |  | 1 | 0.167 | 0.188 | 0.167 | 0.149 | 0.135 | 0.16
1.2 | 0.56 |  | 2 | 0.148 | 0.167 | 0.148 | 0.132 | 0.12 | 0.145
1.3 | 0.63 |  | 3 | 0.128 | 0.145 | 0.128 | 0.114 | 0.104 | 0.123
2.1 | 0.71 |  | 4 | 0.109 | 0.124 | 0.109 | 0.097 | 0.088 | 0.108
2.2 | 0.8 |  | 5 | 0.0876 | 0.0993 | 0.0876 | 0.078 | 0.0705 | 0.086
2.3 | 0.9 |  | 6 | 0.0749 | 0.0847 | 0.0749 | 0.067 | 0.0609 | 0.076
3.1 | 1 |  | 7 | 0.0674 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | 0.066
3.2 | 1.12 |  | 8 | 0.0525 | 0.0593 | 0.0525 | 0.047 | 0.0427 | 0.055
3.3 | 1.25 |  | 9 | 0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | 0.045
4.1 | 1.4 |  | 10 | 0.0371 | 0.04 | 0.0371 | 0.0347 | 0.0329 | 0.039
4.2 | 1.6 |  | 11 | 0.0322 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | 0.033
4.3 | 1.8 |  | 12 | 0.0246 | 0.0269 | 0.0246 | 0.0233 | 0.022 | 0.027
5.1 | 2 |  | 13 | 0.0215 | 0.0232 | 0.0215 | 0.0201 | 0.019 | 0.023
5.2 | 2.24 |  | 14 | 0.0172 | 0.0186 | 0.0172 | 0.0161 | 0.0153 | 0.018
5.3 | 2.5 |  | 15 | 0.0155 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | 0.016
6.1 | 2.8 |  | 16 | 0.013 | 0.014 | 0.013 | 0.0122 | 0.0115 | 0.013
6.2 | 3.15 |  | 17 | 0.011 | 0.0119 | 0.011 | 0.0103 | 0.0097 | 0.011
6.3 | 3.55 |  | 18 | 0.0086 | 0.0093 | 0.0086 | 0.0081 | 0.0077 | 0.009
7.1 | 4 |  | 19 | 0.0066 | 0.0071 | 0.0066 | 0.0062 | 0.0059 | 0.007
7.2 | 4.5
7.3 | 5
TO.10 kV Adj | 65 | 70 | 75 | 80
0.012 | 0.0119 | 0.011 | 0.0103 | 0.0097 |  |  |  |  | .
0.016 | 0.0168 | 0.0155 | 0.0145 | 0.0138
0.023 | 0.0232 | 0.0215 | 0.0201 | 0.019
0.032 | 0.0347 | 0.0322 | 0.0301 | 0.0285
0.045 | 0.0512 | 0.045 | 0.0399 | 0.036
0.066 | 0.0766 | 0.0674 | 0.0599 | 0.0542
0.086 | 0.0993 | 0.0876 | 0.078 | 0.0705
0.123 | 0.145 | 0.128 | 0.114 | 0.104
0.16 | 0.188 | 0.167 | 0.149 | 0.135
0.23 | 0.267 | 0.238 | 0.214 | 0.194
0.35 | 0.399 | 0.36 | 0.327 | 0.299
0.5 | 0.618 | 0.573 | 0.533 | 0.497
0.66 | 0.769 | 0.726 | 0.686 | 0.649
0.93 | 0.969 | 0.954 | 0.936 | 0.915
TO.10 | Diameter (mm) | sqrt(Area) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 11.1 | 9.8371 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
B | 7.9 | 7.0012 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
C | 5.6 | 4.9629 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012
D | 4 | 3.5449 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
E | 2.8 | 2.4814 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
F | 2 | 1.7725 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
G | 1.4 | 1.2407 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
H | 1 | 0.88623 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
J | 0.7 | 0.62036 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
K | 0.5 | 0.44311 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
L | 0.35 | 0.31018 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
M | 0.25 | 0.22156 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
TO.10 Values | 0 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9
A | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
B | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
C | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
D | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
E | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
F | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
G | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
H | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
J | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
K | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
L | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
M | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
Half Value Layer
| DHEC |  | 0.915 2: MUSC; 0.915 6: HVL Graph
| Min HVL |  | 0.915 2: HVL; 0.915 3: Min HVL; 0.915 4: Max HVL; 0.915 6: mm Al; 0.915 7: Avg Exp
kVp | (mm Al) |  | 0.936 1: kVp; 0.915 2: (mm Al); 0.915 3: (mm Al); 0.915 4: (mm Al)
30 | 0.3 |  | 0.936 1: 60; 0.915 2: 2.2; 0.915 3: 1.8; 0.915 4: 2.4
40 | 0.4 |  | 0.936 1: 70; 0.915 2: 2.6; 0.915 3: 2.2; 0.915 4: 2.9
50 | 0.5 |  | 0.936 1: 80; 0.915 2: 3; 0.915 3: 2.6; 0.915 4: 3.5
51 | 1.2 |  | 0.936 1: 90; 0.915 2: 3.5; 0.915 3: 3.1; 0.915 4: 4
60 | 1.3 |  | 0.936 1: 100; 0.915 2: 3.9; 0.915 3: 3.5; 0.915 4: 4.5
70 | 1.5 |  | 0.936 1: 110; 0.915 2: 4.3; 0.915 3: 3.9; 0.915 4: 5.1
71 | 2.1 |  | 0.936 1: 120; 0.915 2: 5.7; 0.915 3: 4.4; 0.915 4: 5.6
80 | 2.3 |  | 0.936 1: 130; 0.915 2: 5.2; 0.915 3: 4.8; 0.915 4: 6.1
90 | 2.5
100 | 2.7
110 | 3
120 | 3.2
130 | 3.5
140 | 3.8
150 | 4.1
Note on A30: Eugene Mah: Nominal contrast values as a function of kVp
Note on A46: Eugene Mah: TO.10 nominal contrast values
Note on A60: Eugene Mah: TO.10 contrast values at the kVp specified in Sheet1!I229

## Data Entry
Patient Entrance Exposure Rate (Fluoroscopy)* | [10 empty] | Measurements | [6 empty] | Measurements
Image | Acrylic
Mode | Attenuator | [9 empty] | Set kV (kV) | Set mA (mA) | Tube voltage (kV) | Exposure (mGy) | Exposure rate (mGy/min) | HVL (mm Al) | Total filtr. (mm Al) | Set kV (kV) | Set mA (mA) | Tube voltage (kV) | Exposure (mGy) | Exposure rate (mGy/min) | HVL (mm Al) | Total filtr. (mm Al) | Set frames/s (FPS) | Frames/s (FPS) | Exposure/frame (µGy/frame) | Frame exp. rate (mGy/s) | Pulse width (ms)
(cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | 60 | 5 | [5 empty] | 60 | 5
Normal | 10 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 20 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 30 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
Mag 1 | 10 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 20 | [9 empty] | 80 | 5 | [5 empty] | 80 | 5
| 30 | [9 empty] | 100 | 5 | [5 empty] | 100 | 5
Mag 2 | 10 | [9 empty] | 120 | 5 | [5 empty] | 120 | 5
| 20 | [9 empty] | 80 | 1 | [5 empty] | 80 | 5
| 30 | [9 empty] | 80 | 10 | [5 empty] | 80 | 5
| Max | [9 empty] | 80 | 7.5 | [5 empty] | 80 | 5