## Fluoro
1					Revision 1.0-20140806					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Fluoroscopy System Compliance Inspection					Fluoroscopy System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:				Print Area		Change	Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Number of X-ray Tubes:					Survey ID:				Number of X-ray Tubes:					Survey ID:
14					Room Number:				Accession Number:					Room Number:
15			X-Ray Generator					Site Number:
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18				Max kVp:					Max mA:					Model:				Serial Number:					Number of X-ray Tubes:
19					Max kVp:				Max mA:					Accesssion Number:
20			X-Ray Tube 1					Focal Spot Sizes (mm)					X-Ray Generator
21				Tube Designation/Use:					Large:					X-Ray Tube 1					Focal Spot Sizes (mm)					Manufacturer:
22				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Model:
23			Insert					Micro:					Manufacture Date:				Small:					Max kVp:
24				Manufacturer:					Insert					Micro:					Manufacture Date:
25				Model:					Filtration					Manufacturer:					Serial Number:
26				Serial Number:					Inherent:					Model:				Filtration					Max mA:
27			Housing					Added:					Serial Number:				Inherent:					X-Ray Tube 1
28				Manufacturer:					Housing					Added:					Tube Designation/Use:
29				Model:					Manufacturer:					Manufacture Date:
30				Serial Number:					Model:					Manufacturer:
31					Serial Number:					Model:
32			X-Ray Tube 2					Focal Spot Sizes (mm)					Serial Number:
33				Tube Designation/Use:					Large:					X-Ray Tube 2					Focal Spot Sizes (mm)					Manufacturer:
34				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Model:
35			Insert					Micro:					Manufacture Date:				Small:					Serial Number:
36				Manufacturer:					Insert					Micro:					Large:
37				Model:					Filtration					Manufacturer:					Small:
38				Serial Number:					Inherent:					Model:				Filtration					Micro:
39					Added:					Serial Number:				Inherent:					Inherent:
40				Manufacturer:					Added:					Added:
41				Model:					Manufacturer:					X-Ray Tube 2
42				Serial Number:					Model:					Tube Designation/Use:
43					Serial Number:					Manufacture Date:
44					Manufacturer:
45					Model:
46					Inspection Results					Serial Number:
47					Enter 1 for YES, 2 for NO, 3 for NA					Manufacturer:
48					Labels, Notices, Postings		Tube Number:					Model:
49					DHEC Registration sticker is present, clearly visible and legible					**Input only one of the following				Serial Number:
50					DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					numbers or names				Large:
51					Radiation warning label posted on the generator control panel					1	1			Small:
52					Tube loading charts for each focal spot are posted					2	2			Micro:
53					Tube loading charts are correct for installed tube(s).					3	3			Inherent:
54					Operator manuals are available.					AP	AP			Added:
55					Monthly radiation monitoring reports are posted.					Front	Front			Tube Number:
56					Fluoroscopy Systems			Frontal	Frontal			Tube Designation:		Change	1
57					Interlock systems function to prevent fluoro while II is parked					Lat	Lat					Change	2
58					The minimum Source Skin Distance > 38 cm (20 cm for surgical applications)					Lateral	Lateral					Change	3
59					Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels)					Change	AP
60					The useful beam is intercepted by the image assembly at any SID					Change	Front
61					Fluoroscopy patient dose charts are posted					Change	Frontal
62					1	Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Change	Lat
63					3	Max Entrance Exposure Rate <= 176 mGy/min (high level)					Change	Lateral
64					High contrast resolution within MUSC specifications
65		Date:					Inspector:			1	Low contrast resolution within MUSC specifications
66		Room Number					Survey ID:				Transmission through II < 2 uGy/hr for each mGy/min at patient entrance					Receptor Size
1					Medical University of South Carolina			Beam limiting device available (functioning) to restrict the x-ray field					Normal:
2					Inspection Results					Fluoroscopy System Compliance Inspection			The useful beam is centered on the image receptor					Mag 1:
3		Rule Number			Labels, Notices, Postings					Compliance				Collimator edges are visible at full open					Mag 2:
4		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD				Elapsed timer (5 min maximum) functional					Mag 3:
5		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD				Fluoroscopic kV and mA meter present and functional					Mag 4:
6		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Radiation Safety					Pt Dose Measured Fluoro
7				Tube loading charts for each focal spot are posted					TBD				Continuous pressure on exposure switch required to make an exposure with selected tube					Manual/Auto:
8				Tube loading charts are correct for installed tube(s).					TBD				Operator has full visibility of the exposure factors and patient					Lucite/Al:		Change	1
9				Operator manuals are available.					TBD				Lead aprons available					# dose rates:
10				Monthly radiation monitoring reports are posted.					TBD				Lead gloves available					High level present:
11					Properly designed and installed apron rack is present					cm/in:		Change	cm
12					Radiation Safety					Documentation of annual protective apparel integrity inspection is available					Auto/mA:		Change	Auto
13				Continuous pressure on exposure switch required to make an exposure with selected tube					TBD				Patient restraint devices available					Technique 1:
14				Operator has full visibility of the exposure factors and patient					TBD				Gonadal shielding is present and functional					Technique 2:
15				Lead aprons available					TBD				Radiation warning light at room entrance is functional with selected tube					Technique 3:
16				Lead gloves available					TBD					Physical Inspection					Maximum Dose Rates Fluoro
17				Properly designed and installed apron rack is present					TBD				System clean and free of debris					kV:
18				Documentation of annual protective apparel integrity inspection is available					TBD				All cables free from excessive wear or restraints					mA:
19				Patient restraint devices available					TBD				System free of missing or damaged components					mGy/min:
20				Gonadal shielding is present and functional					TBD				Tube support movements and locks are acceptable					kV:
21				Radiation warning light at room entrance is functional with selected tube					TBD				All other movements and locks functional					mA:
22					Automatic Brightness Control / Display Monitor					mGy/min:
23					Fluoroscopy Systems					Automatic Brightness Control maintains image brightness for varying attenuation					kV:
24		DHEC RHB 4.9.2.1.2		Interlock systems function to prevent fluoro while II is parked					TBD				Display monitor is free of artifacts, shadows, bright spots, and burn-in					mA:
25		DHEC RHB 4.9.1		The minimum Source Skin Distance > 38 cm (20 cm for surgical applications)					TBD				Display monitor is free of distortion					mGy/min:
26		DHEC RHB 4.9.8		Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels)					TBD				Display monitor is free of excessive persistance (after-image)					Maximum Dose Rates Pulse/Dig
27		DHEC RHB 4.9.2.1.2		The useful beam is intercepted by the image assembly at any SID					TBD				Display monitor is free of static or rolling nosie bands					kV:
28		DHEC RHB 4.9.4.3.6		Fluoroscopy patient dose charts are posted					TBD				The last image hold (LIH) display is free of jitter or abnormal appearance					mA:
29		DHEC RHB 4.9.4.2.1		Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					YES					mGy/min:
30		DHEC RHB 4.9.4.2.1.2		Max Entrance Exposure Rate <= 176 mGy/min (high level)					NA					kV:
31				High contrast resolution within MUSC specifications					TBD					Inspection Results					mA:
32				Low contrast resolution within MUSC specifications					YES					mGy/min:
33		DHEC RHB 4.9.5.1		Transmission through II < 2 uGy/hr for each mGy/min at patient entrance					TBD			Entrance Exposure Rate				Type 1 for manual, leave blank for auto			Auto	Auto	Auto/mA mode			kV:
34		DHEC RHB 4.9.2.2.1		beam limiting device available (functioning) to restrict the x-ray field					TBD					1	1	Type 1 if using Lucite, leave blank if using Al					mA:
35				The useful beam is centered on the image receptor					TBD					How many dose rate settings are present (exclusing High Level)?					mGy/min:
36		DHEC RHB 4.9.2.2.1		Collimator edges are visible at full open					TBD					Is High Level provided? (1=Yes, 2=No)					HVL
37		DHEC RHB 4.9.7		Elapsed timer (5 min maximum) functional					TBD			Blank					HVL Measured kV:		Change	80
38		DHEC RHB 4.9.6		Fluoroscopic kV and mA meter present and functional					TBD			cm	cm	Unit for image receptor size					HVL Measured:		Change	TBD
39		DHEC RHB 4.3.5		The Measured HVL meeds DHEC minimum requirements					Patient Entrance Exposure Rate (Fluoroscopy)*					Pt Dose Measured Dig Acq
40					Image	Acrylic					Technique 1:
41					Automatic Brightness Control / Display Monitor					Mode	Attenuator	Auto 1					Auto				Technique 2:
42				Automatic Brightness Control maintains image brightness for varying attenuation					TBD			(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min		Technique 3:
43				Display monitor is free of artifacts, shadows, bright spots, and burn-in					TBD			Normal	10.00					Receptor Entr Exp Fluoro
44				Display monitor is free of distortion					TBD				20.00					Auto 1 Normal
45				Display monitor is free of excessive persistance (after-image)					TBD				30.00					Auto 1 Mag 1
46				Display monitor is free of static or rolling nosie bands					TBD			Mag 1	10.00					Auto 1 Mag 2
47				The last image hold (LIH) display is free of jitter or abnormal appearance					TBD				20.00					Auto 1 Mag 3
48					30.00					Auto 1 Mag 4
49					Physical Inspection					Mag 2	10.00					Normal
50				System clean and free of debris					TBD				20.00					Mag 1
51				All cables free from excessive wear or restraints					TBD				30.00					Mag 2
52				System free of missing or damaged components					TBD			Mag 3	10.00					Mag 3
53				Tube support movements and locks are acceptable					TBD				20.00					Mag 4
54				All other movements and locks functional					TBD				30.00					Auto Normal
55					Mag 4	10.00					Auto Mag 1
56					20.00					Auto Mag 2
57					30.00					Auto Mag 3
58					max					Auto Mag 4
59					Last Year:					Receptor Entr Exp Dig Acq
60					*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	NA		Auto 1 Normal
61					Undertable fluoroscopy units: 1 cm above table top					Auto 1 Mag 1
62					Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Auto 1 Mag 2
63					C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					Auto 1 Mag 3
64					Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					Auto 1 Mag 4
65		Date:					Inspector:				Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Normal
66		Room Number					Survey ID:					Max Entrance Exposure Rate < = 176 mGy/min (high level)					Mag 1
1					Medical University of South Carolina			Comments:					Mag 2
2					Inspection Results					Fluoroscopy System Compliance Inspection			Do not exceed 390 chars					Mag 3
3					Half Value Layer Measurement					New:					Mag 4
4					Auto Normal
5				Set kVp:	80.00					Patient Entrance Exposure Rate – Digital Acquisition					Mode:		Enter 1 for Pulse fluoro, 2 for digital acquisition mode					Auto Mag 1
6					Image	Acrylic					Auto Mag 2
7				Al (mm)	mGy/min	mGy/min		Measured HVL:	TBD					Mode	Attenuator	Auto 1					Auto				Auto Mag 3
8				0.0				HVL Acceptable:	TBD					(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min		Auto Mag 4
9					Normal	10.00					Fluoro Scatter/Transmission
10					ACCEPATBLE RESULTS (mm Al)					20.00					Transmission:
11					Min acceptable HVL at 80 kVp:	2.60					30.00					Scatter (eyes):
12					Desired HVL at 80 kVp:	3.10				Mag 1	10.00					Scatter (waist):
13					Criteria:	MUSC Standard					20.00					Dig Acq Scatter
14					30.00					Scatter (eyes):
15					Mag 2	10.00					Scatter (waist):
16					20.00					Leeds Phantom
17					30.00					kVp:		Change	70
18					Mag 3	10.00					Huttner Normal:
19					20.00					Huttner Mag 1:
20					30.00					Huttner Mag 2:
21					Mag 4	10.00					Huttner Mag 3:
22					20.00					Huttner Mag 4:
23					30.00					N3 Normal:
24					max					N3 Mag 1:
25					Last Year:					N3 Mag 2
26					*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	NA		N3 Mag 3
27					Undertable fluoroscopy units: 1 cm above table top					N3 Mag 4:
28					Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Field Size:
29					C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					SID:
30					Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					Front:
31					Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Back:
32					Max Entrance Exposure Rate < = 176 mGy/min (high level)					Left:
33					Comments:					Right:
34					Do not exceed 390 chars					TO.10.A
35					New:					TO.10.B
36					TO.10.C
37					Half Value Layer					TO.10.D
38					2	Use DHEC standards? (Type 1 to base acceptance on DHEC criteria)					TO.10.E
39					TO.10.F
40					kV	mA	mm Al	mGy/min		Data Processing					TO.10.G
41					80		0.0			mm Al	Avg Exp	DATA(1/2)	Data Al	HVL			TO.10.H
42					0.0			0.00				TBD			TO.10.J
43					TO.10.K
44					TO.10.L
45					Data Exp	LN(EXP)			TO.10.M
46					TO.10.A
47					TO.10.B
48					TBD	TBD					TO.10.C
49					Acceptable (mm Al)			HVL Acceptable:	TBD			TO.10.D
50					Low	2.60					TO.10.E
51					Desired	3.10		Previous HVL @ kV:				TO.10.F
52					High	3.50					TO.10.G
53					TO.10.H
54					Image Receptor Entrance Exposure Rate – Fluoroscopy					TO.10.J
55					TO.10.K
56					Dose	Receptor					Previous Year					TO.10.L
57					Level	Size (cm)	kVp	mA	uGy/min	nGy/frame	nGy/frame					TO.10.M
58					Auto 1					Measure using pancake probe placed against				TO.10.A
59					image receptor input window				TO.10.B
60					Use 1.5mm Cu filtration				TO.10.C
61					TO.10.D
62					TO.10.E
63					TO.10.F
64					TO.10.G
65		Date:					Inspector:					TO.10.H
66		Room Number					Survey ID:					TO.10.J
1					Medical University of South Carolina					TO.10.K
2					Inspection Results					Fluoroscopy System Compliance Inspection		Auto					TO.10.L
3		Patient Entrance Exposure Rate – Fluoroscopy					TO.10.M
4		Image	Acrylic					TO.10.A
5		Mode	Attenuator	Auto 1					Auto					TO.10.B
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min					TO.10.C
7		cm	10.00					TO.10.D
8			20.00					Comments:					TO.10.E
9			30.00					TO.10.F
10		cm	10.00					New:					TO.10.G
11			20.00					TO.10.H
12			30.00					Image Receptor Entrance Exposure Rate – Digital Acquisition					TO.10.J
13		cm	10.00					TO.10.K
14			20.00					Dose	Receptor					Previous Year					TO.10.L
15			30.00					Level	Size (cm)	kVp	mA	uGy/min	nGy/frame	nGy/frame					TO.10.M
16		cm	10.00					Auto 1					Measure using pancake probe placed against				TO.10.A
17			20.00					image receptor input window				TO.10.B
18			30.00					Use 1.5mm Cu filtration				TO.10.C
19		cm	10.00					TO.10.D
20			20.00					TO.10.E
21			30.00					TO.10.F
22			max					TO.10.G
23					TO.10.H
24			*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	NA					TO.10.J
25			Undertable fluoroscopy units: 1 cm above table top					TO.10.K
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Auto					TO.10.L
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					TO.10.M
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					Comments
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Pt Entr Exp Fluoro:
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)					Pt Entr Exp Dig Acq:
31		Comments:					Receptor Entr Exp Fluoro:
32					Comments:					Receptor Entr Exp Dig Acq:
33					Do not exceed 390 chars					Scatter/Transmission:
34					New:					Leeds Phantom:
35		Image Receptor Entrance Exposure Rate – Fluoroscopy					TO.10:
36					Scatter and Transmission					Beam Limitation:
37		Dose	Receptor					Patient					Comment Page 1
38		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame					Entrance
39		Auto 1					Measure using pancake probe placed against					kVp	mA	mGy/min	uGy/hr	Last Year		Acceptable
40					image receptor input window					Transmission* @ 10cm					TBD
41					Use 1.5mm Cu filtration					Scatter – Fluoro	Eye level
42					Waist level
43					Scatter – Digital acq	Eye level
44					Waist level
45					*Measured with 8” Lucite in beam
46					Largest field size mode, highest dose rate mode
47					Criteria:	Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure
48
49		Auto					Comments:
50					Do not exceed 390 chars
51					New:
52
53					Leeds Phantom
54		Comments:
55					Receptor Size	MS1	MS3	MS4	Huttner	N3	kVp:	70	70
56					MS1-4		Huttner: Enter col.row
57		Scatter and Transmission – Fluoroscopy				Pt Entrance					1	Clearly visible	N3: Enter # of visible disks
58				kVp	mA	mGy/min	uGy/hr		Acceptable					2	Visible
59			Transmission* @ 10cm				0		TBD					3	Barely visible
60		Scatter – Fluoro	Eye level				0					4	Not visible
61			Waist level				0					Acceptable:	TBD	TBD
62		Comments:					Comments:
63					Do not exceed 390 chars
64					New:
65		Date:					Inspector:
66		Room Number					Survey ID:			TO.10		Details Detected
1					Medical University of South Carolina			Row
2					Inspection Results					Fluoroscopy System Compliance Inspection			A
3		Patient Entrance Exposure Rate – Digital Acquisition					B
4		Image	Acrylic					C
5		Mode	Attenuator	Auto 1					Auto					D
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min			E
7		cm	10.00					F
8			20.00					G					Comment Page 2
9			30.00					H
10		cm	10.00					J
11			20.00					K
12			30.00					L
13		cm	10.00					M
14			20.00
15			30.00					Comments:
16		cm	10.00					Do not exceed 390 chars
17			20.00					New:
18			30.00
19		cm	10.00					Field Size
20			20.00					Horizontal					Type 1 for circular field, 2 for square field
21			30.00					Vertical
22			max					Diagonal
23					Acceptable	NA	NA	NA	NA	NA
24		Comments:					Criteria:	Measured field size is at least 85% of nominal field size
25					Useful Beam Limitation
26					SID (cm)
27
28					Measured	Difference	Sum	Last Year			T	B	L	R
29					Top					Visible
30					Bottom					X-ray
31					Left					SID
32		Image Receptor Entrance Exposure Rate – Digital Acquisition					Right					SOD
33					Criteria:	Difference between x-ray field and visible image area is less than 2% SID					SID/SOD
34		Dose	Receptor					Sum of length and width differences is less than 3% SID
35		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame					Comments:
36		Auto 1					Measure using pancake probe placed against					Do not exceed 390 chars
37					image receptor input window					New:
38					Use 1.5mm Cu filtration
39
40					Comments Page 1
41					Previous Year Comments
42					Comments:
43					Do not exceed 390 chars					**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
44					Additional Comments:
45					Do not exceed 390 chars					**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
46		Auto					Additional Comments:
47					Do not exceed 390 chars					A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
48					Additional Comments:
49					Do not exceed 390 chars					All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
50					Additional Comments:
51		Comments:					Do not exceed 390 chars					**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
52					Additional Comments:
53					Do not exceed 390 chars
54					Additional Comments:
55		Scatter and Transmission – Digital Acquisition					Do not exceed 390 chars
56					Additional Comments:
57				kVp	mA	uGy/hr					Do not exceed 390 chars
58		Scatter – Digital acq	Eye level			0					Additional Comments:
59			Waist level			0					Do not exceed 390 chars
60					Additional Comments:
61					Do not exceed 390 chars
62		Comments:					Additional Comments:
63					Do not exceed 390 chars
64					Additional Comments:
65		Date:					Inspector:				Do not exceed 390 chars
66		Room Number					Survey ID:				Additional Comments:
1					Medical University of South Carolina			Do not exceed 390 chars
2					Inspection Results					Fluoroscopy System Compliance Inspection			Additional Comments:
3		Leeds Test Objects					High contrast resolution		Low contrast resolution					Do not exceed 390 chars
4			Receptor Size	MS1	MS3	MS4	lp/mm	Acceptable	N3	Acceptable					Additional Comments:
5					Do not exceed 390 chars
6					Additional Comments:
7					Do not exceed 390 chars
8					Additional Comments:
9					Do not exceed 390 chars
10			Criteria:	Resolution: >1.0 lp/mm Normal mode, >1.25 lp/mm Mag 1, >1.4 lp/mm Mag 2, > 2 lp/mm Mag 3, > 2.5 lp/mm Mag 4					Additional Comments:
11				N3: Minimum contrast level < 0.04 in all II modes					Do not exceed 390 chars
12		Comments:					Additional Comments:
13					Do not exceed 390 chars
14
15
16		TO.10 Contrast Resolution					Comments Page 2
17			Contrast Detail					Threshold Index					Previous Year Comments
18		Row					Comments:
19		A					Do not exceed 390 chars
20		B					Additional Comments:
21		C					Do not exceed 390 chars
22		D					Additional Comments:
23		E					Do not exceed 390 chars
24		F					Additional Comments:
25		G					Do not exceed 390 chars
26		H					Additional Comments:
27		J					Do not exceed 390 chars
28		K					Additional Comments:
29		L					Do not exceed 390 chars
30		M					Additional Comments:
31		Comments:					Do not exceed 390 chars
32					Additional Comments:
33					Do not exceed 390 chars
34					Additional Comments:
35		Field Size					Do not exceed 390 chars
36					Additional Comments:
37			Horizontal					Do not exceed 390 chars
38			Vertical					Additional Comments:
39			Diagonal					Do not exceed 390 chars
40			Acceptable	NA	NA	NA	NA	NA					Additional Comments:
41			Criteria:	Measured field size is at least 85% of nominal field size					Do not exceed 390 chars
42					Additional Comments:
43		Useful Beam Limitation					Do not exceed 390 chars
44			SID (cm)					Additional Comments:
45				Measured	Difference	Sum					Do not exceed 390 chars
46			Top					Additional Comments:
47			Bottom					Do not exceed 390 chars
48			Left					Additional Comments:
49			Right					Do not exceed 390 chars
50			Criteria:	Difference between x-ray field and visible image area is less than 2% SID					Additional Comments:
51				Sum of length and width differences is less than 3% SID					Do not exceed 390 chars
52		Comments:					Additional Comments:
53					Do not exceed 390 chars
54					Additional Comments:
55					Do not exceed 390 chars
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Comments Page 1					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Comments Page 2					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Radiation Output – Fluoroscopy					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Radiation Output – Digital Acquisition					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Contrast Resolution					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
Half Value Layer
mm Al
Exposure Rate (mGy/min)
Radiation Output
kV
mGy/mA-min
Fluoro Patient Entrance Exposure
Attenuator thickness (cm)
mGy/min
Radiation Output
kV
mGy/mA-min
Fluoro Patient Entrance Exposure
Attenuator thickness (cm)
mGy/min
Contrast Detail Diagram
Detail diameter (mm)
Threshold Detection Index
sqrt(Detail area)

## ExpCharts
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Fluoroscopy
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	R/min	kVp	mA	R/min	kVp	mA	R/min
7		cm	10.00
8		0.00	20.00
9		0.00	30.00
10		cm	10.00
11		0.00	20.00
12		0.00	30.00
13		cm	10.00
14		0.00	20.00
15		0.00	30.00
16		cm	10.00
17		0.00	20.00
18		0.00	30.00
19		cm	10.00
20		0.00	20.00
21		0.00	30.00
22			max
23
24			*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	NA
25			Undertable fluoroscopy units: 1 cm above table top
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)
31			Comments:
32
33
34
35		Image Receptor Entrance Exposure Rate – Fluoroscopy
36
37		Dose	Receptor
38		Level	Size (cm)	kVp	mA	mR/min	nGy/frame
39		Auto 1					Measure using pancake probe placed against
40		0.00					image receptor input window
41		0.00					Use 1.5mm Cu filtration
42		0.00
43		0.00
44
45		0.00
46		0.00
47		0.00
48		0.00
49		Auto
50		0.00
51		0.00
52		0.00
53		0.00
54			Comments:
55
56
57		Scatter and Transmission – Fluoroscopy				Pt Entrance
58				kVp	mA	mGy/min	uGy/hr		Acceptable
59			Transmission* @ 10cm				0.00		TBD
60		Scatter – Fluoro	Eye level				0.00
61		0.00	Waist level				0.00
62			Comments:
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Digital Acquisition
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	R/min	kVp	mA	R/min	kVp	mA	R/min
7		cm	10.00
8		0.00	20.00
9		0.00	30.00
10		cm	10.00
11		0.00	20.00
12		0.00	30.00
13		cm	10.00
14		0.00	20.00
15		0.00	30.00
16		cm	10.00
17		0.00	20.00
18		0.00	30.00
19		cm	10.00
20		0.00	20.00
21		0.00	30.00
22			max
23
24			Comments:
25
26
27
28
29
30
31
32		Image Receptor Entrance Exposure Rate – Digital Acquisition
33
34		Dose	Receptor
35		Level	Size (cm)	kVp	mA	mR/min	nGy/frame
36		Auto 1					Measure using pancake probe placed against
37		0.00					image receptor input window
38		0.00					Use 1.5mm Cu filtration
39		0.00
40		0.00
41
42		0.00
43		0.00
44		0.00
45		0.00
46		Auto
47		0.00
48		0.00
49		0.00
50		0.00
51			Comments:
52
53
54
55		Scatter and Transmission – Digital Acquisition
56
57				kVp	mA	uGy/hr
58		Scatter – Digital acq	Eye level			0.00
59		0.00	Waist level			0.00
60
61
62			Comments:
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Fluoroscopy
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min
7		cm	10.00
8		0.00	20.00
9		0.00	30.00
10		cm	10.00
11		0.00	20.00
12		0.00	30.00
13		cm	10.00
14		0.00	20.00
15		0.00	30.00
16		cm	10.00
17		0.00	20.00
18		0.00	30.00
19		cm	10.00
20		0.00	20.00
21		0.00	30.00
22			max
23
24			*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	NA
25			Undertable fluoroscopy units: 1 cm above table top
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)
31			Comments:
32
33
34
35		Image Receptor Entrance Exposure Rate – Fluoroscopy
36
37		Dose	Receptor
38		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame
39		Auto 1					Measure using pancake probe placed against
40		0.00					image receptor input window
41		0.00					Use 1.5mm Cu filtration
42		0.00
43		0.00
44
45		0.00
46		0.00
47		0.00
48		0.00
49		Auto
50		0.00
51		0.00
52		0.00
53		0.00
54			Comments:
55
56
57		Scatter and Transmission – Fluoroscopy				Pt Entrance
58				kVp	mA	mGy/min	uGy/hr		Acceptable
59			Transmission* @ 10cm				0.00		TBD
60		Scatter – Fluoro	Eye level				0.00
61		0.00	Waist level				0.00
62			Comments:
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Digital Acquisition
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min
7		cm	10.00
8		0.00	20.00
9		0.00	30.00
10		cm	10.00
11		0.00	20.00
12		0.00	30.00
13		cm	10.00
14		0.00	20.00
15		0.00	30.00
16		cm	10.00
17		0.00	20.00
18		0.00	30.00
19		cm	10.00
20		0.00	20.00
21		0.00	30.00
22			max
23
24			Comments:
25
26
27
28
29
30
31
32		Image Receptor Entrance Exposure Rate – Digital Acquisition
33
34		Dose	Receptor
35		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame
36		Auto 1					Measure using pancake probe placed against
37		0.00					image receptor input window
38		0.00					Use 1.5mm Cu filtration
39		0.00
40		0.00
41
42		0.00
43		0.00
44		0.00
45		0.00
46		Auto
47		0.00
48		0.00
49		0.00
50		0.00
51			Comments:
52
53
54
55		Scatter and Transmission – Digital Acquisition
56
57				kVp	mA	uGy/hr
58		Scatter – Digital acq	Eye level			0.00
59		0.00	Waist level			0.00
60
61
62			Comments:
63
64
65		Date:					Inspector:
66		Room Number					Survey ID:

## Tables
Leeds MS					Sorted by kV					0.00			Sorted by kV
1	Clearly visible					kV	Fluoro	High Level	kV	Fluoro	High Level			kV	Fluoro	High Level	kV	Fluoro	High Level
2	Visible					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3	Barely visible					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
4	Not visible					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
				#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
Huttner Grating			Leeds N3		65.00	70.00	75.00	80.00	Nominal					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
1.1	0.50		1	0.1670	0.1880	0.1670	0.1490	0.1350	0.1600					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
1.2	0.56		2	0.1480	0.1670	0.1480	0.1320	0.1200	0.1450					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
1.3	0.63		3	0.1280	0.1450	0.1280	0.1140	0.1040	0.1230					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
2.1	0.71		4	0.1090	0.1240	0.1090	0.0970	0.0880	0.1080					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
2.2	0.80		5	0.0876	0.0993	0.0876	0.0780	0.0705	0.0860					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
2.3	0.90		6	0.0749	0.0847	0.0749	0.0670	0.0609	0.0760					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3.1	1.00		7	0.0674	0.0766	0.0674	0.0599	0.0542	0.0660					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3.2	1.12		8	0.0525	0.0593	0.0525	0.0470	0.0427	0.0550					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3.3	1.25		9	0.0450	0.0512	0.0450	0.0399	0.0360	0.0450					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
4.1	1.40		10	0.0371	0.0400	0.0371	0.0347	0.0329	0.0390					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
4.2	1.60		11	0.0322	0.0347	0.0322	0.0301	0.0285	0.0330				Max	#N/A	#N/A	#N/A					Max	#N/A	#N/A	#N/A
4.3	1.80		12	0.0246	0.0269	0.0246	0.0233	0.0220	0.0270					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
5.1	2.00		13	0.0215	0.0232	0.0215	0.0201	0.0190	0.0230					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
5.2	2.24		14	0.0172	0.0186	0.0172	0.0161	0.0153	0.0180					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
5.3	2.50		15	0.0155	0.0168	0.0155	0.0145	0.0138	0.0160					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
6.1	2.80		16	0.0130	0.0140	0.0130	0.0122	0.0115	0.0130					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
6.2	3.15		17	0.0110	0.0119	0.0110	0.0103	0.0097	0.0110					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
6.3	3.55		18	0.0086	0.0093	0.0086	0.0081	0.0077	0.0090					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
7.1	4.00		19	0.0066	0.0071	0.0066	0.0062	0.0059	0.0070					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
7.2	4.50					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
7.3	5.00					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
				#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
TO.10 kV Adj	65.00	70.00	75.00	80.00					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0120	0.0119	0.0110	0.0103	0.0097					.					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0160	0.0168	0.0155	0.0145	0.0138					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0230	0.0232	0.0215	0.0201	0.0190					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0320	0.0347	0.0322	0.0301	0.0285					Max	#N/A	#N/A	#N/A					Max	#N/A	#N/A	#N/A
0.0450	0.0512	0.0450	0.0399	0.0360					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0660	0.0766	0.0674	0.0599	0.0542					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0860	0.0993	0.0876	0.0780	0.0705					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.1230	0.1450	0.1280	0.1140	0.1040					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.1600	0.1880	0.1670	0.1490	0.1350					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.2300	0.2670	0.2380	0.2140	0.1940					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.3500	0.3990	0.3600	0.3270	0.2990					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.5000	0.6180	0.5730	0.5330	0.4970					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.6600	0.7690	0.7260	0.6860	0.6490					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.9300	0.9690	0.9540	0.9360	0.9150					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
				#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
TO.10	Diameter (mm)	sqrt(Area)	1	2	3	4	5	6	7	8	9			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
A	11.10	9.837	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
B	7.90	7.001	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
C	5.60	4.963	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
D	4.00	3.545	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016		Max	#N/A	#N/A	#N/A					Max	#N/A	#N/A	#N/A
E	2.80	2.481	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016
F	2.00	1.772	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016
G	1.40	1.241	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
H	1.00	0.886	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
J	0.70	0.620	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
K	0.50	0.443	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066
L	0.35	0.310	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066
M	0.25	0.222	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066

TO.10 Values	0	1	1.5	2	2.5	3	3.5	4	4.5	5	5.5	6	6.5	7	7.5	8	8.5	9
A	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
B	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
C	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
D	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
E	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
F	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
G	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
H	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
J	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
K	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067
L	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067
M	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067

Half Value Layer
	DHEC			MUSC				HVL Graph
	Min HVL			HVL	Min HVL	Max HVL		mm Al	Avg Exp
kVp	(mm Al)		kVp	(mm Al)	(mm Al)	(mm Al)		TBD	TBD
30.00	0.3		60	2.2	1.8	2.4		TBD	TBD
40.00	0.4		70	2.6	2.2	2.9		TBD	TBD
50.00	0.5		80	3.0	2.6	3.5		TBD	TBD
51.00	1.2		90	3.5	3.1	4.0		TBD	TBD
60.00	1.3		100	3.9	3.5	4.5		TBD	TBD
70.00	1.5		110	4.3	3.9	5.1
71.00	2.1		120	5.7	4.4	5.6
80.00	2.3		130	5.2	4.8	6.1
90.00	2.5
100.00	2.7
110.00	3.0
120.00	3.2
130.00	3.5
140.00	3.8
150.00	4.1

Fluoro Patient Entrance Exposure Rate
Acrylic	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min
Attenuator	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.
(cm)
10.00	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A
20.00	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A
30.00	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A

Dig Acq Patient Entrance Exposure Rate
Acrylic	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min
Attenuator	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.
(cm)
10.00	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A
20.00	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A
30.00	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A	#N/A

## DataEntry
Patient Entrance Exposure Rate (Fluoroscopy)*
Image	Aluminum
Mode	Attenuator
#NAME?	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min
Normal	10.00
	20.00
	30.00
Mag 1	10.00
	20.00
	30.00
Mag 2	10.00
	20.00
	30.00
Mag 3	10.00
	20.00
	30.00
Mag 4	10.00
	20.00
	30.00
	max

Patient Entrance Exposure Rate (Digital Acquisition)*
Image	Aluminum
Mode	Attenuator
#NAME?	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min
Normal	10.00
	20.00
	30.00
Mag 1	10.00
	20.00
	30.00
Mag 2	10.00
	20.00
	30.00
Mag 3	10.00
	20.00
	30.00
Mag 4	10.00
	20.00
	30.00
	max

## Changelog
Date	Comment
06-Aug-14	Initial Revision