## Fluoro
1					Revision 1.2-20151125					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Fluoroscopy System Compliance Inspection					Fluoroscopy System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Number of X-ray Tubes:					Survey ID:				Number of X-ray Tubes:					Survey ID:
14					Room Number:				Accession Number:					Room Number:
15			X-Ray Generator					Site Number:
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18				Max kVp:					Max mA:					Model:				Serial Number:					Number of X-ray Tubes:
19					Software Version:					Max kVp:				Max mA:					Accesssion Number:
20			X-Ray Tube 1					Software Version:					X-Ray Generator
21				Tube Designation/Use:					Focal Spot Sizes (mm)					X-Ray Tube 1					Manufacturer:
22				Manufacture Date:					Large:					Tube Designation/Use:				Focal Spot Sizes (mm)					Model:
23			Insert					Small:					Manufacture Date:				Large:					Max kVp:
24				Manufacturer:					Micro:					Insert					Small:					Manufacture Date:
25				Model:					Manufacturer:				Micro:					Serial Number:
26				Serial Number:					Filtration					Model:					Max mA:
27			Housing					Inherent:					Serial Number:				Filtration					Software Version:
28				Manufacturer:					Added:					Housing					Inherent:					X-Ray Tube 1
29				Model:					Manufacturer:				Added:					Tube Designation/Use:
30				Serial Number:					Model:					Manufacture Date:
31					Serial Number:					Manufacturer:
32			X-Ray Tube 2					Model:
33				Tube Designation/Use:					Focal Spot Sizes (mm)					X-Ray Tube 2					Serial Number:
34				Manufacture Date:					Large:					Tube Designation/Use:				Focal Spot Sizes (mm)					Manufacturer:
35			Insert					Small:					Manufacture Date:				Large:					Model:
36				Manufacturer:					Micro:					Insert					Small:					Serial Number:
37				Model:					Manufacturer:				Micro:					Large:
38				Serial Number:					Filtration					Model:					Small:
39					Inherent:					Serial Number:				Filtration					Micro:
40				Manufacturer:					Added:					Housing					Inherent:					Inherent:
41				Model:					Manufacturer:				Added:					Added:
42				Serial Number:					Model:					X-Ray Tube 2
43					Serial Number:					Tube Designation/Use:
44					Manufacture Date:
45					Manufacturer:
46					Inspection Results					Model:
47					Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
48					Labels, Notices, Postings		Tube Number:					Manufacturer:
49					DHEC Registration sticker is present, clearly visible and legible					**Input only one of the following				Model:
50					DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					numbers or names				Serial Number:
51					Radiation warning label posted on the generator control panel					1	1			Large:
52					Tube loading charts for each focal spot are posted					2	2			Small:
53					Tube loading charts are correct for installed tube(s).					3	3			Micro:
54					Operator manuals are available.					AP	AP			Inherent:
55					Monthly radiation monitoring reports are posted.					Front	Front			Added:
56					Fluoroscopy Systems			Frontal	Frontal			Tube Number:
57					Interlock systems function to prevent fluoro while II is parked					Lat	Lat			Tube Designation:		Change	1
58					The minimum Source Skin Distance > 38 cm (20 cm for surgical applications)					Lateral	Lateral					Change	2
59					Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels)					Change	3
60					The useful beam is intercepted by the image assembly at any SID					Change	AP
61					Fluoroscopy patient dose charts are posted					Change	Front
62					1	Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Change	Frontal
63					1	Max Entrance Exposure Rate <= 176 mGy/min (high level)					Change	Lat
64					3	High contrast resolution within MUSC specifications					Change	Lateral
65		Date:					Inspector:	Eugene Mah			Low contrast resolution within MUSC specifications
66		Room Number					Survey ID:			1	Transmission through II < 2 uGy/hr for each mGy/min at patient entrance
1					Medical University of South Carolina			Beam limiting device available (functioning) to restrict the x-ray field					Receptor Size
2					Inspection Results					Fluoroscopy System Compliance Inspection			The useful beam is centered on the image receptor					Normal:
3		Rule Number			Labels, Notices, Postings					Compliance				Collimator edges are visible at full open					Mag 1:
4		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD				Elapsed timer (5 min maximum) functional					Mag 2:
5		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD				Fluoroscopic kV and mA meter present and functional					Mag 3:
6		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Radiation Safety					Mag 4:
7				Tube loading charts for each focal spot are posted					TBD				Continuous pressure on exposure switch required to make an exposure with selected tube					Pt Dose Measured Fluoro
8				Tube loading charts are correct for installed tube(s).					TBD				Operator has full visibility of the exposure factors and patient					Manual/Auto:
9				Operator manuals are available.					TBD				Lead aprons available					Lucite/Al:		Change	1
10				Monthly radiation monitoring reports are posted.					TBD				Lead gloves available					# dose rates:
11					Properly designed and installed apron rack is present					High level present:
12					Radiation Safety					Documentation of annual protective apparel integrity inspection is available					cm/in:		Change	cm
13				Continuous pressure on exposure switch required to make an exposure with selected tube					TBD				Patient restraint devices available					Auto/mA:		Change	Auto
14				Operator has full visibility of the exposure factors and patient					TBD				Gonadal shielding is present and functional					Technique 1:
15				Lead aprons available					TBD				Radiation warning light at room entrance is functional with selected tube					Technique 2:
16				Lead gloves available					TBD					Physical Inspection					Technique 3:
17				Properly designed and installed apron rack is present					TBD				System clean and free of debris					Maximum Dose Rates Fluoro
18				Documentation of annual protective apparel integrity inspection is available					TBD				All cables free from excessive wear or restraints					kV:
19				Patient restraint devices available					TBD				System free of missing or damaged components					mA:
20				Gonadal shielding is present and functional					TBD				Tube support movements and locks are acceptable					mGy/min:			0.00
21				Radiation warning light at room entrance is functional with selected tube					TBD				All other movements and locks functional					kV:
22					Automatic Brightness Control / Display Monitor					mA:
23					Fluoroscopy Systems					Automatic Brightness Control maintains image brightness for varying attenuation					mGy/min:			0.00
24		DHEC RHB 4.9.2.1.2		Interlock systems function to prevent fluoro while II is parked					TBD				Display monitor is free of artifacts, shadows, bright spots, and burn-in					kV:
25		DHEC RHB 4.9.1		The minimum Source Skin Distance > 38 cm (20 cm for surgical applications)					TBD				Display monitor is free of distortion					mA:
26		DHEC RHB 4.9.8		Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels)					TBD				Display monitor is free of excessive persistence (after-image)					mGy/min:			0.00
27		DHEC RHB 4.9.2.1.2		The useful beam is intercepted by the image assembly at any SID					TBD				Display monitor is free of static or rolling noise bands					Maximum Dose Rates Pulse/Dig
28		DHEC RHB 4.9.4.3.6		Fluoroscopy patient dose charts are posted					TBD				The last image hold (LIH) display is free of jitter or abnormal appearance					kV:
29		DHEC RHB 4.9.4.2.1		Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					YES					mA:
30		DHEC RHB 4.9.4.2.1.2		Max Entrance Exposure Rate <= 176 mGy/min (high level)					YES					mGy/min:
31				High contrast resolution within MUSC specifications					NA					Inspection Results					kV:
32				Low contrast resolution within MUSC specifications					TBD					mA:
33		DHEC RHB 4.9.5.1		Transmission through II < 2 uGy/hr for each mGy/min at patient entrance					YES			Entrance Exposure Rate				Type 1 for manual, leave blank for auto			Auto	Auto	Auto/mA mode			mGy/min:
34		DHEC RHB 4.9.2.2.1		beam limiting device available (functioning) to restrict the x-ray field					TBD					1	1	Type 1 if using Lucite, leave blank if using Al					kV:
35				The useful beam is centered on the image receptor					TBD					How many dose rate settings are present (exclusing High Level)?					mA:
36		DHEC RHB 4.9.2.2.1		Collimator edges are visible at full open					TBD					Is High Level provided? (1=Yes, 2=No)					mGy/min:
37		DHEC RHB 4.9.7		Elapsed timer (5 min maximum) functional					TBD			Blank					HVL
38		DHEC RHB 4.9.6		Fluoroscopic kV and mA meter present and functional					TBD			cm	cm	Unit for image receptor size					HVL Measured kV:		Change	80.8
39		DHEC RHB 4.3.5		The Measured HVL meeds DHEC minimum requirements					YES			Patient Entrance Exposure Rate (Fluoroscopy)*					HVL Measured:		Change	6.12
40				Kerma area product display is accurate to within 20%					NA			Image	Acrylic					Pt Dose Measured Dig Acq
41				Air kerma display is accurate to within 20%					NA			Mode	Attenuator	Auto 1					Auto				Technique 1:
42					(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min		Technique 2:
43					Automatic Brightness Control / Display Monitor					Normal	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Technique 3:
44				Automatic Brightness Control maintains image brightness for varying attenuation					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Receptor Entr Exp Fluoro
45				Display monitor is free of artifacts, shadows, bright spots, and burn-in					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Normal
46				Display monitor is free of distortion					TBD			Mag 1	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 1
47				Display monitor is free of excessive persistance (after-image)					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 2
48				Display monitor is free of static or rolling nosie bands					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 3
49				The last image hold (LIH) display is free of jitter or abnormal appearance					TBD			Mag 2	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 4
50					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Normal
51					Physical Inspection					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 1
52				System clean and free of debris					TBD			Mag 3	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 2
53				All cables free from excessive wear or restraints					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 3
54				System free of missing or damaged components					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 4
55				Tube support movements and locks are acceptable					TBD			Mag 4	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Normal
56				All other movements and locks functional					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 1
57					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 2
58					max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 3
59					Last Year:					Auto Mag 4
60					*Measured at:		Acceptable:	YES		Acceptable:	YES		Acceptable:	YES		Receptor Entr Exp Dig Acq
61					Undertable fluoroscopy units: 1 cm above table top					Auto 1 Normal
62					Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Auto 1 Mag 1
63					C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					Auto 1 Mag 2
64					Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					Auto 1 Mag 3					0.07
65		Date:					Inspector:	Eugene Mah			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Auto 1 Mag 4					0.24
66		Room Number					Survey ID:					Max Entrance Exposure Rate < = 176 mGy/min (high level)					Normal
1					Medical University of South Carolina			Comments:					Mag 1
2					Inspection Results					Fluoroscopy System Compliance Inspection			Do not exceed 390 chars		0					Mag 2
3					Half Value Layer Measurement					New:					Mag 3
4					Mag 4
5				Set kVp:	80.80					Patient Entrance Exposure Rate – Pulse Fluroscopy					Mode:	1	Enter 1 for Pulse fluoro, 2 for digital acquisition mode					Auto Normal
6					Image	Acrylic					Auto Mag 1
7				Al (mm)	mGy/min	mGy/min		Measured HVL:	6.12					Mode	Attenuator	Auto 1					Auto				Auto Mag 2
8				0.0	85.14	85.20		HVL Acceptable:	YES					(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min		Auto Mag 3
9				5.5	45.40	45.37					Normal	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 4
10				6.0	43.12	43.13			ACCEPTABLE RESULTS (mm Al)					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Fluoro Scatter/Transmission
11				6.5	40.92	40.96			Min acceptable HVL at 80.8 kVp:	2.60					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Transmission:			0.00
12					Desired HVL at 80.8 kVp:	3.10				Mag 1	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (eyes):
13					Criteria:	MUSC Standard					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (waist):
14					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (eyes):
15					Mag 2	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (waist):
16					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Dig Acq Scatter
17					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (eyes):
18					Mag 3	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (waist):
19					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (eyes):
20					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (waist):
21					Mag 4	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Leeds Phantom
22					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		kVp:		Change	70
23					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Huttner Normal:
24					max					Huttner Mag 1:
25					Last Year:					Huttner Mag 2:
26					*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	NA		Huttner Mag 3:
27					Undertable fluoroscopy units: 1 cm above table top					Huttner Mag 4:
28					Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					N3 Normal:
29					C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					N3 Mag 1:
30					Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					N3 Mag 2
31					Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					N3 Mag 3
32					Max Entrance Exposure Rate < = 176 mGy/min (high level)					N3 Mag 4:
33					Comments:					Field Size:
34					Do not exceed 390 chars		0					SID:
35					New:					Front:
36					Back:
37					Half Value Layer					Left:
38					2	Use DHEC standards? (Type 1 to base acceptance on DHEC criteria)					Right:
39					TO.10.A
40					kV	mA	mm Al	mGy/min		Data Processing					TO.10.B
41					80.8	122.8	0.0	85.14		mm Al	Avg Exp	DATA(1/2)	Data Al	HVL			TO.10.C
42					0.0	85.20		0.00	85.17	42.59	6.00	6.12			TO.10.D
43					5.5	45.40		5.50	45.39	2.80	6.50				TO.10.E
44					5.5	45.37		6.00	43.13	0.54					TO.10.F
45					6.0	43.12		6.50	40.94	1.65	Data Exp	LN(EXP)			TO.10.G
46					6.0	43.13					43.13	3.76			TO.10.H
47					6.5	40.92					40.94	3.71			TO.10.J
48					6.5	40.96		6.12	42.59					TO.10.K
49					Acceptable (mm Al)			HVL Acceptable:	YES			TO.10.L
50					Low	2.60					TO.10.M
51					Desired	3.10		Previous HVL @ kV:				TO.10.A
52					High	3.50					TO.10.B
53					TO.10.C
54					Image Receptor Entrance Exposure Rate – Fluoroscopy					TO.10.D
55					TO.10.E
56					Dose	Receptor					Previous Year					TO.10.F
57					Level	Size (cm)	kVp	mA	uGy/min	nGy/frame	nGy/frame					TO.10.G
58					Auto 1					Measure using pancake probe placed against				TO.10.H
59					image receptor input window				TO.10.J
60					Use 1.5mm Cu filtration				TO.10.K
61					TO.10.L
62					TO.10.M
63					TO.10.A
64					TO.10.B
65		Date:					Inspector:	Eugene Mah					TO.10.C
66		Room Number					Survey ID:					TO.10.D
1					Medical University of South Carolina					TO.10.E
2					Inspection Results					Fluoroscopy System Compliance Inspection		Auto					TO.10.F
3		Patient Entrance Exposure Rate – Fluoroscopy					TO.10.G
4		Image	Acrylic					TO.10.H
5		Mode	Attenuator	Auto 1					Auto					TO.10.J
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min					TO.10.K
7		cm	10.00			0.00			0.00			0.00					TO.10.L
8			20.00			0.00			0.00			0.00			Comments:					TO.10.M
9			30.00			0.00			0.00			0.00					0					TO.10.A
10		cm	10.00			0.00			0.00			0.00			New:					TO.10.B
11			20.00			0.00			0.00			0.00					TO.10.C
12			30.00			0.00			0.00			0.00		Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy					TO.10.D
13		cm	10.00			0.00			0.00			0.00					TO.10.E
14			20.00			0.00			0.00			0.00		Dose	Receptor					Previous Year					TO.10.F
15			30.00			0.00			0.00			0.00		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame	nGy/frame					TO.10.G
16		cm	10.00			0.00			0.00			0.00		Auto 1					Measure using pancake probe placed against				TO.10.H
17			20.00			0.00			0.00			0.00					image receptor input window				TO.10.J
18			30.00			0.00			0.00			0.00					Use 1.5mm Cu filtration				TO.10.K
19		cm	10.00			0.00			0.00			0.00					TO.10.L
20			20.00			0.00			0.00			0.00					TO.10.M
21			30.00			0.00			0.00			0.00					TO.10.A
22			max			0.00			0.00			0.00					TO.10.B
23					TO.10.C
24			*Measured at:		Acceptable:	YES		Acceptable:	YES		Acceptable:	YES					TO.10.D
25			Undertable fluoroscopy units: 1 cm above table top					TO.10.E
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Auto					TO.10.F
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					TO.10.G
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					TO.10.H
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					TO.10.J
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)					TO.10.K
31		Comments:					TO.10.L
32					Comments:					TO.10.M
33					Do not exceed 390 chars		0					Comments
34					New:					Pt Entr Exp Fluoro:
35		Image Receptor Entrance Exposure Rate – Fluoroscopy					Pt Entr Exp Dig Acq:
36					Scatter and Transmission					Receptor Entr Exp Fluoro:
37		Dose	Receptor					Pt Entrance					Receptor Entr Exp Dig Acq:
38		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame					kVp	mA	mGy/min	uGy/hr	Last Year		Acceptable				Scatter/Transmission:
39		Auto 1					Measure using pancake probe placed against					Transmission* @ 10cm			0.00	0.00			YES				Leeds Phantom:
40					image receptor input window					Unshielded	Shielded	Unshielded	Shielded	Unshielded	Shielded			TO.10:
41					Use 1.5mm Cu filtration					uGy/min	uGy/min	uGy/hr	uGy/hr	uGy/hr	uGy/hr			Beam Limitation:
42					Scatter – Fluoro	Eye level					Comment Page 1		Change	The patient skin entrance exposure for AP and lateral DSA exposures can get very high (up to 1.3 Gy/min). Operators should be aware of this and that unnecessary prolonged or multiple cine exposures may impart sufficient radiation dose to cause radiation injury to the patient.
43					Waist level
44					Scatter - Pulse	Eye level					Change	Scatter radiation at eye level 1 m away from the beam center was 0.5 mGy/h and 0.1 mGy/h at waist level during AP plane fluoroscopy. When the hanging shield was in place, scatter radiation level was 0.03 uGy/h at eye and 0.04 uGy/h at waist level.
45					Waist level
46					*Measured with 8” Lucite in beam					Change	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
47					Largest field size mode, highest dose rate mode
48					Criteria:	Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure					Change	Scatter radiation 1 m away from the beam center was 0.2 mGy/h at eye level and 0.3 mGy/h at waist level during lateral plane fluoroscopy. When the hanging shield was in place, scatter radiation level was 0.01 at eye and waist level.
49		Auto					Comments:
50					Do not exceed 390 chars		0					Change	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
51					New:
52					Change	Scatter radiation during AP plane DSA exposures at eye level 1 m away from the beam center without the hanging shield in place was 4.4 mGy/h and 1.8 mGy/h at waist level. With the hanging shield in place, scatter radiation levels were 0.3 at eye and waist level respectively.
53					Kerma Area Product/Cumulative Air Kerma
54		Comments:					0	Source-PDC distance (cm)				Source-Isocenter					Change	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
55					0	Source-Ref Pt distance (cm)				Table-Isocenter
56					#DIV/0!	CAK Reference Point correction factor				RefPt-Isocenter					Change	Scatter radiation during lateral plane DSA exposures at eye level 1 m away from the beam center without the hanging shield in place was 2.5 mGy/h and 3.4 mGy/h at waist level. With the hanging shield in place, scatter radiation level was 0.13 mGy/h at eye and waist level.
57		Scatter and Transmission – Fluoroscopy				Pt Entrance					2	Apply criteria
58				kVp	mA	mGy/min	uGy/hr	Unshielded	Shielded		Acceptable					Start KAP	End KAP		Start AK	End AK					Change	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
59			Transmission* @ 10cm			0	0.00	uGy/hr	uGy/hr		YES			kV	mA	Gy·cm²		KAP	mGy		CAK
60		Scatter – Fluoro	Eye level					Change	AP plane tube replaced 21-May-2018
61			Waist level
62		Comments:
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection				Radcal PDC
3		Patient Entrance Exposure Rate – Pulse Fluroscopy					kV	mA	Sample	Time	mGy	mGy/Min	µGy·m²	µGy·m²/Min
4		Image	Acrylic
5		Mode	Attenuator	Auto 1					Auto
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min
7		cm	10.00			0.00			0.00			0.00
8			20.00			0.00			0.00			0.00
9			30.00			0.00			0.00			0.00
10		cm	10.00			0.00			0.00			0.00
11			20.00			0.00			0.00			0.00
12			30.00			0.00			0.00			0.00		Indicated KAP µGy·m²	Measured KAP µGy·m²				Indicated AK @ Ref Pt	Measured AK @ Ref Pt
13		cm	10.00			0.00			0.00			0.00				KAP Deviation					CAK Deviation					Comment Page 2
14			20.00			0.00			0.00			0.00					#DIV/0!	#DIV/0!
15			30.00			0.00			0.00			0.00					Acceptable:			#DIV/0!	#DIV/0!	Acceptable:
16		cm	10.00			0.00			0.00			0.00					NA			#DIV/0!	#DIV/0!	NA
17			20.00			0.00			0.00			0.00					#DIV/0!	#DIV/0!
18			30.00			0.00			0.00			0.00					#DIV/0!	#DIV/0!
19		cm	10.00			0.00			0.00			0.00					#DIV/0!	#DIV/0!
20			20.00			0.00			0.00			0.00					#DIV/0!	#DIV/0!
21			30.00			0.00			0.00			0.00					#DIV/0!	#DIV/0!
22			max					Criteria:	Measured KAP/AK should be within 20% of the displayed values
23					Leeds Phantom
24		Comments:
25					Receptor Size	MS1	MS3	MS4	Huttner	N3	kVp:	70	70
26					MS1-4		Huttner: Enter col.row
27					1	Clearly visible	N3: Enter # of visible disks
28					2	Visible
29					3	Barely visible
30					4	Not visible
31					Acceptable:	TBD	TBD
32		Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy					Comments:
33					Do not exceed 390 chars		0
34		Dose	Receptor					New:
35		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame
36		Auto 1					Measure using pancake probe placed against					TO.10		Details Detected
37					image receptor input window					Row
38					Use 1.5mm Cu filtration					A
39					B
40					C
41					D
42					E
43					F
44					G
45					H
46		Auto					J
47					K
48					L
49					M
50
51		Comments:					Comments:
52					Do not exceed 390 chars		0
53					New:
54
55		Scatter and Transmission – Pulse Fluoroscopy					Field Size
56					Unshielded	Shielded					Horizontal					Type 1 for circular field, 2 for square field
57				kVp	mA	uGy/hr	uGy/hr					Vertical
58		Scatter - Pulse	Eye level					Diagonal
59			Waist level					Acceptable	NA	NA	NA	NA	NA
60					Criteria:	Measured field size is at least 85% of nominal field size
61					Useful Beam Limitation
62		Comments:					SID (cm)
63
64					Measured	Difference	Sum	Last Year			T	B	L	R
65		Date:					Inspector:	Eugene Mah		Top					Visible
66		Room Number					Survey ID:			Bottom					X-ray
1					Medical University of South Carolina		Left					Acceptable	SID
2					Inspection Results					Fluoroscopy System Compliance Inspection		Right					SOD
3		Leeds Test Objects					High contrast resolution		Low contrast resolution					Criteria:	Difference between x-ray field and visible image area is less than 2% SID					SID/SOD
4			Receptor Size	MS1	MS3	MS4	lp/mm	Acceptable	N3	Acceptable					Sum of length and width differences is less than 4% SID
5					Comments:
6					Do not exceed 390 chars		0
7					New:
8
9
10			Criteria:	Resolution: >1.0 lp/mm Normal mode, >1.25 lp/mm Mag 1, >1.4 lp/mm Mag 2, > 2 lp/mm Mag 3, > 2.5 lp/mm Mag 4					Comments Page 1
11				N3: Minimum contrast level < 0.04 in all II modes					Previous Year Comments
12		Comments:					Comments:	The patient skin entrance exposure for AP and lateral DSA exposures can get very high (up to 1.3 Gy/min). Operators should be aware of this and that unnecessary prolonged or multiple cine exposures may impart sufficient radiation dose to cause radiation injury to the patient.
13					Do not exceed 390 chars		276			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
14					Additional Comments:	Scatter radiation at eye level 1 m away from the beam center was 0.5 mGy/h and 0.1 mGy/h at waist level during AP plane fluoroscopy. When the hanging shield was in place, scatter radiation level was 0.03 uGy/h at eye and 0.04 uGy/h at waist level.
15					Do not exceed 390 chars		249			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
16		TO.10 Contrast Resolution					Additional Comments:	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
17			Contrast Detail					Threshold Index					Do not exceed 390 chars		205			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
18		Row					Additional Comments:	Scatter radiation 1 m away from the beam center was 0.2 mGy/h at eye level and 0.3 mGy/h at waist level during lateral plane fluoroscopy. When the hanging shield was in place, scatter radiation level was 0.01 at eye and waist level.
19		A					Do not exceed 390 chars		234			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
20		B					Additional Comments:	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
21		C					Do not exceed 390 chars		205			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
22		D					Additional Comments:	Scatter radiation during AP plane DSA exposures at eye level 1 m away from the beam center without the hanging shield in place was 4.4 mGy/h and 1.8 mGy/h at waist level. With the hanging shield in place, scatter radiation levels were 0.3 at eye and waist level respectively.
23		E					Do not exceed 390 chars		275
24		F					Additional Comments:	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
25		G					Do not exceed 390 chars		205
26		H					Additional Comments:	Scatter radiation during lateral plane DSA exposures at eye level 1 m away from the beam center without the hanging shield in place was 2.5 mGy/h and 3.4 mGy/h at waist level. With the hanging shield in place, scatter radiation level was 0.13 mGy/h at eye and waist level.
27		J					Do not exceed 390 chars		273
28		K					Additional Comments:	It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile shield should also be used to reduce scatter exposure to other room personnel.
29		L					Do not exceed 390 chars		205
30		M					Additional Comments:	AP plane tube replaced 21-May-2018
31		Comments:					Do not exceed 390 chars		34
32					Additional Comments:
33					Do not exceed 390 chars		0
34					Additional Comments:
35		Field Size					Do not exceed 390 chars		0
36					Additional Comments:
37			Horizontal					Do not exceed 390 chars		0
38			Vertical					Additional Comments:
39			Diagonal					Do not exceed 390 chars		0
40			Acceptable	NA	NA	NA	NA	NA					Additional Comments:
41			Criteria:	Measured field size is at least 85% of nominal field size					Do not exceed 390 chars		0
42					Additional Comments:
43		Useful Beam Limitation					Do not exceed 390 chars		0
44			SID (cm)					Additional Comments:
45				Measured	Difference	Sum					Do not exceed 390 chars		0
46			Top					Additional Comments:
47			Bottom					Acceptable					Do not exceed 390 chars		0
48			Left					Additional Comments:
49			Right					Do not exceed 390 chars		0
50			Criteria:	Difference between x-ray field and visible image area is less than 2% SID
51				Sum of length and width differences is less than 4% SID
52		Comments:					Comments Page 2
53					Previous Year Comments
54					Comments:
55					Do not exceed 390 chars		0
56					Additional Comments:
57					Do not exceed 390 chars		0
58					Additional Comments:
59					Do not exceed 390 chars		0
60					Additional Comments:
61					Do not exceed 390 chars		0
62					Additional Comments:
63					Do not exceed 390 chars		0
64					Additional Comments:
65		Date:					Inspector:	Eugene Mah			Do not exceed 390 chars		0
66		Room Number					Survey ID:				Additional Comments:
1					Medical University of South Carolina			Do not exceed 390 chars		0
2					Comments Page 1					Fluoroscopy System Compliance Inspection			Additional Comments:
3					Do not exceed 390 chars		0
4			The patient skin entrance exposure for AP and lateral DSA exposures can get very high (up to 1.3 Gy/min). Operators should be aware					Additional Comments:
5			of this and that unnecessary prolonged or multiple cine exposures may impart sufficient radiation dose to cause radiation injury					Do not exceed 390 chars		0
6			to the patient.					Additional Comments:
7			Scatter radiation at eye level 1 m away from the beam center was 0.5 mGy/h and 0.1 mGy/h at waist level during AP plane fluoroscopy.					Do not exceed 390 chars		0
8			When the hanging shield was in place, scatter radiation level was 0.03 uGy/h at eye and 0.04 uGy/h at waist level.					Additional Comments:
9					Do not exceed 390 chars		0
10			It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile					Additional Comments:
11			shield should also be used to reduce scatter exposure to other room personnel.					Do not exceed 390 chars		0
12					Additional Comments:
13			Scatter radiation 1 m away from the beam center was 0.2 mGy/h at eye level and 0.3 mGy/h at waist level during lateral plane					Do not exceed 390 chars		0
14			fluoroscopy. When the hanging shield was in place, scatter radiation level was 0.01 at eye and waist level.					Additional Comments:
15					Do not exceed 390 chars		0
16			It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile					Additional Comments:
17			shield should also be used to reduce scatter exposure to other room personnel.					Do not exceed 390 chars		0
18					Additional Comments:
19			Scatter radiation during AP plane DSA exposures at eye level 1 m away from the beam center without the hanging shield in place was					Do not exceed 390 chars		0
20			4.4 mGy/h and 1.8 mGy/h at waist level. With the hanging shield in place, scatter radiation levels were 0.3 at eye and waist					Additional Comments:
21			level respectively.					Do not exceed 390 chars		0
22			It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile					Additional Comments:
23			shield should also be used to reduce scatter exposure to other room personnel.					Do not exceed 390 chars		0
24					Additional Comments:
25			Scatter radiation during lateral plane DSA exposures at eye level 1 m away from the beam center without the hanging shield in place					Do not exceed 390 chars		0
26			was 2.5 mGy/h and 3.4 mGy/h at waist level. With the hanging shield in place, scatter radiation level was 0.13 mGy/h at eye
27			and waist level.
28			It is recommended that the hanging shield be used whenever practical to minimize exposure to the operator. The rolling mobile
29			shield should also be used to reduce scatter exposure to other room personnel.
30
31			AP plane tube replaced 21-May-2018
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Comments Page 2					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Radiation Output – Fluoroscopy					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Radiation Output - Pulse Fluoroscopy					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Contrast Resolution					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
Half Value Layer
mm Al
Exposure Rate (mGy/min)
Radiation Output
kV
mGy/mA-min
Fluoro Patient Entrance Exposure
Attenuator thickness (cm)
mGy/min
Radiation Output
kV
mGy/mA-min
Fluoro Patient Entrance Exposure
Attenuator thickness (cm)
mGy/min
Contrast Detail Diagram
Detail diameter (mm)
Threshold Detection Index
sqrt(Detail area)

## ExpCharts
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Fluoroscopy
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	R/min	kVp	mA	R/min	kVp	mA	R/min
7		cm	10.00			0.00			0.00			0.00
8		0.00	20.00			0.00			0.00			0.00
9		0.00	30.00			0.00			0.00			0.00
10		cm	10.00			0.00			0.00			0.00
11		0.00	20.00			0.00			0.00			0.00
12		0.00	30.00			0.00			0.00			0.00
13		cm	10.00			0.00			0.00			0.00
14		0.00	20.00			0.00			0.00			0.00
15		0.00	30.00			0.00			0.00			0.00
16		cm	10.00			0.00			0.00			0.00
17		0.00	20.00			0.00			0.00			0.00
18		0.00	30.00			0.00			0.00			0.00
19		cm	10.00			0.00			0.00			0.00
20		0.00	20.00			0.00			0.00			0.00
21		0.00	30.00			0.00			0.00			0.00
22			max			0.00			0.00			0.00
23
24			*Measured at:		Acceptable:	YES		Acceptable:	YES		Acceptable:	YES
25			Undertable fluoroscopy units: 1 cm above table top
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)
31			Comments:
32
33
34
35		Image Receptor Entrance Exposure Rate – Fluoroscopy
36
37		Dose	Receptor
38		Level	Size (cm)	kVp	mA	mR/min	nGy/frame
39		Auto 1					Measure using pancake probe placed against
40		0.00					image receptor input window
41		0.00					Use 1.5mm Cu filtration
42		0.00
43		0.00
44
45		0.00
46		0.00
47		0.00
48		0.00
49		Auto
50		0.00
51		0.00
52		0.00
53		0.00
54			Comments:
55
56
57		Scatter and Transmission – Fluoroscopy				Pt Entrance	Unshielded	Shielded
58				kVp	mA	mGy/min	uGy/hr	uGy/hr		Acceptable
59			Transmission* @ 10cm			0.00				YES
60		Scatter – Fluoro	Eye level
61		0.00	Waist level
62			Comments:
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Pulse Fluroscopy
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	R/min	kVp	mA	R/min	kVp	mA	R/min
7		cm	10.00			0.00			0.00			0.00
8		0.00	20.00			0.00			0.00			0.00
9		0.00	30.00			0.00			0.00			0.00
10		cm	10.00			0.00			0.00			0.00
11		0.00	20.00			0.00			0.00			0.00
12		0.00	30.00			0.00			0.00			0.00
13		cm	10.00			0.00			0.00			0.00
14		0.00	20.00			0.00			0.00			0.00
15		0.00	30.00			0.00			0.00			0.00
16		cm	10.00			0.00			0.00			0.00
17		0.00	20.00			0.00			0.00			0.00
18		0.00	30.00			0.00			0.00			0.00
19		cm	10.00			0.00			0.00			0.00
20		0.00	20.00			0.00			0.00			0.00
21		0.00	30.00			0.00			0.00			0.00
22			max
23
24			Comments:
25
26
27
28
29
30
31
32		Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy
33
34		Dose	Receptor
35		Level	Size (cm)	kVp	mA	mR/min	nGy/frame
36		Auto 1					Measure using pancake probe placed against
37		0.00					image receptor input window
38		0.00					Use 1.5mm Cu filtration
39		0.00
40		0.00
41
42		0.00
43		0.00
44		0.00
45		0.00
46		Auto
47		0.00
48		0.00
49		0.00
50		0.00
51			Comments:
52
53
54
55		Scatter and Transmission – Pulse Fluoroscopy
56					Unshielded	Shielded
57				kVp	mA	uGy/hr	uGy/hr
58		Scatter - Pulse	Eye level
59		0.00	Waist level
60
61
62			Comments:
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:

## Tables
Leeds MS					Patient Entrance Exposure Rate (Fluoroscopy)*			Sorted by kV					Patient Entrance Exposure Rate – Pulse Fluroscopy			Sorted by kV
1	Clearly visible					kV	Fluoro	High Level	kV	Fluoro	High Level			kV	Fluoro	High Level	kV	Fluoro	High Level
2	Visible					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3	Barely visible					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
4	Not visible					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
				#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
Huttner Grating			Leeds N3		65.00	70.00	75.00	80.00	Nominal					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
1.1	0.50		1	0.1670	0.1880	0.1670	0.1490	0.1350	0.1600					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
1.2	0.56		2	0.1480	0.1670	0.1480	0.1320	0.1200	0.1450					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
1.3	0.63		3	0.1280	0.1450	0.1280	0.1140	0.1040	0.1230					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
2.1	0.71		4	0.1090	0.1240	0.1090	0.0970	0.0880	0.1080					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
2.2	0.80		5	0.0876	0.0993	0.0876	0.0780	0.0705	0.0860					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
2.3	0.90		6	0.0749	0.0847	0.0749	0.0670	0.0609	0.0760					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3.1	1.00		7	0.0674	0.0766	0.0674	0.0599	0.0542	0.0660					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3.2	1.12		8	0.0525	0.0593	0.0525	0.0470	0.0427	0.0550					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
3.3	1.25		9	0.0450	0.0512	0.0450	0.0399	0.0360	0.0450					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
4.1	1.40		10	0.0371	0.0400	0.0371	0.0347	0.0329	0.0390					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
4.2	1.60		11	0.0322	0.0347	0.0322	0.0301	0.0285	0.0330				Max	#N/A	#N/A	#N/A					Max	#N/A	#N/A	#N/A
4.3	1.80		12	0.0246	0.0269	0.0246	0.0233	0.0220	0.0270					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
5.1	2.00		13	0.0215	0.0232	0.0215	0.0201	0.0190	0.0230					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
5.2	2.24		14	0.0172	0.0186	0.0172	0.0161	0.0153	0.0180					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
5.3	2.50		15	0.0155	0.0168	0.0155	0.0145	0.0138	0.0160					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
6.1	2.80		16	0.0130	0.0140	0.0130	0.0122	0.0115	0.0130					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
6.2	3.15		17	0.0110	0.0119	0.0110	0.0103	0.0097	0.0110					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
6.3	3.55		18	0.0086	0.0093	0.0086	0.0081	0.0077	0.0090					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
7.1	4.00		19	0.0066	0.0071	0.0066	0.0062	0.0059	0.0070					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
7.2	4.50					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
7.3	5.00					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
				#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
TO.10 kV Adj	65.00	70.00	75.00	80.00					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0120	0.0119	0.0110	0.0103	0.0097					.					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0160	0.0168	0.0155	0.0145	0.0138					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0230	0.0232	0.0215	0.0201	0.0190					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0320	0.0347	0.0322	0.0301	0.0285					Max	#N/A	#N/A	#N/A					Max	#N/A	#N/A	#N/A
0.0450	0.0512	0.0450	0.0399	0.0360					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0660	0.0766	0.0674	0.0599	0.0542					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.0860	0.0993	0.0876	0.0780	0.0705					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.1230	0.1450	0.1280	0.1140	0.1040					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.1600	0.1880	0.1670	0.1490	0.1350					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.2300	0.2670	0.2380	0.2140	0.1940					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.3500	0.3990	0.3600	0.3270	0.2990					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.5000	0.6180	0.5730	0.5330	0.4970					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.6600	0.7690	0.7260	0.6860	0.6490					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
0.9300	0.9690	0.9540	0.9360	0.9150					#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
				#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
TO.10	Diameter (mm)	sqrt(Area)	1	2	3	4	5	6	7	8	9			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
A	11.10	9.837	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
B	7.90	7.001	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
C	5.60	4.963	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			#N/A	#N/A	#N/A					#N/A	#N/A	#N/A
D	4.00	3.545	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016		Max	#N/A	#N/A	#N/A					Max	#N/A	#N/A	#N/A
E	2.80	2.481	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016
F	2.00	1.772	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016
G	1.40	1.241	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
H	1.00	0.886	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
J	0.70	0.620	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
K	0.50	0.443	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066
L	0.35	0.310	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066
M	0.25	0.222	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066

TO.10 Values	0	1	1.5	2	2.5	3	3.5	4	4.5	5	5.5	6	6.5	7	7.5	8	8.5	9
A	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
B	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
C	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
D	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
E	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
F	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
G	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
H	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
J	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
K	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067
L	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067
M	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067

Half Value Layer
	DHEC			MUSC				HVL Graph
	Min HVL			HVL	Min HVL	Max HVL		mm Al	Avg Exp
kVp	(mm Al)		kVp	(mm Al)	(mm Al)	(mm Al)		5.5	45.39
30.00	0.3		60	2.2	1.8	2.4		6.0	43.13
40.00	0.4		70	2.6	2.2	2.9		6.1	42.59
50.00	0.5		80	3.0	2.6	3.5		6.5	40.94
51.00	1.2		90	3.5	3.1	4.0		#N/A	#N/A
60.00	1.3		100	3.9	3.5	4.5		#N/A	#N/A
70.00	1.5		110	4.3	3.9	5.1
71.00	2.1		120	5.7	4.4	5.6
80.00	2.3		130	5.2	4.8	6.1
90.00	2.5
100.00	2.7
110.00	3.0
120.00	3.2
130.00	3.5
140.00	3.8
150.00	4.1

Fluoro Patient Entrance Exposure Rate
Acrylic	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min
Attenuator	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.
(cm)
10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00

Dig Acq Patient Entrance Exposure Rate
Acrylic	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min
Attenuator	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.
(cm)
10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00

KAP Conversion Table
	µGy·m²
cGy·cm²	1.00
Gy·cm²	100.00
mGy·cm²	0.01
µGy·m²	1.00

## DataEntry
Patient Entrance Exposure Rate (Fluoroscopy)*
Image	Aluminum
Mode	Attenuator
(cm)	(cm)	kVp	mA	uGy/pls	kVp	mA	uGy/pls	kVp	mA	uGy/pls
Normal	10.00
	20.00
	30.00
Mag 1	10.00
	20.00
	30.00
Mag 2	10.00
	20.00
	30.00
Mag 3	10.00
	20.00
	30.00
Mag 4	10.00
	20.00
	30.00
	max

Patient Entrance Exposure Rate (Digital Acquisition)*
Image	Aluminum
Mode	Attenuator
(cm)	(cm)	kVp	mA	uGy/pls	kVp	mA	uGy/pls	kVp	mA	uGy/pls
Normal	10.00
	20.00
	30.00
Mag 1	10.00
	20.00
	30.00
Mag 2	10.00
	20.00
	30.00
Mag 3	10.00
	20.00
	30.00
Mag 4	10.00
	20.00
	30.00
	max

## DataPage
Sheet	FLUORO
SurveyID
HVL	80.80	6.12
Field size

Dose mode	Auto 1

	Auto
SEE	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
Max SEE	max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
Dose mode	Auto 1

	Auto
SEE	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
Max SEE	max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
IEER		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
IEER		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
TO.N3

TO.10	Field size	A	B	C	D	E	F	G	H	J	K	L	M

Resolution