## QC Test Summary
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY

Site Name					Report Date	13-Nov-2015
Address					Survey Date	30-Dec-1899
X-Ray Unit Manufacturer					Model
Date of Installation					Room ID
Film (mfr & type)					Screen (mfr & type)
Film Processor Mfr					Model
Digital Image Receptor Mfr					Model
Medical Physicist's Name			Eugene Mah					Signature

Medical Physicist's QC Tests
					PASS/FAIL
1. Stereotactic Breast Biopsy Unit Assembly Evaluation
2. Collimation Assessment
		A. X-ray field adequately matches image receptor
		B. Biopsy window generally centered over digital image receptor (NA if film used)
3. Focal Spot Performance and System Limiting Resolution
		A. Focal spot performance acceptable (NA if digital used)
		B. Digital system spatial resolution acceptable (NA if film used)
4. kVp Accuracy and Reproducibility
		Measured average kVp within ±5% of indicated kVp
		kVp coefficient of variation ≤ 0.02
5. Beam Quality Assessment (Half-Value Layer Measurement)
		HVL is within acceptable lower and upper limits at all kVp values tested
6. AEC System or Manual Exposure Performance Assessment
		Optical density or signal range acceptable
7. Receptor Speed Uniformity
		Screen speed uniformity or digital receptor uniformity acceptable
8. Breast Entrance Exposure, Average Glandular Dose and Exposure Reproducibility
		Exposure reproducibility is within acceptable limits
		Average glandular dose to a 4.2 cm thick breast is ≤ 3 mGy (300 mrad)					mGy
9. Image Quality Evaluation
		Phantom image quality is acceptable
		Phantom type:	ACR Mammography Accreditation Phantom					Mini Phantom
		Phantom image quality scores:				Fibers		Specks		Masses
10. Artifact Evaluation
		Artifacts were not apparent or not significant:

		Artifacts identified:
11. Localization Accuracy Test
		Localization and sampling accurate/object captured

*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***

## Tech QC Eval
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY (continued)

Evaluation of Site's Technologist QC Program

					Frequency		PASS/FAIL
1.	Localization Accuracy Test					Daily
2.	Darkroom Cleanliness (NA if digital used)					Daily
3.	Processor Quality Control (NA if digital used)					Daily
4.	Phantom Images					Weekly
5.	Screen Cleanliness (NA if digital used)					Weekly
6.	Viewboxes and Viewing Conditions (NA if digital used)					Weekly
7.	Hardcopy Output Quality (if hardcopy produced from digital data)					Monthly
8.	Visual Checklist					Monthly
9.	Analysis of Fixer Retention in Film (NA if digital used)					Quarterly
10.	Compression					Semi-annually
11.	Repeat Analysis					Semi-annually
12.	Screen-Film Contact (NA if digital used)					Semi-annually
13.	Darkroom Fog (NA if digital used)					Semi-annually
14.	Zero Alignment Test (if required by manufacturer)					Before each patient
15.	Any additional tests required by manufacturer					As required by manufacturer

Medical Physicist's Recommendations for Quality Improvement

Comments:

## Sheet1
1					Revision 1.0-20140918					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Stereotactic Breast Biopsy System Compliance Inspection					Stereotactic Breast Biopsy System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector	Eugene Mah		Eugene Mah				kV Set	mAs Set	Add Filt	Target	Filter	kVp	ms	Dose (mGy)	Dose rate (mGy/s)	Corr kV
10				Facility:					Site Number:					Facility:				Site Number:					Facility:					24	50	0.0	Mo	Mo
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:					24	50	0.0	Mo	Mo
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:					24	50	0.2	Mo	Mo
13				Survey ID:					Number of X-ray Tubes:					Survey ID:				Number of X-ray Tubes:					Survey ID:					24	50	0.2	Mo	Mo
14					Room Number:				Accession Number:					Room Number:					24	50	0.3	Mo	Mo
15			X-Ray Generator					Site Number:					24	50	0.3	Mo	Mo
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:					25	50	0.0	Mo	Mo
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:					25	50	0.0	Mo	Mo
18				Max kVp:					Max mA:					Model:				Serial Number:					Number of X-ray Tubes:					25	50	0.2	Mo	Mo
19					Max kVp:				Max mA:					Accesssion Number:					25	50	0.2	Mo	Mo
20			X-Ray Tube 1					X-Ray Generator					25	50	0.3	Mo	Mo
21				Tube Designation/Use:					Target:	Mo					X-Ray Tube 1					Target:	Mo	Mo				Manufacturer:					25	50	0.3	Mo	Mo
22				Manufacture Date:					Tube Designation/Use:					Model:					26	50	0.0	Mo	Mo
23			Insert					Filter:	Mo					Manufacture Date:					Max kVp:					28	20	0.0	Mo	Mo
24				Manufacturer:					Insert					Filter:	Mo	Mo				Manufacture Date:					28	50	0.0	Mo	Mo
25				Model:					Manufacturer:					Serial Number:					28	50	0.0	Mo	Mo
26				Serial Number:				Focal Spot Sizes (mm)					Model:					Max mA:					28	50	0.0	Mo	Mo
27			Housing					Large:					Serial Number:				Focal Spot Sizes (mm)					X-Ray Tube 1					28	50	0.0	Mo	Mo
28				Manufacturer:					Small:					Housing					Large:					Tube Designation/Use:					28	50	0.3	Mo	Mo
29				Model:				Filtration	Inherent:					Manufacturer:				Small:					Manufacture Date:					28	50	0.3	Mo	Mo
30				Serial Number:					Added:					Model:					Manufacturer:					28	50	0.4	Mo	Mo
31					Serial Number:				Filtration					Model:					28	50	0.4	Mo	Mo
32					Inherent:					Serial Number:					28	100	0.0	Mo	Mo
33					Technique Chart					Added:					Manufacturer:					28	300	0.0	Mo	Mo
34					Model:					30	50	0.0	Mo	Mo
35		Compressed	Fatty Breast			50% Fatty -			Dense Breast					Serial Number:					32	50	0.0	Mo	Mo
36		Breast				50% Dense Breast					Large:					32	50	0.0	Mo	Mo
37		Thickness	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs					Inspection Results					Small:					32	50	0.3	Mo	Mo
38		<3 cm					Enter 1 for YES, 2 for NO, 3 for NA					Target 1:		Change	Mo				32	50	0.3	Mo	Mo
39		3 to 5 cm					DHEC Registration sticker is present, clearly visible and legible					Target 2:					32	50	0.4	Mo	Mo
40		5 to 7 cm					DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Filter 1:		Change	Mo				32	50	0.4	Mo	Mo
41		>7 cm					Radiation warning label posted on the generator control panel					Filter 2:					34	50	0.0	Mo	Mo
42					Operator manuals are available.					Inherent:
43					Inspection Results					Monthly radiation monitoring reports are posted.					Added
44		Rule Number					Compliance					Technique Chart
45		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Mammography Unit Assembly Evaluation					Fatty Breast
46		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD				Free-standing unit is mechanically stable					Target/Filter:
47		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD				All moving parts move smoothly, without obstructions to motion					kVp:
48				Operator manuals are available.					TBD				All locks and detents work properly					Density:
49				Monthly radiation monitoring reports are posted.					TBD				Needle holder and guides are firmly attached and support the needle					Target/Filter:
50					Mammography Unit Assembly Evaluation					Image receptor, compression plate and biopsy window are free from vibration					kVp:
51				Free-standing unit is mechanically stable					TBD				Image receptor is held securely by assembly in any orientation					Density:
52				All moving parts move smoothly, without obstructions to motion					TBD				Compressed breast thickness scale accurate to +0.5 cm, reproducible to +2 mm					Target/Filter:
53				All locks and detents work properly					TBD				Patient or operator is not exposed to sharp or rough edges, or other hazards					kVp:
54				Needle holder and guides are firmly attached and support the needle					TBD				Operator technique control charts are posted					Density:
55				Image receptor, compression plate and biopsy window are free from vibration					TBD				Operator protected during exposure by adequate radiation shielding					Target/Filter:
56				Image receptor is held securely by assembly in any orientation					TBD				All indicator lights working properly					kVp:
57				Compressed breast thickness scale accurate to 0.5 cm, reproducible to 2 mm					TBD				Localization accuracy test passes					Density:
58				Patient or operator is not exposed to sharp or rough edges, or other hazards					TBD					50/50 Breast
59				Operator technique control charts are posted					TBD					Target/Filter:
60				Operator protected during exposure by adequate radiation shielding					TBD					MQSA Equipment Standards					kVp:
61				All indicator lights working properly					TBD				Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position					Density:
62				Localization accuracy test passes					TBD				Tube-receptor assembly locking mechanism does not fail during power interruption					Target/Filter:
63					Image Receptor Sizes 18 x 24 & 24 x 30 are available					kVp:
64					Moving grids are present for all image receptor sizes					Density:
65					Magnification can be done without grid					Target/Filter:
66					Beam limiting device allows useful beam to extend beyond chest wall edge of receptor					kVp:
67		Date:					Inspector:	Eugene Mah		1	Average illumination of the light field is at least 160 lux					Density:
68		Room Number					Survey ID:				Magnification mode is available for non-interventional problem solving procedures					Target/Filter:
1					Medical University of South Carolina			System provides magnification factor between 1.4 – 2.0					kVp:
2					Inspection Results					Stereotactic Breast Biopsy System Compliance Inspection			Prior indication of selected focal spot is provided					Density:
3		Rule Number					MQSA Equipment Standards					Prior indication of selected target is provided					Dense Breast
4		900.12(b)(3i)		Tube-receptor can be fixed into any operating position and does not move					TBD				Post exposure display of target/focal spot is provided					Target/Filter:
5		900.12(b)(3ii)		Tube-receptor locking mechanism does not fail during power interruption					TBD				Initial power driven compression can be activated hands-free from both sides					kVp:
6		900.12(b)(4i)		Image Receptor Sizes 18 x 24 & 24 x 30 are available					TBD				Fine adjustment of compression can be activated from both sides					Density:
7		900.12(b)(4ii)		Moving grids are present for all image receptor sizessizes					TBD				Compression paddles match the sizes of all image receptors					Target/Filter:
8		900.12(b)(4iii)		Magnification can be done without grid					TBD				Compression paddles do not flex by more than 1 cm					kVp:
9		900.12(b)(5i)		Beam limiting device allows useful beam to extend beyond chest wall edge					TBD				Compression paddles are flat and parallel to the breast support table					Density:
10		900.12(b)(5ii)		Average illumination of the light field is at least 160 lux					YES				Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor					Target/Filter:
11		900.12(b)(6i)		Magnification mode available for non-interventional problem solving					TBD				Chest wall edge may be bent upwards for patient comfort and does not appear in the image					kVp:
12		900.12(b)(6ii)		System provides magnification factor between 1.4 – 2.0					TBD				Manual selection of mAs or mA and/or time is available					Density:
13		900.12(b)(7i)		Prior indication of selected focal spot is provided					TBD				Prior indication of selected technique factors is provided					Target/Filter:
14		900.12(b)(7i)		Prior indication of selected target is provided					TBD				Post exposure display of technique factors is provided for AEC exposures					kVp:
15		900.12(b)(7iii)		Post exposure display of target/focal spot is provided					TBD				AEC mode is operable in all combinations of equipment configuration					Density:
16		900.12(b)(8)(iB)		Initial power driven compression can be activated hands-free from both sides					TBD				AEC sensor positioning is flexible
17				Fine adjustment of compression can be activated from both sides					TBD				AEC sensor position and size is indicated on the compression paddle					MGD:
18		900.12(b)(8)(iiA)		Compression paddles match the sizes of all image receptors					TBD				Selected position of the AEC sensor is clearly indicated					mGy/mAs:
19		900.12(b)(8)(iiB)		Compression paddles do not flex by more than 1 cm					TBD				AEC density control setting is provided					mGy/s:
20		900.12(b)(8)(iiC)		Compression paddles are flat and parallel to the breast support table					TBD			3	X-ray film used is appropriate for mammography					Res Par 512:
21		900.12(b)(8)(iiD)		Chest wall edge of the compression paddle is straight and parallel to receptor					TBD			3	Intensifying screens used are appropriate for mammography and matched to the film used					Res Perp 512:
22		900.12(b)(8)(iiE)		Chest wall edge may be bent upwards and does not appear in the image					TBD			3	Film processing chemicals are appropriate for the film being used					Res Par 1024:
23		900.12(b)(9i)		Manual selection of mAs or mA and/or time is available					TBD				Lights for film illumination i.e. hot lights are available					Res Perp 1024:
24		900.12(b)(9ii)		Prior indication of selected technique factors is provided					TBD				Film masking devices are available to all interpreting physicians
25		900.12(b)(9iii)		Post exposure display of technique factors is provided for AEC exposures					TBD					Image Quality (Acq Unit)
26		900.12(b)(10i)		AEC mode is operable in all combinations of equipment configuration					TBD					kV:			0
27		900.12(b)(10ii)		AEC sensor positioning is flexible					TBD					Mammography Phototimer Technique Chart					mA:			0
28				AEC sensor position and size is indicated on the compression paddle					TBD					Exp Ind:			0
29				Selected position of the AEC sensor is clearly indicated					TBD			Compressed	Fatty Breast			50% Fatty -			Dense Breast					Fibers:
30		900.12(b)(10iii)		AEC density control setting is provided					TBD			Breast				50% Dense Breast					Specks:
31		900.12(b)(11)		X-ray film used is appropriate for mammography					NA			Thickness	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs			Masses:
32		900.12(b)(12)		Intensifying screens used are appropriate for mammography and matched to film					NA			<3 cm					Average:
33		900.12(b)(13)		Film processing chemicals are appropriate for the film being used					NA			3 to 5 cm					SD:
34		900.12(b)(14)		Lights for film illumination i.e. hot lights are available					TBD			5 to 7 cm					kV:			0
35		900.12(b)(15)		Film masking devices are available to all interpreting physicians					TBD			>7 cm					mA:			0
36					Exp Ind:			0
37					Compressed	Fatty Breast			50% Fatty -			Dense Breast					Fibers:
38					Breast				50% Dense Breast					Specks:
39					Thickness	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs			Masses:
40					<3 cm					Average:
41					3 to 5 cm					SD:
42					5 to 7 cm
43					>7 cm
44
45					Inspection Results					Comment Page 1
46					Compression Force Indicator
47					Indicated	Measured	After 20s	Slip
48					Power
49					Manual
50					Criteria:	Minimum compression force – 25 lbs
51					Maximum compression force – 45 lbs
52
53					Compression Thickness Indicator
54					BR-12 (cm)	Indicated	Difference
55					1
56					2
57					4			Mean:
58					4			StDev:
59					4
60					4				Criteria:	Difference between indicated and actual thickness is <= 2mm
61					6
62					8
63
64					Paddle flex is less than 1 cm				Compression thickness indicator:
65					Overall unit assembly
66
67		Date:					Inspector:	Eugene Mah		Artifact Evaluation
68		Room Number					Survey ID:					kV:	28
1					Medical University of South Carolina			Target:	Mo	mAs:
2					Inspection Results					Stereotactic Breast Biopsy System Compliance Inspection			Filter:	Mo	EI:
3		Compression Force Indicator					Focal spot:	Large	Acceptable:
4				Indicated	Measured	After 20s	Slip
5			Power					Comments:
6			Manual					Acceptable:					Do not exceed 390 chars
7			Criteria:	Minimum compression force – 25 lbs					New:
8				Maximum compression force – 45 lbs
9					X-ray Field/Image Receptor Alignment
10		Compression Thickness Indicator					SID (cm)
11			BR-12 (cm)	Indicated	Difference				Paddle flex is less than 1 cm:	TBD					Enter deviation in cm
12			1					Overall unit assembly:	TBD					% Deviation
13			2					Compression thickness indicator:	TBD					Left:
14			4			Mean:					Right:
15			4			StDev:					Anterior:
16			4			Criteria:	Difference between indicated and actual thickness is <= 2mm					Chest Wall:
17			4
18			6					Criteria:	X-ray field does not extend beyond image receptor by more than 5 mm on any side
19			8					Biopsy window is centered in image
20					Biopsy window centered in image
21		Artifact Evaluation
22					System Resolution			512		1024
23			Target:	Mo					Par	Perp	Par	Perp
24			Filter:	Mo					kV	lp/mm:
25			Focal spot:	Large					mAs	Previous:
26			kV:	28					Acceptable:
27			mAs:					Criteria:	Note any significant degradation from previous measurement
28			EI:
29			Acceptable:					AEC Thickness Tracking
30			Criteria:	No visible artifacts in the images					AEC Mode:			AEC Position:
31			Comments:					Thickness					Reported
32					(cm)	Filter	kV	mAs	ROI	MGD (mGy)	Deviation
33					2
34					4
35		X-ray Field/Image Receptor Alignment					4	0	0
36					4	0	0
37			% Deviation					4	0	0
38		Left:				Biopsy window is centered in image					6
39		Right:			Criteria:	X-ray field does not extend beyond image receptor by more than 5 mm on any side					8
40		Anterior:				Biopsy window is centered in image					Average:
41		Chest Wall:					Criteria:	Mean pixel value of each image is within 10% of the average pixel value
42
43		System Resolution		512		1024
44				Perp	Par	Perp	Par					AEC Density Control
45			lp/mm:					Acceptable:					AEC Mode:	0		AEC Position:	0	Reported
46			Criteria:	Note any significant degradation from previous measurement					Density	kV	mAs	ROI	MGD (mGy)	Change			Low	High
47		Detector Uniformity					-3					0.5	0.61
48		kV:		mA:					-2					0.63	0.77
49			ROI	SD	SNR	Deviation					-1					0.77	0.94
50		Center					0
51		ROI 1					1					1.04	1.27
52		ROI 2					2					1.17	1.43
53		ROI 3					3					1.31	1.6
54		ROI 4					4					1.44	1.76
55		Average					Criteria:	Change in pixel value at each density step is within specified ranges
56		SD
57					Detector Uniformity
58		Criteria:	SNR values are within 15% of the SNR in the center					kV:
59					mAs:
60					ROI	SD	SNR	Deviation
61					Center
62					ROI 1
63					ROI 2
64					ROI 3
65					ROI 4
66					Average
67		Date:					Inspector:	Eugene Mah			SD
68		Room Number					Survey ID:
1					Medical University of South Carolina			Criteria:	SNR values are within 15% of the SNR in the center
2					Inspection Results					Stereotactic Breast Biopsy System Compliance Inspection
3		AEC Thickness Tracking
4		AEC Mode:			AEC Position:
5			Thickness					Reported					Mean Glandular Dose
6			(cm)	Filter	kV	mAs	EI	MGD (mGy)	Deviation					Phantom:			kV:
7			2					Target/Filter:			mA:
8			4					Reported		Air KERMA
9			4	0	0					mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:
10			4	0	0					ESE (mGy):
11			4	0	0					DGN (mrad/R):
12			6					MGD (mGy):
13			8					Prev MGD (mGy):
14					Average:					Average:					% Change:
15					Coeff of Var:
16			Criteria:	Mean pixel value of each image is within 10% of the average pixel value
17					Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:
18		AEC Density Control					Deviation from reported MGD is less than15%					mAs at 3 mGy:
19		AEC Mode:	0		AEC Position:	0		kV:
20					Reported					kVp Accuracy and Output
21			Density	mAs	EI	MGD (mGy)	Change	Acceptable					Meter:			Calibration Date:
22			-3					Serial #:			Calibration Due:
23			-2
24			-1					Measured Values
25			0					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
26			1					Mo	Mo	24	50
27			2					Mo	Mo	25	50
28			3					Mo	Mo	26	50
29			4					Mo	Mo	28	50
30			Criteria:	Change in pixel value at each density step is within specified ranges					Mo	Mo	30	50
31					Mo	Mo	32	50
32		Mean Glandular Dose					Mo	Mo	34	50
33			Phantom:		kV:
34			Target/Filter:		mA:					Criteria:	kV accuracy is <= 5%
35					Reported		Air KERMA					kV, Exposure, and Output Reproducibility
36				mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:
37					ESE (mGy):					Measured Values
38					DGN (mrad/R):					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
39					MGD (mGy):					Mo	Mo	28	50
40					Prev MGD (mGy):					Mo	Mo	28	50
41			Average:					% Change:					Mo	Mo	28	50
42			Coeff of Var:					Mo	Mo	28	50
43			Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:					Mean:
44				Deviation from reported MGD is less than15%					mAs at 3 mGy:					Std Dev:
45					Coeff of Var:
46		Signal to Noise Ratio					Previous:
47			kV:				Current	Previous	Change	Acceptable					Criteria:	Coefficient of variation <= 5%
48			mA:			SNR:					Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
49			Target:			CNR:
50			Filter:					Exposure Linearity
51			Criteria:	SNR must be >= 40					Measured Values
52				CNR does not change by more than 15%					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
53					Mo	Mo	28	20
54		Image Quality Evaluation					Mo	Mo	28	50
55					Mo	Mo	28	100
56				Acq Workstation					Mo	Mo	28	300
57				Normal					Criteria:	Coefficient of linearity is <= 10%					Coeff of Linearity:
58			kV:	0
59			mA:	0					Unit Can Maintain Output for 3 sec at 28 kVp
60			Exposure Index:	0
61			Fibers:					Half Value Layer
62			Specks:					Target/Filter	Mo/Mo	Mo/Mo	Mo/Mo	Mo/Mo
63			Masses:					kV	24	25	28	32
64					Al (mm)	Air Kerma (mGy)
65					0.0
66					0.2
67		Date:					Inspector:	Eugene Mah			0.3
68		Room Number					Survey ID:				0.4
1					Medical University of South Carolina			0.5
2					Inspection Results					Stereotactic Breast Biopsy System Compliance Inspection			0.0
3		Meter:		SN:			Calibration Date:					HVL
4					Calibration Due:					0 mm
5		kVp Accuracy					kVp and Output Reproducibility					Min	0.24	0.25	0.28	0.32
6		Target:	Mo	Filter:	Mo		Target:	Mo	Filter:	Mo					Max	0.36	0.37	0.40	0.44
7		mAs:	50				mAs:	50
8		Indicated	Measured				Indicated	Measured					Criteria:	HVL is between the minimum and maximum limits
9		kVp	kVp	Error (%)			kVp	kVp	mGy	mGy/mAs	mGy/s
10		24					28					Image Quality Evaluation
11		25					Acquisition Unit		1	Mini phantom used?			Previous
12		26					512	1024			512	1024
13		28					kV:	0	0
14		30					Mean:					mAs:	0	0
15		32					Std Dev:					Exposure Index:	0	0
16		34					Coeff of Var:					Fibers:
17					Acceptable:					Specks:
18			Acceptable:				Criteria:	Coefficient of variation <= 2%					Masses:
19					Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target					Average PV:
20			Criteria:	kVp accurate to within 5% of indicated				Minimum acceptable output is 2 mGy/s at 28 kVp for W target					SD:
21					Criteria:	Must see at least 3 fibers, 3 specks, 2.5 masses
22		Exposure Linearity					Previous Year Comments
23		Target:	Mo	Filter:	Mo					Comments:
24					Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
25		Set mAs	kVp	mGy	mGy/mAs	mGy/s					Additional Comments:
26		20					Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
27		50					Additional Comments:
28		100					Do not exceed 390 chars		0			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
29		300					Additional Comments:
30				Coeff of Linearity:					Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
31			Criteria:	Coefficient of linearity is <= 10%					Additional Comments:
32					Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
33		Half Value Layer					Additional Comments:
34		Target/Filter	Mo/Mo	Mo/Mo	Mo/Mo	Mo/Mo					Do not exceed 390 chars		0
35		kV	24	25	28	32					Additional Comments:
36		HVL					Do not exceed 390 chars		0
37		0 mm
38		Min	0.24	0.25	0.28	0.32
39		Max	0.36	0.37	0.40	0.44
40		Acceptable:
41			Criteria:	HVL is between the minimum and maximum limits
42
43					Comments
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast					Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast					Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
	kVp					kVp					kVp
HVL	23	24	25	26	27	28	29	30	31	HVL	23	24	25	26	27	28	29	30	31	HVL	23	24	25	26	27	28	29	30	31
0.23	116					0.23					0.23
0.24	121	124					0.24					0.24
0.25	126	129	131					0.25					0.25
0.26	130	133	135	138					0.26					0.26
0.27	135	138	140	142	143					0.27					0.27
0.28	140	142	144	146	147	149				0.28			149	151	154					0.28			150	155	159
0.29	144	146	148	150	151	153	154			0.29			154	156	158	159				0.29			155	160	164	168
0.3	149	151	153	155	156	157	158	159		0.3			158	160	162	162	163			0.3			160	164	168	172	176
0.31	154	156	157	159	160	161	162	163	164	0.31			163	164	166	166	167	167		0.31			165	168	172	174	180	182
0.32	158	160	162	163	164	166	167	168	168	0.32			167	169	171	171	171	172	172	0.32			169	173	177	181	184	186	188
0.33	163	165	166	168	169	170	171	173	173	0.33			171	173	175	176	176	176	176	0.33			174	178	181	185	188	190	192
0.34	168	170	171	172	173	174	175	176	177	0.34			176	178	179	179	180	180	180	0.34			179	183	186	190	193	195	196
0.35		174	175	176	177	178	179	180	181	0.35			180	181	183	183	184	185	185	0.35			184	187	190	194	197	199	201
0.36			179	181	182	183	184	185	185	0.36			185	186	187	187	188	188	189	0.36			189	192	195	198	201	204	205
0.37				185	186	187	188	189	190	0.37			189	190	191	191	192	193	193	0.37			193	196	199	202	205	207	209
0.38					190	191	192	193	194	0.38			193	194	196	196	197	197	197	0.38			198	201	204	207	209	211	213
0.39					196	197	198	198	0.39			198	199	200	200	201	201	202	0.39			203	206	208	211	214	216	217
0.4					201	202	203	0.4			202	203	204	204	205	205	206	0.4			208	211	213	216	218	220	221
0.41					206	207	0.41			206	207	208	208	209	209	210	0.41			213	215	217	220	222	224	225
0.42					211	0.42			211	211	212	212	213	213	214	0.42			218	220	222	224	226	228	229

Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast					Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
	kVp					kVp
HVL	22	23	24	25	26	27	28	29	30	31	32	33	HVL	27	28	29	30	31	32	33	34	35	36	37	38	39
0.3	152	157	163	166	170	173	175	177	179	182	184	187	0.3	222	226	229	231	234	236	239	241	244	246	248	250	252
0.325	163	169	174	177	181	183	186	188	190	192	195	197	0.325	233	236	239	242	244	246	248	251	253	256	258	260	262
0.35	175	180	185	188	191	194	196	198	200	202	205	207	0.35	244	247	249	252	254	256	258	260	263	265	267	269	271
0.375	186	191	196	199	202	205	207	209	211	213	215	218	0.375	254	257	260	262	264	266	268	270	273	275	277	279	281
0.4	198	203	207	210	213	215	217	219	221	223	226	228	0.4	265	267	270	272	274	276	278	280	282	284	286	288	290
0.425	209	214	218	221	224	226	228	230	232	234	236	238	0.425	275	278	280	282	284	286	288	290	292	294	296	298	300
0.45	221	226	230	232	235	237	238	240	242	244	246	248	0.45	286	288	290	292	294	296	298	299	301	303	305	307	309
0.475	233	237	241	243	245	247	249	251	253	254	256	258	0.475	296	298	3090	302	304	305	307	309	311	313	315	317	318
0.5	244	248	252	254	256	258	260	261	263	265	267	269	0.5	306	308	310	312	313	315	317	319	320	322	324	326	328
0.525	256	260	263	265	267	269	270	272	273	275	277	279	0.525	316	318	320	322	323	325	326	328	330	332	333	335	337
0.55	267	271	274	276	278	279	281	282	284	285	287	289	0.55	326	328	330	331	333	334	336	338	339	341	343	344	346
0.575	279	282	285	287	288	290	291	292	294	296	297	299	0.575	336	338	339	341	342	344	345	347	349	350	352	354	355
0.6	290	293	296	297	299	300	301	303	304	306	308	310	0.6	346	348	349	350	352	353	355	356	358	359	361	363	364
0.625	301	304	306	308	309	310	312	313	315	316	318	320	0.625	356	357	358	360	361	362	364	365	367	368	370	372	373
0.65	312	314	317	318	320	321	322	323	325	326	328	330	0.65	365	367	368	369	370	372	373	375	376	378	379	381	382
0.675	322	325	327	328	330	331	333	333	335	336	338	340	0.675	374	376	377	378	379	381	382	383	385	386	388	390	391
0.7	333	335	337	339	340	341	342	343	345	346	348	350	0.7	384	385	386	387	388	390	391	392	394	395	397	398	400
0.725	342	345	347	348	349	351	352	353	354	356	358	360	0.725	393	394	395	396	397	399	400	401	403	404	406	407	408
0.75	352	355	357	358	359	360	361	363	364	366	368	369	0.75	402	403	404	405	406	407	409	410	411	413	414	416	417
0.775	361	365	367	368	369	370	371	372	374	375	377	379	0.775	410	411	412	413	415	416	417	418	420	421	423	424	425
0.8	369	374	376	377	378	379	380	382	383	385	387	389	0.8	419	420	421	422	423	424	425	427	428	429	431	432	434

Siemens Inspiration
	PMMA thickness (mm)	Equiv Breast thickness (mm)	HVL (mm Al)
			0.25	0.3	0.35	0.4	0.45	0.5	0.55	0.6	Slope	Intercept
g-factor	45	53	0.13	0.155	0.177	0.198	0.22	0.245	0.272	0.295	0.4681	0.0126
c-factor	45	53		1.109	1.105	1.102	1.099	1.096	1.091	1.088	-0.0693	1.1298
g-factor	50	60	0.112	0.135	0.154	0.172	0.192	0.214	0.236	0.261	0.4162	0.0076
c-factor	50	60		1.164	1.160	1.151	1.150	1.144	1.139	1.134	-0.0993	1.1935

	Mo/Mo	Mo/Rh	W/Rh
s-factor	1	1.017	1.042

kV Accuracy/Output		Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV^2					Radcal kV correction for Hologic Selenia/Dimension
Target	kV	kV^2	mGy/mAs	mGy/s					Selenia		Dimension
Mo	24	576					Measured	W/Rh	W/Ag	W/Rh	W/Ag	W/Al
Filter	25	625
Mo	26	676
	28	784
	30	900
	32	1024
	34	1156
		Slope:	#VALUE!	#VALUE!
		Intercept:	#VALUE!	#VALUE!

Half Value Layer
Target	kV	HVL					Target	Filter	C
Mo	24					Mo	Mo	0.12
Filter	25					Mo	Rh	0.19
Mo	28					Rh	Rh	0.22
	32					Critiera:	Min HVL > kV/100 (no paddle)
	Slope:	#VALUE!					Min HVL > kV/100 + 0.03 (with paddle)
	Intercept:	#VALUE!					Max HVL < kV/100 + C

DGN	#VALUE!
Mo/Mo	Mo/Rh	Rh/Rh	W/Rh	W/Ag	Siemens Inspiration
-95.4	-152.0	-192.1	#N/A	#N/A	#VALUE!

Radcal kV Correction for Hologic Selenia
	kV		kV^0	kV^1	kV^2	kV^3	kV^4	kV^5	kV^6
W/Rh	<	27.586	-8375.0727645926	975.9254356043	-37.913729682	0.4908658347	0	0	0
	>=	27.586	-9984.6167916494	1436.5245457141	-82.5051021853	2.3655908176	-0.0338672434	0.0001936869	0
		Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
		Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
	kV		kV^0	kV^1	kV^2	kV^3	kV^4	kV^5	kV^6
W/Ag	<	30.1	-540847.695500773	100186.233642731	-7418.4790179813	274.4766092958	-5.0743695436	0.0375005748	0
	>=	30.1	-11057.7739361992	1297.2285673767	-56.9891889897	1.1115828564	-0.0081233997	0	0
		Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
		Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.

Radcal kV Correction for Hologic Dimension
	kV		kV^0	kV^1	kV^2	kV^3	kV^4
W/Rh	<	26.9	138.88667	-10.72639	0.26216	-0.00082
	>=	26.9	-5009.7751652	605.732006	-27.3018617	0.5467114	-0.0040986

W/Ag	kV		kV^0	kV^1	kV^2	kV^3	kV^4
	<	28.7	296.34185	-31.62925	1.18025	-0.01417
	>=	28.7	4.834469	0.919242

W/Al			kV^0	kV^1	kV^2	kV^3	kV^4
	<	28.7	49.31115	-2.93017	0.07379
		28.7 < kV < 30.1	-24.875	1.8032
	>=	30.1	-4.83461	1.15715

## Corrected kV
GENERAL
	Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
	Read the true corrected kV in the adjacent cell to the right.
	If that cell is light green, the corrected kV is accurate.
	If the cell is red, the entered kV is out of range and the correction should not be used.
	If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.

Tungsten-Rhodium (W / Rh)
	Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
	Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
	Radcal kV	True kV
	25.03	21.81

Tungsten-Silver (W / Ag)
	Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
	Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
	Radcal kV	True kV
	37.1	38.97

27.586	-8375.0727645926	975.9254356043	-37.913729682	0.4908658347	0	0	0
	-9984.6167916494	1436.5245457141	-82.5051021853	2.3655908176	-0.0338672434	0.0001936869	0

30.1	-540847.695500773	100186.233642731	-7418.4790179813	274.4766092958	-5.0743695436	0.0375005748	0
	-11057.7739361992	1297.2285673767	-56.9891889897	1.1115828564	-0.0081233997	0	0

## dropdowns
Pass/Fail
Pass
Fail

Pass/Fail/NA
Pass
Fail
NA

Phantom Scores
6.0
5.5
5.0
4.5
4.0
3.5
3.0
2.5
2.0
1.5
1.0
0.5