## Sheet1
1 | [11 empty] | Revision 1.4-20190606 | [11 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [6 empty] | Page1
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [23 empty] | All: | Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5 | [5 empty] | Dental CT System Compliance Inspection | [10 empty] | Dental CT System Compliance Inspection
6 | [23 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah |  | Print Area |  | Change | Page1 |  | Half Value Layer
8 | [5 empty] | System Information | [5 empty] | Previous Date: | [11 empty] | Date |  |  |  |  | DHEC |  |  | MUSC
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  | Inspector |  | Change | Eugene Mah |  |  | Min HVL |  |  | HVL | Min HVL | Max HVL
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  |  | kVp | (mm Al) |  | kVp | (mm Al) | (mm Al) | (mm Al)
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  |  | 30 | 0.3 |  | 60 | 2.2 | 1.8 | 2.4
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  |  | 40 | 0.4 |  | 70 | 2.6 | 2.2 | 2.9
13 |  |  | Survey ID: |  |  |  |  | Accession Number: | [5 empty] | Survey ID: |  |  |  | Accession Number: | [5 empty] | Survey ID: |  |  |  |  | 50 | 0.5 |  | 80 | 3 | 2.6 | 3.5
14 | [13 empty] | Room Number: | [9 empty] | Room Number: |  |  |  |  | 51 | 1.2 |  | 90 | 3.5 | 3.1 | 4
15 |  | CT Scanner | [21 empty] | Site Number: |  |  |  |  | 60 | 1.3 |  | 100 | 3.9 | 3.5 | 4.5
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | CT Scanner | [10 empty] | Authorized Use: |  |  |  |  | 70 | 1.5 |  | 110 | 4.3 | 3.9 | 5.1
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: | [5 empty] | Date of Installation: |  |  |  |  | 71 | 2.1 |  | 120 | 5.7 | 4.4 | 5.6
18 | [13 empty] | Model: |  |  |  | Serial Number: | [5 empty] | Accesssion Number: |  |  |  |  | 80 | 2.3 |  | 130 | 5.2 | 4.8 | 6.1
19 |  | X-Ray Tube |  |  |  |  | Focal Spot Sizes (mm) | [21 empty] | 90 | 2.5
20 |  |  | Tube Designation/Use: |  |  |  |  | Large: |  |  |  |  | X-Ray Tube |  |  |  |  | Focal Spot Sizes (mm) | [5 empty] | X-Ray Generator |  |  |  |  | 100 | 2.7
21 |  |  | Manufacture Date: |  |  |  |  | Small: | [5 empty] | Tube Designation/Use: |  |  |  | Large: | [5 empty] | Manufacturer: |  |  |  |  | 110 | 3
22 |  | Insert | [5 empty] | Micro: | [5 empty] | Manufacture Date: |  |  |  | Small: | [5 empty] | Model: |  |  |  |  | 120 | 3.2
23 |  |  | Manufacturer: | [9 empty] | Insert |  |  |  |  | Micro: | [5 empty] | Manufacture Date: |  |  |  |  | 130 | 3.5
24 |  |  | Model: |  |  |  |  | Filtration | [5 empty] | Manufacturer: | [9 empty] | Serial Number: |  |  |  |  | 140 | 3.8
25 |  |  | Serial Number: |  |  |  |  | Inherent: | [5 empty] | Model: |  |  |  | Filtration | [10 empty] | 150 | 4.1
26 |  | Housing | [5 empty] | Added: | [5 empty] | Serial Number: |  |  |  | Inherent: | [5 empty] | X-Ray Tube 1
27 |  |  | Manufacturer: | [9 empty] | Housing |  |  |  |  | Added: | [5 empty] | Tube Designation/Use:
28 |  |  | Model: | [10 empty] | Manufacturer: | [9 empty] | Manufacture Date: |  |  |  | Set kV | Set mA | kV | Time (ms) | mGy | mGy/s | HVL | Tot Filtr | Pulses
29 |  |  | Serial Number: | [10 empty] | Model: | [9 empty] | Manufacturer: |  |  |  | 60 | 5
30 | [13 empty] | Serial Number: | [9 empty] | Model: |  |  |  | 70 | 5
31 | [23 empty] | Serial Number: |  |  |  | 80 | 5
32 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance | [13 empty] | Manufacturer: |  |  |  | 80 | 5
33 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Inspection Results | [6 empty] | Model: |  |  |  | 80 | 5
34 | DHEC RHB 10.2.1 |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA | [11 empty] | Serial Number: |  |  |  | 80 | 5
35 |  |  | Pregnancy Warning sign is posted | [6 empty] | TBD | [13 empty] | Large: |  |  |  | 80 | 5
36 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD | [6 empty] | Labels, Notices, Postings | [6 empty] | Small: |  |  |  | 80 | 1
37 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | DHEC Registration sticker is present, clearly visible and legible | [10 empty] | Micro: |  |  |  | 80 | 10
38 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [10 empty] | Inherent: |  |  |  | 80 | 16
39 | [12 empty] | Pregnancy Warning sign is posted | [10 empty] | Added: |  |  |  | 90 | 5
40 |  |  |  | Radiation Safety | [8 empty] | Radiation warning label posted on the generator control panel
41 |  |  | Radiation warning light at room entrance is functional with selected tube | [6 empty] | TBD |  |  | Operator manuals are available. | [10 empty] | CTDI Measurements
42 |  |  | Lead aprons available. | [6 empty] | TBD |  |  | Monthly radiation monitoring reports are posted. | [10 empty] | kV:
43 |  |  | A properly designed and installed apron rack is present. | [6 empty] | TBD | [13 empty] | mA:
44 |  |  | Documentation of annual protective apparel integrity inspection is available. | [6 empty] | TBD | [6 empty] | Radiation Safety | [6 empty] | Time (s):
45 |  |  | Patient restraint devices available. | [6 empty] | TBD |  |  | Radiation warning light at room entrance is functional with selected tube | [10 empty] | Beam width (mm):
46 | [12 empty] | Lead aprons available. | [10 empty] | Chamber length (mm): |  | Change | 100
47 |  |  | Reconstructed images are free of artifacts | [6 empty] | TBD |  |  | A properly designed and installed apron rack is present. | [10 empty] | Chamber CF: |  | Change | 1
48 | DHEC RHB 4.11.1.1 |  | Tomographic plane is indicated by a light source | [6 empty] | TBD |  |  | Documentation of annual protective apparel integrity inspection is available. | [10 empty] | CTDIw:
49 | DHEC RHB 4.11.1.3.1 |  | Visual indication of x-ray production is present and functional | [6 empty] | TBD |  |  | Patient restraint devices available.
50 | DHEC RHB 4.11.1.6.1 |  | Positioning with the alignment light is accurate to within 2mm | [6 empty] | TBD | [13 empty] | HVL Measured kV:
51 | DHEC RHB 4.11.2.5.1 |  | Technologist has a clear view of the patient during the scan | [6 empty] | TBD |  |  | Reconstructed images are free of artifacts | [10 empty] | HVL Measured: |  | #DIV/0! | #DIV/0!
52 |  |  | Patient communication system is operational | [6 empty] | TBD |  |  | Tomographic plane is indicated by a light source | [10 empty] | mGy/mAs:
53 | [12 empty] | Visual indication of x-ray production is present and functional | [10 empty] | High contrast resolution:
54 | [12 empty] | Positioning with the alignment light is accurate to within 2mm
55 | [12 empty] | Technologist has a clear view of the patient during the scan | [10 empty] | Comment Page 1
56 | [12 empty] | Patient communication system is operational
57
58
59 | [16 empty] | Measurement Protocols
60 | [11 empty] | CTDI Measurements
61 | [12 empty] | kVp: |  |  |  | Center (mGy) | Periph (mGy)
62 | [12 empty] | mA: |  |  | 1 |  |  |  | Indicated CTDI (mGy): |  |  | Ind CTDI
63 | [12 empty] | Time (s): |  |  | 2 |  |  |  | CTDIw (mGy): |  |  | 14.9
64 | [12 empty] | Beam width (mm): |  |  | 3 |  |  |  | Previous CTDIw (mGy) |  |  | 14.9
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Chamber length (mm): | 100 |  | Average: |  |  |  | CTDI change:
66 | Room Number | [8 empty] | Survey ID: |  |  | Chamber CF: | 1 |  | CTDI (mGy): |  |  |  | Acceptable: | NA
1 | [10 empty] | Medical University of South Carolina |  | Criteria: | Change in CTDI from previous year is less than 15%
2 | [5 empty] | Inspection Results |  |  |  |  | Dental CT System Compliance Inspection
3 | CTDI Measurements | [10 empty] | Half Value Layer
4 |  | kVp: | [9 empty] | 2 | Use DHEC standards? (Type 1 to base acceptance on DHEC criteria)
5 |  | mA: |  |  | CTDIw (mGy):
6 |  | Time (s): |  |  | CTDI change: | [6 empty] | kV | mA | mm Al | mGy |  | Data Processing
7 |  | Beam width (mm): |  |  | Acceptable: | NA | [7 empty] | 0 |  |  | mm Al | Avg Exp | DATA(1/2) | Data Al | HVL
8 |  | Chamber length (mm): | 100 | [10 empty] | 0 |  |  | 0 |  |  |  | #DIV/0!
9 |  | Chamber CF: | 1 |  | Criteria: | Change in CTDI from previous year is less than 15%
10
11 | Half Value Layer | [18 empty] | Data Exp | LN(EXP)
12 | Set kVp:
13
14 | Al (mm) | mGy | mGy |  | Measured HVL: | #DIV/0! | [10 empty] | #DIV/0! | #DIV/0!
15 | 0 |  |  |  | HVL Acceptable: | #DIV/0! | [10 empty] | Acceptable (mm Al) |  |  | HVL Acceptable: | #DIV/0!
16 | [16 empty] | Low | #VALUE!
17 | [5 empty] | ACCEPTABLE RESULTS (mm Al) | [10 empty] | Desired | #VALUE! |  | Previous HVL @ kV:
18 | [5 empty] | Min acceptable HVL at kVp: | #VALUE! | [9 empty] | High | #VALUE!
19 | [5 empty] | Desired HVL at kVp: | #VALUE!
20 |  |  |  |  | Criteria: | MUSC Standard | [5 empty] | Exposure Linearity
21 | [11 empty] | kV: | 80
22
23 | [12 empty] | Meas kV | mA | Time (s) | mAs | mGy | mGy/mAs |  | mGy/mAs:
24 | [13 empty] | 5 | [5 empty] | Coeff of Linearity:
25 | [11 empty] | ` |  | 1 | [5 empty] | Previous mGy/mAs:
26 | [13 empty] | 10 | [5 empty] | Acceptable:
27 | [13 empty] | 16
28 | [12 empty] | Critiera: | Coefficient of linearity <= 10%.
29 | [13 empty] | kVp accurate to within 5% of indicated. | [9 empty] | Field sizes
30 | [13 empty] | kVp coefficient of variation (std dev./mean) <= 5%. | [9 empty] | 3D Teeth | 18 | 20
31
32 | [11 empty] | kV Accuracy, Reproducibility and Output
33 | [12 empty] | kV | mA | Measuerd values |  |  | kV % err
34 | [14 empty] | kV | mGy | mGy/s
35 | [12 empty] | 60 | 5
36 | [12 empty] | 70 | 5
37 | [12 empty] | 80 | 5
38 | [12 empty] | 80 | 5
39 | [12 empty] | 80 | 5
40 | [12 empty] | 80 | 5
41 | [12 empty] | 80 | 5
42 | [12 empty] | 90 | 5
43
44 | [13 empty] | Mean: |  |  | Acceptable:
45 | [13 empty] | St Dev:
46 | [13 empty] | Coeff of Var:
47 | kV Accuracy, Reproducibility and Output, Exposure Linearity
48 | kVp: | 80 |  |  | kVp | mA | Measuerd values |  |  | kV % err
49 | Meas kV | mA | mGy | mGy/mAs |  |  | kV | mGy | mGy/s |  |  | High Contrast Resolution
50 |  | 5 |  |  | 60 | 5 | [6 empty] | lp/cm | Acceptable | Previous
51 |  | 1 |  |  | 70 | 5 | [9 empty] | Critiera: | 7 lp/cm
52 |  | 10 |  |  | 80 | 5 |  |  | #DIV/0! | #DIV/0!
53 |  | 16 |  |  | 90 | 5
54 |  |  | mGy/mAs: |  |  | Mean: |  |  | Acceptable:
55 |  |  | Coeff of Linearity: |  |  | St Dev: | [10 empty] | Comments Page 1
56 |  |  | Acceptable: |  |  | Coeff of Var: | [9 empty] | Previous Year Comments
57 | [12 empty] | Comments:
58 | [5 empty] | Critiera: | Coefficient of linearity <= 10%. | [5 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
59 | [6 empty] | kVp accurate to within 5% of indicated. | [5 empty] | Additional Comments:
60 | High Contrast Resolution | [5 empty] | kVp coefficient of variation (std dev./mean) <= 5%. | [5 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
61 | lp/cm: | [11 empty] | Additional Comments:
62 | Acceptable: | [11 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
63 | Critiera: | 7 lp/cm | [10 empty] | Additional Comments:
64 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Additional Comments:
66 | Room Number | [8 empty] | Survey ID: |  |  | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
1 | [10 empty] | Medical University of South Carolina |  | Additional Comments:
2 | [5 empty] | Comments Page 1 |  |  |  |  | Dental CT System Compliance Inspection |  | Do not exceed 390 chars |  | 0
3 | [12 empty] | Additional Comments:
4 | [12 empty] | Do not exceed 390 chars |  | 0
5 | [12 empty] | Additional Comments:
6 | [12 empty] | Do not exceed 390 chars |  | 0
7 | [12 empty] | Additional Comments:
8 | [12 empty] | Do not exceed 390 chars |  | 0
9 | [12 empty] | Additional Comments:
10 | [12 empty] | Do not exceed 390 chars |  | 0
11 | [12 empty] | Additional Comments:
12 | [12 empty] | Do not exceed 390 chars |  | 0
13 | [12 empty] | Additional Comments:
14 | [12 empty] | Do not exceed 390 chars |  | 0
15 | [12 empty] | Additional Comments:
16 | [12 empty] | Do not exceed 390 chars |  | 0
17 | [12 empty] | Additional Comments:
18 | [12 empty] | Do not exceed 390 chars |  | 0
19 | [12 empty] | Additional Comments:
20 | [12 empty] | Do not exceed 390 chars |  | 0
21 | [12 empty] | Additional Comments:
22 | [12 empty] | Do not exceed 390 chars |  | 0
23 | [12 empty] | Additional Comments:
24 | [12 empty] | Do not exceed 390 chars |  | 0
25 | [12 empty] | Additional Comments:
26 | [12 empty] | Do not exceed 390 chars |  | 0
27 | [12 empty] | Additional Comments:
28 | [12 empty] | Do not exceed 390 chars |  | 0
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID: