## Sheet1
1					Revision 1.1-201160706					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Dental CT System Compliance Inspection					Dental CT System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1			Half Value Layer
8					System Information					Previous Date:					Date					DHEC			MUSC
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah				Min HVL			HVL	Min HVL	Max HVL
10				Facility:					Site Number:					Facility:				Site Number:					Facility:					kVp	(mm Al)		kVp	(mm Al)	(mm Al)	(mm Al)
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:					30.00	0.3		60	2.2	1.8	2.4
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:					40.00	0.4		70	2.6	2.2	2.9
13				Survey ID:					Accession Number:					Survey ID:				Accession Number:					Survey ID:					50.00	0.5		80	3.0	2.6	3.5
14					Room Number:					Room Number:					51.00	1.2		90	3.5	3.1	4.0
15			CT Scanner					Site Number:					60.00	1.3		100	3.9	3.5	4.5
16				Manufacturer:					Manufacture Date:					CT Scanner					Authorized Use:					70.00	1.5		110	4.3	3.9	5.1
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:					71.00	2.1		120	5.7	4.4	5.6
18					Model:				Serial Number:					Accesssion Number:					80.00	2.3		130	5.2	4.8	6.1
19			X-Ray Tube					Focal Spot Sizes (mm)					90.00	2.5
20				Tube Designation/Use:					Large:					X-Ray Tube					Focal Spot Sizes (mm)					X-Ray Generator					100.00	2.7
21				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Manufacturer:					110.00	3.0
22			Insert					Micro:					Manufacture Date:				Small:					Model:					120.00	3.2
23				Manufacturer:					Insert					Micro:					Manufacture Date:					130.00	3.5
24				Model:					Filtration					Manufacturer:					Serial Number:					140.00	3.8
25				Serial Number:					Inherent:					Model:				Filtration					150.00	4.1
26			Housing					Added:					Serial Number:				Inherent:					X-Ray Tube 1
27				Manufacturer:					Housing					Added:					Tube Designation/Use:
28				Model:					Manufacturer:					Manufacture Date:					kV	mGy	mGy/s	Tot Filt	HVL	mm Al
29				Serial Number:					Model:					Manufacturer:					0.00
30					Serial Number:					Model:
31					Serial Number:
32		Rule Number			Labels, Notices, Postings					Compliance					Manufacturer:
33		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Inspection Results					Model:					0.00
34		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
35				Pregnancy Warning sign is posted					TBD					Large:
36		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Labels, Notices, Postings					Small:
37				Operator manuals are available.					TBD				DHEC Registration sticker is present, clearly visible and legible					Micro:
38				Monthly radiation monitoring reports are posted.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Inherent:
39					Pregnancy Warning sign is posted					Added:
40					Radiation Safety					Radiation warning label posted on the generator control panel
41				Radiation warning light at room entrance is functional with selected tube					TBD				Operator manuals are available.					CTDI Measurements
42				Lead aprons available.					TBD				Monthly radiation monitoring reports are posted.					kV:
43				A properly designed and installed apron rack is present.					TBD					mA:
44				Documentation of annual protective apparel integrity inspection is available.					TBD					Radiation Safety					Time (s):
45				Patient restraint devices available.					TBD				Radiation warning light at room entrance is functional with selected tube					Beam width (mm):
46					Lead aprons available.					Chamber length (mm):		Change	100.00
47				Reconstructed images are free of artifacts					TBD				A properly designed and installed apron rack is present.					Chamber CF:		Change	1.00
48		DHEC RHB 4.11.1.1		Tomographic plane is indicated by a light source					TBD				Documentation of annual protective apparel integrity inspection is available.					CTDIw:
49		DHEC RHB 4.11.1.3.1		Visual indication of x-ray production is present and functional					TBD				Patient restraint devices available.
50		DHEC RHB 4.11.1.6.1		Positioning with the alignment light is accurate to within 2mm					TBD					HVL Measured kV:
51		DHEC RHB 4.11.2.5.1		Technologist has a clear view of the patient during the scan					TBD				Reconstructed images are free of artifacts					HVL Measured:
52				Patient communication system is operational					TBD				Tomographic plane is indicated by a light source					mGy/mAs:
53					Visual indication of x-ray production is present and functional					High contrast resolution:
54					Positioning with the alignment light is accurate to within 2mm
55					Technologist has a clear view of the patient during the scan					Comment Page 1
56					Patient communication system is operational
57
58
59					Measurement Protocols
60					CTDI Measurements
61					kVp:				Center (mGy)	Periph (mGy)
62					mA:			1				Indicated CTDI (mGy):
63					Time (s):			2				CTDIw (mGy):
64					Beam width (mm):			3				Previous CTDIw (mGy)
65		Date:					Inspector:	Eugene Mah			Chamber length (mm):	100.00		Average:				CTDI change:
66		Room Number					Survey ID:				Chamber CF:	1.00		CTDI (mGy):				Acceptable:	NA
1					Medical University of South Carolina			Criteria:	Change in CTDI from previous year is less than 15%
2					Inspection Results					Dental CT System Compliance Inspection
3		CTDI Measurements					Half Value Layer
4			kVp:					2	Use DHEC standards? (Type 1 to base acceptance on DHEC criteria)
5			mA:			CTDIw (mGy):
6			Time (s):			CTDI change:					kV	mA	mm Al	mGy		Data Processing
7			Beam width (mm):			Acceptable:	NA					0.0			mm Al	Avg Exp	DATA(1/2)	Data Al	HVL
8			Chamber length (mm):	100.00					0.0			0.0
9			Chamber CF:	1.00		Criteria:	Change in CTDI from previous year is less than 15%
10
11		Half Value Layer					Data Exp	LN(EXP)
12		Set kVp:
13
14		Al (mm)	mGy	mGy		Measured HVL:
15		0.0				HVL Acceptable:					Acceptable (mm Al)			HVL Acceptable:
16					Low	Err:502
17					ACCEPTABLE RESULTS (mm Al)					Desired	Err:502		Previous HVL @ kV:
18					Min acceptable HVL at kVp:	Err:502					High	Err:502
19					Desired HVL at kVp:	Err:502
20					Criteria:	MUSC Standard					Exposure Linearity
21					kV:
22
23					Meas kV	mA	Time (s)	mAs	mGy	mGy/mAs		kV Accuracy:
24					kV Coeff of Var:
25					mGy/mAs:
26					Coeff of Linearity:
27					Previous mGy/mAs:
28					Critiera:	Coefficient of linearity <= 10%.					Acceptable:
29					kVp accurate to within 5% of indicated.
30					kVp coefficient of variation (std dev./mean) <= 5%.
31
32					High Contrast Resolution
33					lp/cm	Acceptable	Previous
34					Critiera:	7 lp/cm
35
36
37
38					Comments Page 1
39					Previous Year Comments
40					Comments:
41					Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
42					Additional Comments:
43					Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
44					Additional Comments:
45					Do not exceed 390 chars		0			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
46					Additional Comments:
47		Exposure Linearity					Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
48		kVp:					Additional Comments:
49					Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
50			Meas kV	mAs	mGy	mGy/mAs		kV Accuracy:					Additional Comments:
51					kV Coeff of Var:					Do not exceed 390 chars		0
52					mGy/mAs:					Additional Comments:
53					Coeff of Linearity:					Do not exceed 390 chars		0
54					Acceptable:					Additional Comments:
55			Critiera:	Coefficient of linearity <= 10%.					Do not exceed 390 chars		0
56				kVp accurate to within 5% of indicated.					Additional Comments:
57				kVp coefficient of variation (std dev./mean) <= 5%.					Do not exceed 390 chars		0
58					Additional Comments:
59		High Contrast Resolution					Do not exceed 390 chars		0
60		lp/cm:					Additional Comments:
61		Acceptable:					Do not exceed 390 chars		0
62					Additional Comments:
63		Critiera:	7 lp/cm					Do not exceed 390 chars		0
64					Additional Comments:
65		Date:					Inspector:	Eugene Mah			Do not exceed 390 chars		0
66		Room Number					Survey ID:				Additional Comments:
1					Medical University of South Carolina			Do not exceed 390 chars		0
2					Comments Page 1					Dental CT System Compliance Inspection			Additional Comments:
3					Do not exceed 390 chars		0
4					Additional Comments:
5					Do not exceed 390 chars		0
6					Additional Comments:
7					Do not exceed 390 chars		0
8					Additional Comments:
9					Do not exceed 390 chars		0
10					Additional Comments:
11					Do not exceed 390 chars		0
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
Half Value Layer
mm Al
Exposure (mGy)