## Fluoro
1					Revision 1.3-20180816					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Fluoroscopy System Compliance Inspection					Fluoroscopy System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Number of X-ray Tubes:					Survey ID:				Number of X-ray Tubes:					Survey ID:
14					Room Number:				Accession Number:					Room Number:
15			X-Ray Generator					Site Number:
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18				Max kVp:					Max mA:					Model:				Serial Number:					Number of X-ray Tubes:
19					Software Version:					Max kVp:				Max mA:					Accesssion Number:
20			X-Ray Tube 1					Software Version:					X-Ray Generator
21				Tube Designation/Use:					Focal Spot Sizes (mm)					X-Ray Tube 1					Manufacturer:
22				Manufacture Date:					Large:					Tube Designation/Use:				Focal Spot Sizes (mm)					Model:
23			Insert					Small:					Manufacture Date:				Large:					Max kVp:
24				Manufacturer:					Micro:					Insert					Small:					Manufacture Date:
25				Model:					Manufacturer:				Micro:					Serial Number:
26				Serial Number:					Filtration					Model:					Max mA:
27			Housing					Inherent:					Serial Number:				Filtration					Software Version:
28				Manufacturer:					Added:					Housing					Inherent:					X-Ray Tube 1
29				Model:					Manufacturer:				Added:					Tube Designation/Use:
30				Serial Number:					Model:					Manufacture Date:
31					Serial Number:					Manufacturer:
32			X-Ray Tube 2					Model:
33				Tube Designation/Use:					Focal Spot Sizes (mm)					X-Ray Tube 2					Serial Number:
34				Manufacture Date:					Large:					Tube Designation/Use:				Focal Spot Sizes (mm)					Manufacturer:
35			Insert					Small:					Manufacture Date:				Large:					Model:
36				Manufacturer:					Micro:					Insert					Small:					Serial Number:
37				Model:					Manufacturer:				Micro:					Large:
38				Serial Number:					Filtration					Model:					Small:
39					Inherent:					Serial Number:				Filtration					Micro:
40				Manufacturer:					Added:					Housing					Inherent:					Inherent:
41				Model:					Manufacturer:				Added:					Added:
42				Serial Number:					Model:					X-Ray Tube 2
43					Serial Number:					Tube Designation/Use:
44					Manufacture Date:
45					Manufacturer:
46					Inspection Results					Model:
47					Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
48					Labels, Notices, Postings		Tube Number:					Manufacturer:
49					DHEC Registration sticker is present, clearly visible and legible					**Input only one of the following				Model:
50					DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					numbers or names				Serial Number:
51					Radiation warning label posted on the generator control panel					1	1			Large:
52					Tube loading charts for each focal spot are posted					2	2			Small:
53					Tube loading charts are correct for installed tube(s).					3	3			Micro:
54					Operator manuals are available.					AP	AP			Inherent:
55					Monthly radiation monitoring reports are posted.					Front	Front			Added:
56					Fluoroscopy Systems			Frontal	Frontal			Tube Number:
57					Interlock systems function to prevent fluoro while II is parked					Lat	Lat			Tube Designation:		Change	1
58					The minimum Source Skin Distance > 38 cm (20 cm for surgical applications)					Lateral	Lateral					Change	2
59					Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels)					Change	3
60					The useful beam is intercepted by the image assembly at any SID					Change	AP
61					Fluoroscopy patient dose charts are posted					Change	Front
62					Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Change	Frontal
63					1	Max Entrance Exposure Rate <= 176 mGy/min (high level)					Change	Lat
64					3	High contrast resolution within MUSC specifications					Change	Lateral
65		Date:					Inspector:	Eugene Mah			Low contrast resolution within MUSC specifications
66		Room Number					Survey ID:			1	Transmission through II < 2 uGy/hr for each mGy/min at patient entrance
1					Medical University of South Carolina			Beam limiting device available (functioning) to restrict the x-ray field					Receptor Size
2					Inspection Results					Fluoroscopy System Compliance Inspection			The useful beam is centered on the image receptor					Normal:
3		Rule Number			Labels, Notices, Postings					Compliance				Collimator edges are visible at full open					Mag 1:
4		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD				Elapsed timer (5 min maximum) functional					Mag 2:
5		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD				Fluoroscopic kV and mA meter present and functional					Mag 3:
6		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Radiation Safety					Mag 4:
7				Tube loading charts for each focal spot are posted					TBD				Continuous pressure on exposure switch required to make an exposure with selected tube					Pt Dose Measured Fluoro
8				Tube loading charts are correct for installed tube(s).					TBD				Operator has full visibility of the exposure factors and patient					Manual/Auto:
9				Operator manuals are available.					TBD				Lead aprons available					Lucite/Al:		Change	1
10				Monthly radiation monitoring reports are posted.					TBD				Lead gloves available					# dose rates:
11					Properly designed and installed apron rack is present					High level present:
12					Radiation Safety					Documentation of annual protective apparel integrity inspection is available					cm/in:		Change	cm
13				Continuous pressure on exposure switch required to make an exposure with selected tube					TBD				Patient restraint devices available					Auto/mA:		Change	Auto
14				Operator has full visibility of the exposure factors and patient					TBD				Gonadal shielding is present and functional					Technique 1:
15				Lead aprons available					TBD				Radiation warning light at room entrance is functional with selected tube					Technique 2:
16				Lead gloves available					TBD					Physical Inspection					Technique 3:
17				Properly designed and installed apron rack is present					TBD				System clean and free of debris					Maximum Dose Rates Fluoro
18				Documentation of annual protective apparel integrity inspection is available					TBD				All cables free from excessive wear or restraints					kV:			0.00
19				Patient restraint devices available					TBD				System free of missing or damaged components					mA:			0.00
20				Gonadal shielding is present and functional					TBD				Tube support movements and locks are acceptable					mGy/min:			0.00
21				Radiation warning light at room entrance is functional with selected tube					TBD				All other movements and locks functional					kV:			0.00
22					Automatic Brightness Control / Display Monitor					mA:			0.00
23					Fluoroscopy Systems					Automatic Brightness Control maintains image brightness for varying attenuation					mGy/min:			0.00
24		DHEC RHB 4.9.2.1.2		Interlock systems function to prevent fluoro while II is parked					TBD				Display monitor is free of artifacts, shadows, bright spots, and burn-in					kV:			0.00
25		DHEC RHB 4.9.1		The minimum Source Skin Distance > 38 cm (20 cm for surgical applications)					TBD				Display monitor is free of distortion					mA:			0.00
26		DHEC RHB 4.9.8		Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels)					TBD				Display monitor is free of excessive persistence (after-image)					mGy/min:			0.00
27		DHEC RHB 4.9.2.1.2		The useful beam is intercepted by the image assembly at any SID					TBD				Display monitor is free of static or rolling noise bands					Maximum Dose Rates Pulse/Dig
28		DHEC RHB 4.9.4.3.6		Fluoroscopy patient dose charts are posted					TBD				The last image hold (LIH) display is free of jitter or abnormal appearance					kV:			0.00
29		DHEC RHB 4.9.4.2.1		Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					TBD					mA:			0.00
30		DHEC RHB 4.9.4.2.1.2		Max Entrance Exposure Rate <= 176 mGy/min (high level)					YES					mGy/min:			0.00
31				High contrast resolution within MUSC specifications					NA					Inspection Results					kV:			0.00
32				Low contrast resolution within MUSC specifications					TBD					mA:			0.00
33		DHEC RHB 4.9.5.1		Transmission through II < 2 uGy/hr for each mGy/min at patient entrance					YES			Entrance Exposure Rate				Type 1 for manual, leave blank for auto			Auto	Auto	Auto/mA mode			mGy/min:			0.00
34		DHEC RHB 4.9.2.2.1		Beam limiting device available (functioning) to restrict the x-ray field					TBD					1	1	Type 1 if using Lucite, leave blank if using Al					kV:			0.00
35				The useful beam is centered on the image receptor					TBD					How many dose rate settings are present (exclusing High Level)?					mA:			0.00
36		DHEC RHB 4.9.2.2.1		Collimator edges are visible at full open					TBD					Is High Level provided? (1=Yes, 2=No)					mGy/min:			0.00
37		DHEC RHB 4.9.7		Elapsed timer (5 min maximum) functional					TBD			Blank					HVL
38		DHEC RHB 4.9.6		Fluoroscopic kV and mA meter present and functional					TBD			cm	cm	Unit for image receptor size					HVL Measured kV:
39		DHEC RHB 4.3.5		The Measured HVL meeds DHEC minimum requirements					TBD			Patient Entrance Exposure Rate (Fluoroscopy)*					HVL Measured:		Change	TBD
40				Kerma area product display is accurate to within 20%					NA			Image	Acrylic					Pt Dose Measured Dig Acq
41				Air kerma display is accurate to within 30%					NA			Mode	Attenuator	Auto 1					Auto				Technique 1:
42				Air kerma rate display is accurate to within 30%					NA			(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min		Technique 2:
43					Normal	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Technique 3:
44					Automatic Brightness Control / Display Monitor					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Receptor Entr Exp Fluoro
45				Automatic Brightness Control maintains image brightness for varying attenuation					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Normal
46				Display monitor is free of artifacts, shadows, bright spots, and burn-in					TBD			Mag 1	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 1
47				Display monitor is free of distortion					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 2
48				Display monitor is free of excessive persistence (after-image)					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 3
49				Display monitor is free of static or rolling noise bands					TBD			Mag 2	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1 Mag 4
50				The last image hold (LIH) display is free of jitter or abnormal appearance					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Normal
51					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 1
52					Physical Inspection					Mag 3	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 2
53				System clean and free of debris					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 3
54				All cables free from excessive wear or restraints					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Mag 4
55				System free of missing or damaged components					TBD			Mag 4	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Normal
56				Tube support movements and locks are acceptable					TBD				20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 1
57				All other movements and locks functional					TBD				30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 2
58					max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 3
59					Last Year:					Auto Mag 4
60					*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	YES		Receptor Entr Exp Dig Acq
61					Undertable fluoroscopy units: 1 cm above table top					Auto 1 Normal
62					Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Auto 1 Mag 1
63					C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					Auto 1 Mag 2
64					Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					Auto 1 Mag 3
65		Date:					Inspector:	Eugene Mah			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					Auto 1 Mag 4
66		Room Number					Survey ID:					Max Entrance Exposure Rate < = 176 mGy/min (high level)					Normal
1					Medical University of South Carolina			Comments:					Mag 1
2					Inspection Results					Fluoroscopy System Compliance Inspection			Do not exceed 390 chars		0					Mag 2
3					Half Value Layer Measurement					New:					Mag 3
4					Mag 4
5				Set kVp:					Patient Entrance Exposure Rate – Pulse Fluroscopy					Mode:	1	Enter 1 for Pulse fluoro, 2 for digital acquisition mode					Auto Normal
6					Image	Acrylic					Auto Mag 1
7				Al (mm)	mGy/min	mGy/min		Measured HVL:	TBD					Mode	Attenuator	Auto 1					Auto				Auto Mag 2
8				0.0				HVL Acceptable:	TBD					(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min		Auto Mag 3
9					Normal	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto Mag 4
10					ACCEPTABLE RESULTS (mm Al)					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Fluoro Scatter/Transmission
11					Min acceptable HVL at kVp:	Err:502					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Transmission:			0.00
12					Desired HVL at kVp:	Err:502				Mag 1	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (eyes):
13					Criteria:	MUSC Standard					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (waist):
14					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (eyes):
15					Mag 2	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (waist):
16					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Dig Acq Scatter
17					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (eyes):
18					Mag 3	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Unshielded Scatter (waist):
19					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (eyes):
20					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Shielded Scatter (waist):
21					Mag 4	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Leeds Phantom
22					20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		kVp:		Change	70
23					30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Huttner Normal:
24					max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Huttner Mag 1:
25					Last Year:					Huttner Mag 2:
26					*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	YES		Huttner Mag 3:
27					Undertable fluoroscopy units: 1 cm above table top					Huttner Mag 4:
28					Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					N3 Normal:
29					C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					N3 Mag 1:
30					Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					N3 Mag 2
31					Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					N3 Mag 3
32					Max Entrance Exposure Rate < = 176 mGy/min (high level)					N3 Mag 4:
33					Comments:					Field Size:
34					Do not exceed 390 chars		0					SID:
35					New:					Front:
36					Back:
37					Half Value Layer					Left:
38					2	Use DHEC standards? (Type 1 to base acceptance on DHEC criteria)					Right:
39					TO.10.A
40					kV	mA	mm Al	mGy/min		Data Processing					TO.10.B
41					0.0			mm Al	Avg Exp	DATA(1/2)	Data Al	HVL			TO.10.C
42					0.0			0.00				TBD			TO.10.D
43					TO.10.E
44					TO.10.F
45					Data Exp	LN(EXP)			TO.10.G
46		Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate					TO.10.H
47				Indicated KAP µGy·m²	Measured KAP µGy·m²	KAP Deviation	Ind AK @ Ref Pt (mGy)	Meas AK @ Ref Pt (mGy)	AK Deviation	Ind AKR	Meas AKR	AKR Deviation					TO.10.J
48		kV	mA					(mGy/min)	(mGy/min)					TBD	TBD					TO.10.K
49					Acceptable (mm Al)			HVL Acceptable:	TBD			TO.10.L
50					Low	Err:502					TO.10.M
51					Desired	Err:502		Previous HVL @ kV:				TO.10.A
52					High	Err:502					TO.10.B
53					TO.10.C
54					Image Receptor Entrance Exposure Rate – Fluoroscopy					TO.10.D
55					TO.10.E
56					Dose	Receptor					Previous Year					TO.10.F
57					Acceptable:			Acceptable:			Acceptable:			Level	Size (cm)	kVp	mA	uGy/min	nGy/frame	nGy/frame					TO.10.G
58			Criteria:	Measured KAP/AK/AKR should be within 40% of the displayed values					Auto 1					Measure using pancake probe placed against				TO.10.H
59					image receptor input window				TO.10.J
60					Use 1.5mm Cu filtration				TO.10.K
61					TO.10.L
62					TO.10.M
63					TO.10.A
64					TO.10.B
65		Date:					Inspector:	Eugene Mah					TO.10.C
66		Room Number					Survey ID:					TO.10.D
1					Medical University of South Carolina					TO.10.E
2					Inspection Results					Fluoroscopy System Compliance Inspection		Auto					TO.10.F
3		Patient Entrance Exposure Rate – Fluoroscopy					TO.10.G
4		Image	Acrylic					TO.10.H
5		Mode	Attenuator	Auto 1					Auto					TO.10.J
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min					TO.10.K
7		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.L
8			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00			Comments:					TO.10.M
9			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					0					TO.10.A
10		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00			New:					TO.10.B
11			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.C
12			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy					TO.10.D
13		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.E
14			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Dose	Receptor					Previous Year					TO.10.F
15			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame	nGy/frame					TO.10.G
16		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Auto 1					Measure using pancake probe placed against				TO.10.H
17			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					image receptor input window				TO.10.J
18			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					Use 1.5mm Cu filtration				TO.10.K
19		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.L
20			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.M
21			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.A
22			max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00					TO.10.B
23					TO.10.C
24			*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	YES					TO.10.D
25			Undertable fluoroscopy units: 1 cm above table top					TO.10.E
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)					Auto					TO.10.F
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point					TO.10.G
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)					TO.10.H
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)					TO.10.J
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)					TO.10.K
31		Comments:					TO.10.L
32					Comments:					TO.10.M
33					Do not exceed 390 chars		0					Comments
34					New:					Pt Entr Exp Fluoro:
35		Image Receptor Entrance Exposure Rate – Fluoroscopy					Pt Entr Exp Dig Acq:
36					Scatter and Transmission					Receptor Entr Exp Fluoro:
37		Dose	Receptor					Pt Entrance					Receptor Entr Exp Dig Acq:
38		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame					kVp	mA	mGy/min	uGy/hr	Last Year		Acceptable				Scatter/Transmission:
39		Auto 1					Measure using pancake probe placed against					Transmission* @ 10cm	0	0	0.00	0.00			YES				Leeds Phantom:
40					image receptor input window					Unshielded	Shielded	Unshielded	Shielded	Unshielded	Shielded			TO.10:
41					Use 1.5mm Cu filtration					uGy/min	uGy/min	uGy/hr	uGy/hr	uGy/hr	uGy/hr			Beam Limitation:
42					Scatter – Fluoro	Eye level	0	0					Comment Page 1
43					Waist level	0	0
44					Scatter - Pulse	Eye level	0	0
45					Waist level	0	0
46					*Measured with 8” Lucite in beam
47					Largest field size mode, highest dose rate mode
48					Criteria:	Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure
49		Auto					Comments:
50					Do not exceed 390 chars		0
51					New:
52
53					Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate
54		Comments:					0	Source-PDC distance (cm)				Source-Isocenter
55					0	Source-Ref Pt distance (cm)				Table-Isocenter
56					CAK Reference Point correction factor				RefPt-Isocenter
57		Scatter and Transmission – Fluoroscopy				Pt Entrance					1	Apply criteria
58				kVp	mA	mGy/min	uGy/hr	Unshielded	Shielded		Acceptable					Start KAP	End KAP		Start AK	End AK		Ind AKR	Meas AKR
59			Transmission* @ 10cm	0	0	0	0.00	uGy/hr	uGy/hr		YES			kV	mA	Gy·cm²		KAP	mGy		CAK	mGy/min	mGy/min
60		Scatter – Fluoro	Eye level	0	0
61			Waist level	0	0
62		Comments:
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection		Radcal PDC				Calibration date:
3		Patient Entrance Exposure Rate – Pulse Fluroscopy					kV	mA	mGy	mGy/Min	µGy·m²	µGy·m²/Min
4		Image	Acrylic
5		Mode	Attenuator	Auto 1					Auto
6		(cm)	(cm)	kVp	mA	mGy/min	kVp	mA	mGy/min	kVp	mA	mGy/min
7		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
8			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
9			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
10		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
11			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
12			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00		Indicated KAP µGy·m²	Measured KAP µGy·m²		Indicated AK @ Ref Pt	Measured AK @ Ref Pt		Ind AKR	Meas AKR
13		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00				KAP Deviation			CAK Deviation	mGy/min	mGy/min	Deviation				Comment Page 2
14			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
15			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
16		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
17			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
18			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
19		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20			20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
21			30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
22			max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00			Acceptable:			Acceptable:			Acceptable:
23
24		Comments:					Criteria:	Measured KAP/AK/AKR should be within 40% of the displayed values
25					Leeds Phantom
26
27					Receptor Size	MS1	MS3	MS4	Huttner	N3	kVp:	70	70
28					MS1-4		Huttner: Enter col.row
29					1	Clearly visible	N3: Enter # of visible disks
30					2	Visible
31					3	Barely visible
32		Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy					4	Not visible
33					Acceptable:	TBD	TBD
34		Dose	Receptor					Comments:
35		Level	Size (cm)	kVp	mA	uGy/min	nGy/frame					Do not exceed 390 chars		0
36		Auto 1					Measure using pancake probe placed against					New:
37					image receptor input window
38					Use 1.5mm Cu filtration					TO.10		Details Detected
39					Row
40					A
41					B
42					C
43					D
44					E
45					F
46		Auto					G
47					H
48					J
49					K
50					L
51		Comments:					M
52
53					Comments:
54					Do not exceed 390 chars		0
55		Scatter and Transmission – Pulse Fluoroscopy					New:
56					Unshielded	Shielded
57				kVp	mA	uGy/hr	uGy/hr					Field Size
58		Scatter - Pulse	Eye level	0	0					Horizontal					Type 1 for circular field, 2 for square field
59			Waist level	0	0					Vertical
60					Diagonal
61					Acceptable	NA	NA	NA	NA	NA
62		Comments:					Criteria:	Measured field size is at least 85% of nominal field size
63					Useful Beam Limitation
64					SID (cm)
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:				Measured	Difference	Sum	Last Year			T	B	L	R
1					Medical University of South Carolina		Top					Visible
2					Inspection Results					Fluoroscopy System Compliance Inspection		Bottom					X-ray
3		Leeds Test Objects					High contrast resolution		Low contrast resolution					Left					Acceptable	SID
4			Receptor Size	MS1	MS3	MS4	lp/mm	Acceptable	N3	Acceptable				Right					SOD
5					Criteria:	Difference between x-ray field and visible image area is less than 2% SID					SID/SOD
6					Sum of length and width differences is less than 4% SID
7					Comments:
8					Do not exceed 390 chars		0
9					New:
10			Criteria:	Resolution: >1.0 lp/mm Normal mode, >1.25 lp/mm Mag 1, >1.4 lp/mm Mag 2, > 2 lp/mm Mag 3, > 2.5 lp/mm Mag 4
11				N3: Minimum contrast level < 0.04 in all II modes
12		Comments:					Comments Page 1
13					Previous Year Comments
14					Comments:
15					Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
16		TO.10 Contrast Resolution					Additional Comments:
17			Contrast Detail					Threshold Index					Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
18		Row					Additional Comments:
19		A					Do not exceed 390 chars		0			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
20		B					Additional Comments:
21		C					Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
22		D					Additional Comments:
23		E					Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
24		F					Additional Comments:
25		G					Do not exceed 390 chars		0
26		H					Additional Comments:
27		J					Do not exceed 390 chars		0
28		K					Additional Comments:
29		L					Do not exceed 390 chars		0
30		M					Additional Comments:
31		Comments:					Do not exceed 390 chars		0
32					Additional Comments:
33					Do not exceed 390 chars		0
34					Additional Comments:
35		Field Size					Do not exceed 390 chars		0
36					Additional Comments:
37			Horizontal					Do not exceed 390 chars		0
38			Vertical					Additional Comments:
39			Diagonal					Do not exceed 390 chars		0
40			Acceptable	NA	NA	NA	NA	NA					Additional Comments:
41			Criteria:	Measured field size is at least 85% of nominal field size					Do not exceed 390 chars		0
42					Additional Comments:
43		Useful Beam Limitation					Do not exceed 390 chars		0
44			SID (cm)					Additional Comments:
45				Measured	Difference	Sum					Do not exceed 390 chars		0
46			Top					Additional Comments:
47			Bottom					Acceptable					Do not exceed 390 chars		0
48			Left					Additional Comments:
49			Right					Do not exceed 390 chars		0
50			Criteria:	Difference between x-ray field and visible image area is less than 2% SID					Additional Comments:
51				Sum of length and width differences is less than 4% SID					Do not exceed 390 chars		0
52		Comments:
53
54					Comments Page 2
55					Previous Year Comments
56					Comments:
57					Do not exceed 390 chars		0
58					Additional Comments:
59					Do not exceed 390 chars		0
60					Additional Comments:
61					Do not exceed 390 chars		0
62					Additional Comments:
63					Do not exceed 390 chars		0
64					Additional Comments:
65		Date:					Inspector:	Eugene Mah			Do not exceed 390 chars		0
66		Room Number					Survey ID:				Additional Comments:
1					Medical University of South Carolina			Do not exceed 390 chars		0
2					Comments Page 1					Fluoroscopy System Compliance Inspection			Additional Comments:
3					Do not exceed 390 chars		0
4					Additional Comments:
5					Do not exceed 390 chars		0
6					Additional Comments:
7					Do not exceed 390 chars		0
8					Additional Comments:
9					Do not exceed 390 chars		0
10					Additional Comments:
11					Do not exceed 390 chars		0
12					Additional Comments:
13					Do not exceed 390 chars		0
14					Additional Comments:
15					Do not exceed 390 chars		0
16					Additional Comments:
17					Do not exceed 390 chars		0
18					Additional Comments:
19					Do not exceed 390 chars		0
20					Additional Comments:
21					Do not exceed 390 chars		0
22					Additional Comments:
23					Do not exceed 390 chars		0
24					Additional Comments:
25					Do not exceed 390 chars		0
26					Additional Comments:
27					Do not exceed 390 chars		0
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Comments Page 2					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Radiation Output – Fluoroscopy					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Radiation Output - Pulse Fluoroscopy					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Contrast Resolution					Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
Half Value Layer
mm Al
Exposure Rate (mGy/min)
Radiation Output
kV
mGy/mA-min
Fluoro Patient Entrance Exposure
Attenuator thickness (cm)
mGy/min
Radiation Output
kV
mGy/mA-min
Fluoro Patient Entrance Exposure
Attenuator thickness (cm)
mGy/min
Contrast Detail Diagram
Detail diameter (mm)
Threshold Detection Index
sqrt(Detail area)

## ExpCharts
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Fluoroscopy
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	R/min	kVp	mA	R/min	kVp	mA	R/min
7		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
8		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
9		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
10		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
11		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
12		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
13		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
14		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
15		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
16		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
17		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
18		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
19		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
21		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
22			max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
23
24			*Measured at:		Acceptable:	NA		Acceptable:	NA		Acceptable:	YES
25			Undertable fluoroscopy units: 1 cm above table top
26			Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27			C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point
28			Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29			Criteria:	Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)
30				Max Entrance Exposure Rate < = 176 mGy/min (high level)
31			Comments:
32
33
34
35		Image Receptor Entrance Exposure Rate – Fluoroscopy
36
37		Dose	Receptor
38		Level	Size (cm)	kVp	mA	mR/min	nGy/frame
39		Auto 1					Measure using pancake probe placed against
40		0.00					image receptor input window
41		0.00					Use 1.5mm Cu filtration
42		0.00
43		0.00
44
45		0.00
46		0.00
47		0.00
48		0.00
49		Auto
50		0.00
51		0.00
52		0.00
53		0.00
54			Comments:
55
56
57		Scatter and Transmission – Fluoroscopy				Pt Entrance	Unshielded	Shielded
58				kVp	mA	mGy/min	uGy/hr	uGy/hr		Acceptable
59			Transmission* @ 10cm	0.00	0.00	0.00				YES
60		Scatter – Fluoro	Eye level	0.00	0.00
61		0.00	Waist level	0.00	0.00
62			Comments:
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:
1					Medical University of South Carolina
2					Inspection Results					Fluoroscopy System Compliance Inspection
3		Patient Entrance Exposure Rate – Pulse Fluroscopy
4		Image	Acrylic
5		Mode	Attenuator	Auto 1	0.00	0.00		0.00	0.00	Auto	0.00	0.00
6		(cm)	(cm)	kVp	mA	R/min	kVp	mA	R/min	kVp	mA	R/min
7		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
8		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
9		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
10		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
11		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
12		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
13		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
14		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
15		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
16		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
17		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
18		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
19		cm	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20		0.00	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
21		0.00	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
22			max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
23
24			Comments:
25
26
27
28
29
30
31
32		Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy
33
34		Dose	Receptor
35		Level	Size (cm)	kVp	mA	mR/min	nGy/frame
36		Auto 1					Measure using pancake probe placed against
37		0.00					image receptor input window
38		0.00					Use 1.5mm Cu filtration
39		0.00
40		0.00
41
42		0.00
43		0.00
44		0.00
45		0.00
46		Auto
47		0.00
48		0.00
49		0.00
50		0.00
51			Comments:
52
53
54
55		Scatter and Transmission – Pulse Fluoroscopy
56					Unshielded	Shielded
57				kVp	mA	uGy/hr	uGy/hr
58		Scatter - Pulse	Eye level	0.00	0.00
59		0.00	Waist level	0.00	0.00
60
61
62			Comments:
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:

## Tables
Leeds MS					Patient Entrance Exposure Rate (Fluoroscopy)*			Sorted by kV					Patient Entrance Exposure Rate – Pulse Fluroscopy			Sorted by kV
1	Clearly visible					kV	Fluoro	High Level	kV	Fluoro	High Level			kV	Fluoro	High Level	kV	Fluoro	High Level
2	Visible					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
3	Barely visible					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
4	Not visible					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
				0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
Huttner Grating			Leeds N3		65.00	70.00	75.00	80.00	Nominal					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
1.1	0.50		1	0.1670	0.1880	0.1670	0.1490	0.1350	0.1600					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
1.2	0.56		2	0.1480	0.1670	0.1480	0.1320	0.1200	0.1450					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
1.3	0.63		3	0.1280	0.1450	0.1280	0.1140	0.1040	0.1230					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
2.1	0.71		4	0.1090	0.1240	0.1090	0.0970	0.0880	0.1080					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
2.2	0.80		5	0.0876	0.0993	0.0876	0.0780	0.0705	0.0860					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
2.3	0.90		6	0.0749	0.0847	0.0749	0.0670	0.0609	0.0760					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
3.1	1.00		7	0.0674	0.0766	0.0674	0.0599	0.0542	0.0660					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
3.2	1.12		8	0.0525	0.0593	0.0525	0.0470	0.0427	0.0550					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
3.3	1.25		9	0.0450	0.0512	0.0450	0.0399	0.0360	0.0450					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
4.1	1.40		10	0.0371	0.0400	0.0371	0.0347	0.0329	0.0390					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
4.2	1.60		11	0.0322	0.0347	0.0322	0.0301	0.0285	0.0330				Max	0.00	#DIV/0!	#N/A					Max	0.00	#DIV/0!	#N/A
4.3	1.80		12	0.0246	0.0269	0.0246	0.0233	0.0220	0.0270					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
5.1	2.00		13	0.0215	0.0232	0.0215	0.0201	0.0190	0.0230					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
5.2	2.24		14	0.0172	0.0186	0.0172	0.0161	0.0153	0.0180					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
5.3	2.50		15	0.0155	0.0168	0.0155	0.0145	0.0138	0.0160					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
6.1	2.80		16	0.0130	0.0140	0.0130	0.0122	0.0115	0.0130					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
6.2	3.15		17	0.0110	0.0119	0.0110	0.0103	0.0097	0.0110					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
6.3	3.55		18	0.0086	0.0093	0.0086	0.0081	0.0077	0.0090					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
7.1	4.00		19	0.0066	0.0071	0.0066	0.0062	0.0059	0.0070					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
7.2	4.50					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
7.3	5.00					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
				0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
TO.10 kV Adj	65.00	70.00	75.00	80.00					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.0120	0.0119	0.0110	0.0103	0.0097					.					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.0160	0.0168	0.0155	0.0145	0.0138					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.0230	0.0232	0.0215	0.0201	0.0190					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.0320	0.0347	0.0322	0.0301	0.0285					Max	0.00	#DIV/0!	#N/A					Max	0.00	#DIV/0!	#N/A
0.0450	0.0512	0.0450	0.0399	0.0360					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.0660	0.0766	0.0674	0.0599	0.0542					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.0860	0.0993	0.0876	0.0780	0.0705					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.1230	0.1450	0.1280	0.1140	0.1040					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.1600	0.1880	0.1670	0.1490	0.1350					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.2300	0.2670	0.2380	0.2140	0.1940					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.3500	0.3990	0.3600	0.3270	0.2990					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.5000	0.6180	0.5730	0.5330	0.4970					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.6600	0.7690	0.7260	0.6860	0.6490					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
0.9300	0.9690	0.9540	0.9360	0.9150					0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
				0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
TO.10	Diameter (mm)	sqrt(Area)	1	2	3	4	5	6	7	8	9			0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
A	11.10	9.837	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
B	7.90	7.001	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
C	5.60	4.963	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016	0.012			0.00	#DIV/0!	#N/A					0.00	#DIV/0!	#N/A
D	4.00	3.545	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016		Max	0.00	#DIV/0!	#N/A					Max	0.00	#DIV/0!	#N/A
E	2.80	2.481	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016
F	2.00	1.772	0.230	0.160	0.123	0.086	0.066	0.045	0.032	0.023	0.016
G	1.40	1.241	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
H	1.00	0.886	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
J	0.70	0.620	0.500	0.350	0.230	0.160	0.123	0.086	0.066	0.045	0.032
K	0.50	0.443	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066
L	0.35	0.310	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066
M	0.25	0.222	0.930	0.660	0.500	0.350	0.230	0.160	0.123	0.086	0.066

TO.10 Values	0	1	1.5	2	2.5	3	3.5	4	4.5	5	5.5	6	6.5	7	7.5	8	8.5	9
A	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
B	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
C	#N/A	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016	0.013	0.011
D	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
E	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
F	#N/A	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032	0.027	0.022	0.019	0.016
G	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
H	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
J	#N/A	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067	0.056	0.045	0.039	0.032
K	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067
L	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067
M	#N/A	0.954	0.840	0.726	0.650	0.573	0.467	0.360	0.299	0.238	0.203	0.167	0.148	0.128	0.108	0.088	0.078	0.067

Half Value Layer
	DHEC			MUSC				HVL Graph
	Min HVL			HVL	Min HVL	Max HVL		mm Al	Avg Exp
kVp	(mm Al)		kVp	(mm Al)	(mm Al)	(mm Al)		TBD	TBD
30.00	0.3		60	2.2	1.8	2.4		TBD	TBD
40.00	0.4		70	2.6	2.2	2.9		TBD	TBD
50.00	0.5		80	3.0	2.6	3.5		TBD	TBD
51.00	1.2		90	3.5	3.1	4.0		TBD	TBD
60.00	1.3		100	3.9	3.5	4.5		TBD	TBD
70.00	1.5		110	4.3	3.9	5.1
71.00	2.1		120	5.7	4.4	5.6
80.00	2.3		130	5.2	4.8	6.1
90.00	2.5
100.00	2.7
110.00	3.0
120.00	3.2
130.00	3.5
140.00	3.8
150.00	4.1

Fluoro Patient Entrance Exposure Rate
Acrylic	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min
Attenuator	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.
(cm)
10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00

Dig Acq Patient Entrance Exposure Rate
Acrylic	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min	mGy/min
Attenuator	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.	Not Avail.
(cm)
10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00

KAP Conversion Table
	µGy·m²
cGy·cm²	1.00
Gy·cm²	100.00
mGy·cm²	0.01
µGy·m²	1.00

## DataEntry
Patient Entrance Exposure Rate (Fluoroscopy)*
Image	Aluminum
Mode	Attenuator
(cm)	(cm)	kVp	mA	uGy/pls	kVp	mA	uGy/pls	kVp	mA	uGy/pls
Normal	10.00
	20.00
	30.00
Mag 1	10.00
	20.00
	30.00
Mag 2	10.00
	20.00
	30.00
Mag 3	10.00
	20.00
	30.00
Mag 4	10.00
	20.00
	30.00
	max

Patient Entrance Exposure Rate (Digital Acquisition)*
Image	Aluminum
Mode	Attenuator
(cm)	(cm)	kVp	mA	uGy/pls	kVp	mA	uGy/pls	kVp	mA	uGy/pls
Normal	10.00
	20.00
	30.00
Mag 1	10.00
	20.00
	30.00
Mag 2	10.00
	20.00
	30.00
Mag 3	10.00
	20.00
	30.00
Mag 4	10.00
	20.00
	30.00
	max

## DataPage
Sheet	FLUORO
SurveyID
HVL		TBD
Field size

Dose mode	Auto 1

	Auto
SEE	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
Max SEE	max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
Dose mode	Auto 1

	Auto
SEE	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	10.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	20.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
	30.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
Max SEE	max	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00	0.00
IEER		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
IEER		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
		0.00	0.00	0.00	0.00
TO.N3

TO.10	Field size	A	B	C	D	E	F	G	H	J	K	L	M

Resolution