## QC Test Summary-Hologic
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Lorad/Hologic
Site Name | [8 empty] | Report Date
Address | [8 empty] | Survey Date | 0
Medical Physicist's Name |  |  | Eugene Mah | [5 empty] | Signature
X-Ray Unit Manufacturer |  |  | Lorad/Hologic | [5 empty] | Model | Selenia Dimensions
Date of Installation | [8 empty] | Room ID
QC Manual Version # |  |  | MAN-03706 Rev 011 |  |  |  |  | (use any version applicable to model; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  |  | Model |  | Location |  | QC Manual Version #
|  | Review Workstation* | Barco |  |  | MDMC-12133 |  |  |  | Barco K5905277 v18
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System (www.accessdata.fda.gov/cdrh_docs/presentations/pghs/Polic_Guidance_Help_System.htm).
Survey Type |  |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [9 empty] | Annual Survey
Features |  |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Mammographic Unit Assembly Evaluation
2. Collimation Assessment
3. Artifact Evaluation
4. kVp Accuracy and Reproducibility
5. Beam Quality Assessment - HVL Measurement
6. Evaluation of System Resolution
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
| Average glandular dose for average breast is ≤3 mGy (300 mrad) (conventional) | [9 empty] | mGy
| Average glandular dose for average breast is ≤3 mGy (300 mrad) (DBT) | [8 empty] | #N/A | mGy
9. Radiation Output Rate
10. Phantom Image Quality Evaluation | [5 empty] | Fibers | Specks | Masses
Phantom image scores (conventional)
Phantom image scores (DBT)
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
|  | SNR (value) | #DIV/0!
|  | CNR (value) | #DIV/0! | (required for new unit MEE and Annual Survey)
CNR should not vary by more than ±15% (NA for MEE)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
13. DICOM Printer QC (if applicable, MEE only) | [11 empty] | NA
14. Detector Flat Field Calibration (MEE only)
15. Geometry Calibration For Tomosynthsis (DBT MEE only)
16. Compression Thickness Indicator (MEE only)
17. Compression (MEE only)
18. Detector Ghosting (troubleshooting only) | [11 empty] | NA
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M22: Click in boxes to use drop-down lists
Note on I34: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Hologic
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Lorad/Hologic, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | DICOM Printer Quality Control (if applicable) | Weekly | NA
2. | Viewboxes and Viewing Conditions | Weekly
3. | Artifact Evaluation | Weekly
4. | Signal-To-Noise and Contrast-To-Noise Measurements | Weekly
5. | Phantom Image Quality Evaluation | Weekly
6. | Detector Flat-Field Calibration | Weekly
7. | Compression Thickness Indicator | Bi-weekly
8. | Visual Checklist | Monthly
9. | Repeat/Reject Analysis | Quarterly
10. | Compression | Semi-annually
11. | Geometry Calibration (Tomosynthsis Option) (DBT) | Semi-annually
12. | Diagnostic Review Workstation QC (NA if only hardcopy read) | See Hologic QC Manual
13. | Mobile Unit Quality Control (if applicable) | After every move
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use any QC manual version that is applicable to the Lorad model surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for Lorad FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## SBB QC Test Summary
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY
Site Name | HCC Breast Imaging Program |  |  |  |  | Report Date
Address | 86 Jonathan Lucas St, Charleston, SC |  |  |  |  | Survey Date | 0
X-Ray Unit Manufacturer | [5 empty] | Model
Date of Installation | [5 empty] | Room ID
Film (mfr & type) |  | NA |  |  |  | Screen (mfr & type) | NA
Film Processor Mfr |  | NA |  |  |  | Model | NA
Digital Image Receptor Mfr | [5 empty] | Model
Medical Physicist's Name |  | Eugene Mah |  |  |  | Signature
Medical Physicist's QC Tests
PASS/FAIL
1. Stereotactic Breast Biopsy Unit Assembly Evaluation
2. Collimation Assessment
A. X-ray field adequately matches image receptor
B. Biopsy window generally centered over digital image receptor (NA if film used)
3. Focal Spot Performance and System Limiting Resolution
| A. Focal spot performance acceptable (NA if digital used) | [7 empty] | NA
B. Digital system spatial resolution acceptable (NA if film used)
4. kVp Accuracy and Reproducibility
Measured average kVp within ±5% of indicated kVp
kVp coefficient of variation ≤ 0.02
5. Beam Quality Assessment (Half-Value Layer Measurement)
HVL is within acceptable lower and upper limits at all kVp values tested
6. AEC System or Manual Exposure Performance Assessment
Optical density or signal range acceptable
7. Receptor Speed Uniformity
Screen speed uniformity or digital receptor uniformity acceptable
8. Breast Entrance Exposure, Average Glandular Dose and Exposure Reproducibility
Exposure reproducibility is within acceptable limits
| Average glandular dose to a 4.2 cm thick breast is ≤ 3 mGy (300 mrad) | [6 empty] | mGy
9. Image Quality Evaluation
Phantom image quality is acceptable
| Phantom type: | ACR Mammography Accreditation Phantom |  |  |  | Mini Phantom
| Phantom image quality scores: |  | Fibers | 0 | Specks | 0 | Masses | 0
10. Artifact Evaluation
Artifacts were not apparent or not significant:
Artifacts identified:
11. Localization Accuracy Test
Localization and sampling accurate/object captured
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M15: Click in boxes to use drop-down lists

## SBB Tech QC Eval
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY (continued)
Evaluation of Site's Technologist QC Program
|  | Frequency | PASS/FAIL
1. | Localization Accuracy Test | Daily
2. | Darkroom Cleanliness (NA if digital used) | Daily | NA
3. | Processor Quality Control (NA if digital used) | Daily | NA
4. | Phantom Images | Weekly
5. | Screen Cleanliness (NA if digital used) | Weekly | NA
6. | Viewboxes and Viewing Conditions (NA if digital used) | Weekly
7. | Hardcopy Output Quality (if hardcopy produced from digital data) | Monthly | NA
8. | Visual Checklist | Monthly
9. | Analysis of Fixer Retention in Film (NA if digital used) | Quarterly | NA
10. | Compression | Semi-annually
11. | Repeat Analysis | Semi-annually
12. | Screen-Film Contact (NA if digital used) | Semi-annually | NA
13. | Darkroom Fog (NA if digital used) | Semi-annually | NA
14. | Zero Alignment Test (if required by manufacturer) | Before each patient | NA
15. | Any additional tests required by manufacturer | As required by manufacturer | Pass
Medical Physicist's Recommendations for Quality Improvement
Comments:
Note on J8: Click in boxes to use drop-down lists

## MQSA Requirements
MEDICAL PHYSICIST'S CHECKLIST MQSA REQUIREMENTS FOR MAMMOGRAPHY EQUIPMENT
Facility Name: | 0
Unit Manufacturer: |  |  | Model:
Serial number: |  |  | Year Mfr:
Medical Physicist: | Eugene Mah |  | Room ID:
Signature: |  |  | Survey Date: | 0
Feature | FDA Rule Section | Requirement | Applies to | Meets FDA Requirements? (if NA, please explain)
Motion of tube-image receptor assembly | 3(i) | The assembly shall be capable of being fixed in any position where it is designed to operate. Once fixed in any such position, it shall not undergo unintended motion. | S-F & FFDM
| 3(ii) | This mechanism shall not fail in the event of power interruption. | S-F & FFDM
Image receptor sizes | 4(i) | Systems using screen-film image receptors shall provide, at a minimum, for operation with image receptors of 18 x 24 cm and 24 x 30 cm. | S-F
| 4(ii) | Systems using screen-film image receptors shall be equipped with moving grids matched to all image receptor sizes provided. | S-F
| 4(iii) | Systems used for magnification procedures shall be capable of operation with the grid removed from between the source and image receptor. | S-F & FFDM
Beam limitation and light fields | 5(i) | All systems shall have beam limiting devices that allow the useful beam to extend to or beyond the chest wall edge of the image receptor. | S-F & FFDM
| 5(ii) | For any mammography system with a light beam that passes through the X-ray beam-limiting device, the light shall provide an average illumination of not less than 160 lux (15 ft-candles) at 100 cm or the maximum source-image receptor distance (SID), whichever is less. | S-F & FFDM (except Fischer)
Magnification | 6(i) | Systems used to perform noninterventional problem-solving procedures shall have radiographic magnification capability available for use by the operator. | S-F & FFDM
| 6(ii) | Systems used for magnification procedures shall provide, at a minimum, at least one magnification value within the range of 1.4 to 2.0. | S-F & FFDM
Focal spot selection | 7(i) | When more than one focal spot is provided, the system shall indicate, prior to exposure, which focal spot is selected. | S-F & FFDM
| 7(ii) | When more than one target material is provided, the system shall indicate, prior to exposure, the preselected target material. | S-F & FFDM
| 7(iii) | When the target material and/or focal spot is selected by a system algorithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure. | S-F & FFDM
Application of compression | 8(i)(A) | Each system shall provide an initial power-driven compression activated by hands-free controls operable from both sides of the patient. | S-F & FFDM
| 8(i)(B) | Each system shall provide fine adjustment compression controls operable from both sides of the patient. | S-F & FFDM
Compression paddle | 8(ii)(A) | Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system. | S-F & FFDM
| 8(ii)(B) | The compression paddle shall be flat and parallel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when compression is applied. | S-F & FFDM (except Fischer)
| 8(ii)(C) | Equipment intended by the manufacturer’s design to not be flat and parallel to the breast support table during compression shall meet the manufacturer’s design specifications and maintenance requirements. | S-F & FFDM
| 8(ii)(D) | The chest wall edge of the compression paddle shall be straight and parallel to the edge of the image receptor. | S-F & FFDM
| 8(ii)(E) | The chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image. | S-F & FFDM
Technique factor selection and display | 9(i) | Manual selection of mAs or at least one of its component parts (mA and/or time) shall be available. | S-F & FFDM
| 9(ii) | The technique factors (kVp and either mA and seconds or mAs) to be used during an exposure shall be indicated before the exposure begins, except when AEC is used, in which case the technique factors that are set prior to the exposure shall be indicated. | S-F & FFDM
| 9(iii) | Following AEC mode use, the system shall indicate the actual kVp and mAs (or mA and time) used during the exposure. | S-F & FFDM
Automatic exposure control | 10(i) | Each screen-film system shall provide an AEC mode that is operable in all combinations of equipment configuration provided, e.g., grid, non-grid; magnification, nonmagnification; and various target-filter combinations. | S-F
| 10(ii) | The positioning or selection of the detector shall permit flexibility in the placement of the detector under the target tissue. The size and the available positions of the detector shall be clearly indicated at the X-ray input surface of the breast compression paddle. The selected position of the detector shall be clearly indicated. | S-F
| 10(iii) | The system shall provide means for the operator to vary the selected optical density from the normal (zero) setting. | S-F
X-ray film* | 11 | The facility shall use X-ray film for mammography that has been designated by the film manufacturer as appropriate for mammography. | S-F
Intensifying screens* | 12 | The facility shall use intensifying screens for mammography that have been designated by the screen manufacturer as appropriate for mammography and shall use film that is matched to the screen’s spectral output as specified by the manufacturer. | S-F
Film processing solutions* | 13 | For processing mammography films, the facility shall use chemical solutions that are capable of developing the films used by the facility in a manner equivalent to the minimum requirements specified by the film manufacturer. | S-F
Lighting* | 14 | The facility shall make special lights for film illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the interpreting physicians. | S-F & FFDM (for hardcopy comparison)
Film masking devices* | 15 | Facilities shall ensure that film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are available to all interpreting physicians interpreting for the facility. | S-F & FFDM (for hardcopy comparison)
* NA is acceptable for new units at existing facilities if these were previously evaluated and have not changed
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA

## Sheet1
1 | [12 empty] | Revision 3.2-20260417 | [11 empty] | Print Area
2 | [6 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [6 empty] | Page1,Page2,Artifact_Coll,AEC_ImgQual,MGD,Gen_HVL,CEDM
3 | [6 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [24 empty] | All: | Page1,Page2,Artifact_Coll,AEC_ImgQual,MGD,SBB,Gen_HVL,CEDM
5 | [6 empty] | Mammography System Compliance Inspection | [10 empty] | Mammography System Compliance Inspection
6 | [24 empty] | Measurement Parameter | Last Year |  | This Year
7 | [12 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah |  | Print Area |  | Change | Page1,Page2,Artifact_Coll,AEC_ImgQual,MGD,Gen_HVL,CEDM
8 | [6 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [6 empty] | Date
9 |  |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  | Inspector |  | Change | Eugene Mah
10 |  |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: | [5 empty] | Facility:
11 |  |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [5 empty] | Department:
12 |  |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division:
13 |  |  |  | Survey ID: |  |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID:
14 | [14 empty] | Room Number: |  |  |  | Accession Number: | [5 empty] | Room Number:
15 |  |  | X-Ray Generator | [21 empty] | Site Number:
16 |  |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [10 empty] | Authorized Use:
17 |  |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: | [5 empty] | Date of Installation:
18 |  |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number: | [5 empty] | Number of X-ray Tubes:
19 | [14 empty] | Max kVp: |  |  |  | Max mA: | [5 empty] | Accesssion Number:
20 |  |  | X-Ray Tube 1 | [21 empty] | X-Ray Generator
21 |  |  |  | Tube Designation/Use: |  |  |  |  | Target: | W |  |  |  | X-Ray Tube 1 |  |  |  |  | Target: | W | W |  |  |  | Manufacturer:
22 |  |  |  | Manufacture Date: | [10 empty] | Tube Designation/Use: | [9 empty] | Model:
23 |  |  | Insert | [5 empty] | Filter: | Rh |  |  |  |  | Manufacture Date: | [9 empty] | Max kVp:
24 |  |  |  | Manufacturer: | [5 empty] | Ag |  |  |  | Insert |  |  |  |  | Filter: | Rh | Rh |  |  |  | Manufacture Date:
25 |  |  |  | Model: | [5 empty] | Al |  |  |  |  | Manufacturer: |  |  |  |  | Ag | Ag |  |  |  | Serial Number:
26 |  |  |  | Serial Number: | [5 empty] | Cu |  |  |  |  | Model: |  |  |  |  | Al | Al |  |  |  | Max mA:
27 |  |  | Housing |  |  |  |  | Focal Spot Sizes (mm) | Large: | [5 empty] | Serial Number: |  |  |  |  | Cu | Cu |  |  |  | X-Ray Tube 1
28 |  |  |  | Manufacturer: |  |  |  |  | Small: |  |  |  |  | Housing |  |  |  |  | Focal Spot Sizes (mm) | [5 empty] | Tube Designation/Use:
29 |  |  |  | Model: |  |  |  | Filtration | Inherent: | [5 empty] | Manufacturer: |  |  |  | Large: | [5 empty] | Manufacture Date:
30 |  |  |  | Serial Number: |  |  |  |  | Added: | [5 empty] | Model: |  |  |  | Small: | [5 empty] | Manufacturer:
31 | [14 empty] | Serial Number: | [9 empty] | Model:
32 | [12 empty] | Enter 1 for YES, 2 for NO | [5 empty] | Filtration | [5 empty] | Serial Number:
33 | [6 empty] | Technique Chart | [6 empty] | Is this a breast tomosynthesis unit |  |  |  |  | Inherent: | [5 empty] | Manufacturer:
34 | [13 empty] | Is this unit used for diagnostic/magnification studies? |  |  |  |  | Added: | [5 empty] | Model:
35 |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | Is this unit used for contrast enhanced studies? | [10 empty] | Serial Number:
36 |  | Breast | [11 empty] | Stereotactic breast biopsy unit or add-on present? | [10 empty] | Large:
37 |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  |  | Is this a new installation? | [10 empty] | Small:
38 |  | <3 cm | [15 empty] | Inspection Results | [6 empty] | Target 1: |  | Change | W
39 |  | 3 to 5 cm | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [11 empty] | Target 2:
40 |  | 5 to 7 cm | [11 empty] | DHEC Registration sticker is present, clearly visible and legible | [10 empty] | Filter 1: |  | Change | Rh
41 |  | >7 cm | [11 empty] | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [10 empty] | Filter 2: |  | Change | Ag
42 | [13 empty] | Radiation warning label posted on the generator control panel | [10 empty] | Filter 3: |  | Change | Al
43 | [6 empty] | Inspection Results | [6 empty] | Operator manuals are available. | [10 empty] | Filter 4: |  | Change | Cu
44 |  | Rule Number | [8 empty] | Compliance |  |  | Monthly radiation monitoring reports are posted. | [10 empty] | Inherent:
45 |  | DES RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD |  |  | Unit installed as shown on shielding plan | [10 empty] | Added
46 |  | DES RHB 11.2.3 |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | TBD | [6 empty] | Mammography Unit Assembly Evaluation | [6 empty] | Technique Chart
47 |  | DES RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD |  |  | Free-standing unit is mechanically stable | [10 empty] | Fatty Breast
48 |  |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | All moving parts move smoothly, without obstructions to motion | [10 empty] | Target/Filter:
49 |  |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | All locks and detents work properly | [10 empty] | kVp:
50 |  |  |  | Unit installed as shown on shielding plan | [6 empty] | TBD |  |  | Image receptor holder assembly is free from vibrations | [10 empty] | Density:
51 | [6 empty] | Mammography Unit Assembly Evaluation | [5 empty] | 3 | Image receptor slides smoothly into holder assembly | [10 empty] | Target/Filter:
52 |  |  |  | Free-standing unit is mechanically stable | [6 empty] | TBD |  |  | Image receptor is held securely by assembly in any orientation | [10 empty] | kVp:
53 |  |  |  | All moving parts move smoothly, without obstructions to motion | [6 empty] | TBD |  |  | Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm | [10 empty] | Density:
54 |  |  |  | All locks and detents work properly | [6 empty] | TBD |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [10 empty] | Target/Filter:
55 |  |  |  | Image receptor holder assembly is free from vibrations | [6 empty] | TBD |  |  | Operator technique control charts are posted | [10 empty] | kVp:
56 |  |  |  | Image receptor slides smoothly into holder assembly | [6 empty] | NA |  |  | Operator protected during exposure by adequate radiation shielding | [10 empty] | Density:
57 |  |  |  | Image receptor is held securely by assembly in any orientation | [6 empty] | TBD |  |  | All indicator lights working properly | [10 empty] | Target/Filter:
58 |  |  |  | Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm | [6 empty] | TBD |  |  | Auto decompression can be overridden to maintain compression (and status displayed) | [10 empty] | kVp:
59 |  |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [6 empty] | TBD |  |  | Manual emergency compression release can be activated in the event of a power failure | [10 empty] | Density:
60 |  |  |  | Operator technique control charts are posted | [6 empty] | TBD | [13 empty] | 50/50 Breast
61 |  |  |  | Operator protected during exposure by adequate radiation shielding | [6 empty] | TBD | [6 empty] | MQSA Equipment Standards | [6 empty] | Target/Filter:
62 |  |  |  | All indicator lights working properly | [6 empty] | TBD |  |  | Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position | [10 empty] | kVp:
63 |  |  |  | Autodecompression can be overridden to maintain compression | [6 empty] | TBD |  |  | Tube-receptor assembly locking mechanism does not fail during power interruption | [10 empty] | Density:
64 |  |  |  | Manual emergency compression release in the event of a power failure | [6 empty] | TBD |  |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [10 empty] | Target/Filter:
65 | [13 empty] | Moving grids are present for all image receptor sizes | [10 empty] | kVp:
66 | [13 empty] | Magnification can be done without grid | [10 empty] | Density:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | Beam limiting device allows useful beam to extend beyond chest wall edge of receptor | [10 empty] | Target/Filter:
68 |  | Room Number | [8 empty] | Survey ID: |  |  | Average illumination of the light field is at least 160 lux | [10 empty] | kVp:
1 | [11 empty] | Medical University of South Carolina |  | Magnification mode is available for non-interventional problem solving procedures | [10 empty] | Density:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | System provides magnification factor between 1.4 – 2.0 | [10 empty] | Target/Filter:
3 |  | Rule Number |  |  |  |  | MQSA Equipment Standards | [6 empty] | Prior indication of selected focal spot is provided | [10 empty] | kVp:
4 |  | 900.12(b)(3i) |  | Tube-receptor can be fixed into any operating position and does not move | [6 empty] | TBD |  |  | Prior indication of selected target is provided | [10 empty] | Density:
5 |  | 900.12(b)(3ii) |  | Tube-receptor locking mechanism does not fail during power interruption | [6 empty] | TBD |  |  | Post exposure display of target/focal spot is provided | [10 empty] | Dense Breast
6 |  | 900.12(b)(4i) |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [6 empty] | TBD |  |  | Initial power driven compression can be activated hands-free from both sides | [10 empty] | Target/Filter:
7 |  | 900.12(b)(4ii) |  | Moving grids are present for all image receptor sizessizes | [6 empty] | TBD |  |  | Fine adjustment of compression can be activated from both sides | [10 empty] | kVp:
8 |  | 900.12(b)(4iii) |  | Magnification can be done without grid | [6 empty] | TBD |  |  | Compression paddles match the sizes of all image receptors | [10 empty] | Density:
9 |  | 900.12(b)(5i) |  | Beam limiting device allows useful beam to extend beyond chest wall edge | [6 empty] | TBD |  |  | Compression paddles do not flex by more than 1 cm | [10 empty] | Target/Filter:
10 |  | 900.12(b)(5ii) |  | Average illumination of the light field is at least 160 lux | [6 empty] | TBD |  |  | Compression paddles are flat and parallel to the breast support table | [10 empty] | kVp:
11 |  | 900.12(b)(6i) |  | Magnification mode available for non-interventional problem solving | [6 empty] | TBD |  |  | Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor | [10 empty] | Density:
12 |  | 900.12(b)(6ii) |  | System provides magnification factor between 1.4 – 2.0 | [6 empty] | TBD |  | 1 | Chest wall edge may be bent upwards for patient comfort and does not appear in the image | [10 empty] | Target/Filter:
13 |  | 900.12(b)(7i) |  | Prior indication of selected focal spot is provided | [6 empty] | TBD |  | 1 | X-ray field extends beyond the chest wall edge of the receptor | [10 empty] | kVp:
14 |  | 900.12(b)(7i) |  | Prior indication of selected target is provided | [6 empty] | TBD |  |  | Manual selection of mAs or mA and/or time is available | [10 empty] | Density:
15 |  | 900.12(b)(7iii) |  | Post exposure display of target/focal spot is provided | [6 empty] | TBD |  |  | Prior indication of selected technique factors is provided | [10 empty] | Target/Filter:
16 |  | 900.12(b)(8)(iB) |  | Initial power driven compression can be activated hands-free from both sides | [6 empty] | TBD |  |  | Post exposure display of technique factors is provided for AEC exposures | [10 empty] | kVp:
17 |  |  |  | Fine adjustment of compression can be activated from both sides | [6 empty] | TBD |  |  | AEC mode is operable in all combinations of equipment configuration | [10 empty] | Density:
18 |  | 900.12(b)(8)(iiA) |  | Compression paddles match the sizes of all image receptors | [6 empty] | TBD |  |  | AEC sensor positioning is flexible
19 |  | 900.12(b)(8)(iiB) |  | Compression paddles do not flex by more than 1 cm | [6 empty] | TBD |  |  | AEC sensor position and size is indicated on the compression paddle | [10 empty] | 2D MGD:
20 |  | 900.12(b)(8)(iiC) |  | Compression paddles are flat and parallel to the breast support table | [6 empty] | TBD |  |  | Selected position of the AEC sensor is clearly indicated | [10 empty] | mGy/mAs:
21 |  | 900.12(b)(8)(iiD) |  | Chest wall edge of the compression paddle is straight and parallel to receptor | [6 empty] | TBD |  |  | AEC density control setting is provided | [10 empty] | mGy/s:
22 |  | 900.12(b)(8)(iiE) |  | Chest wall edge may be bent upwards and does not appear in the image | [6 empty] | YES |  | 3 | X-ray film used is appropriate for mammography | [10 empty] | 3D MGD: |  | #N/A | #N/A
23 |  |  |  | X-ray field extends beyond the chest wall edge of the receptor | [6 empty] | YES |  | 3 | Intensifying screens used are appropriate for mammography and matched to the film used | [10 empty] | Combo 2D MGD:
24 |  | 900.12(b)(9i) |  | Manual selection of mAs or mA and/or time is available | [6 empty] | TBD |  | 3 | Film processing chemicals are appropriate for the film being used | [10 empty] | Combo 3D MGD: |  | #N/A | #N/A
25 |  | 900.12(b)(9ii) |  | Prior indication of selected technique factors is provided | [6 empty] | TBD |  |  | Lights for film illumination i.e. hot lights are available | [10 empty] | Combo 2D+3D MGD: |  | #N/A | #N/A
26 |  | 900.12(b)(9iii) |  | Post exposure display of technique factors is provided for AEC exposures | [6 empty] | TBD |  |  | Film masking devices are available to all interpreting physicians | [10 empty] | CEDM Low MGD:
27 |  | 900.12(b)(10i) |  | AEC mode is operable in all combinations of equipment configuration | [6 empty] | TBD | [13 empty] | CEDM High MGD: |  | #N/A | #N/A
28 |  | 900.12(b)(10ii) |  | AEC sensor positioning is flexible | [6 empty] | TBD | [6 empty] | Mammography Phototimer Technique Chart | [6 empty] | CEDM Low+High MGD: |  | #N/A | #N/A
29 |  |  |  | AEC sensor position and size is indicated on the compression paddle | [6 empty] | TBD | [13 empty] | SBB Stereo MGD:
30 |  |  |  | Selected position of the AEC sensor is clearly indicated | [6 empty] | TBD |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | SBB Tomo MGD:
31 |  | 900.12(b)(10iii) |  | AEC density control setting is provided | [6 empty] | TBD |  | Breast | [11 empty] | System Resolution
32 |  | 900.12(b)(11) |  | X-ray film used is appropriate for mammography | [6 empty] | NA |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  |  | W
33 |  | 900.12(b)(12) |  | Intensifying screens used are appropriate for mammography and matched to film | [6 empty] | NA |  | <3 cm | [11 empty] | 0
34 |  | 900.12(b)(13) |  | Film processing chemicals are appropriate for the film being used | [6 empty] | NA |  | 3 to 5 cm | [11 empty] | W
35 |  | 900.12(b)(14) |  | Lights for film illumination i.e. hot lights are available | [6 empty] | TBD |  | 5 to 7 cm
36 |  | 900.12(b)(15) |  | Film masking devices are available to all interpreting physicians | [6 empty] | TBD |  | >7 cm | [11 empty] | Image Quality (Printer)
37 | [6 empty] | Comments | [17 empty] | kV: |  | Change | 28
38 | [12 empty] | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | mA: |  |  | 0
39 | [12 empty] | Breast | [11 empty] | Bkg Density:
40 | [12 empty] | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  |  | Inside disk:
41 | [12 empty] | <3 cm | [11 empty] | Outside disk:
42 | [12 empty] | 3 to 5 cm | [11 empty] | Difference:
43 | [12 empty] | 5 to 7 cm | [11 empty] | Fibers:
44 | [12 empty] | >7 cm | [11 empty] | Specks:
45 | [24 empty] | Masses:
46 | [17 empty] | Inspection Results | [6 empty] | Image Quality (Acq Unit)
47 | [12 empty] | Compression Force Indicator | [11 empty] | 2D kV: |  | Change | 28
48 | [14 empty] | Indicated | Measured | After 20s | Slip | [6 empty] | mA:
49 | [13 empty] | Power | [10 empty] | Exp Ind:
50 | [13 empty] | Manual | [10 empty] | Fibers:
51 | [13 empty] | Criteria: | Minimum compression force – 25 lbs | [9 empty] | Specks:
52 | [14 empty] | Maximum compression force – 45 lbs | [9 empty] | Masses:
53 | [24 empty] | Mag kV:
54 | [12 empty] | Compression Thickness Indicator | [11 empty] | mA:
55 | [13 empty] | BR-12 (cm) | Indicated | Difference | [8 empty] | Exp Ind:
56 | [13 empty] | 1 | [10 empty] | Fibers:
57 | [13 empty] | 2 | [10 empty] | Specks:
58 | [13 empty] | 4 |  |  | Mean: | [7 empty] | Masses:
59 | [13 empty] | 4 |  |  | StDev: | [7 empty] | 3D kV: |  | Change | 29
60 | [13 empty] | 4 | [10 empty] | mA:
61 | [13 empty] | 4 |  |  |  | Criteria: | Difference between indicated and actual thickness is ± 5mm | [5 empty] | Exp Ind:
62 | [13 empty] | 6 | [10 empty] | Fibers:
63 | [13 empty] | 8 | [10 empty] | Specks:
64 | [24 empty] | Masses:
65 | [13 empty] | Paddle flex is less than 1 cm |  |  |  | Compression thickness indicator: | [6 empty] | SBB kV:
66 | [13 empty] | Overall unit assembly | [10 empty] | mA:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah | [12 empty] | Exp Ind:
68 |  | Room Number | [8 empty] | Survey ID: |  | Artifact Evaluation |  | Acceptable: Enter 1 for YES, 2 for NO, 3 for NA | [9 empty] | Fibers:
1 | [11 empty] | Medical University of South Carolina |  | Target: | W | W | W | W | W | W |  |  |  |  | Specks:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Filter: | Rh | Ag | Al | Cu | Rh | Ag |  |  |  |  | Masses:
3 |  | Compression Force Indicator | [11 empty] | Focal spot: | L | L | L | L | S | S |  |  |  |  | SBB Tomo kV:
4 |  |  |  | Indicated | Measured | After 20s | Slip | [6 empty] | kV: | 28 | 31 | 30 | 49 | 28 | 31 |  |  |  |  | mA:
5 |  |  | Power | [10 empty] | mAs: | [10 empty] | Exp Ind:
6 |  |  | Manual | [5 empty] | Acceptable: |  |  |  |  | EI: | [10 empty] | Fibers:
7 |  |  | Criteria: | Minimum compression force – 25 lbs | [9 empty] | Acceptable: | [10 empty] | Specks:
8 |  |  |  | Maximum compression force – 45 lbs | [10 empty] | *W/Cu for CEDM only. Use manual mode. | [9 empty] | Masses:
9 | [12 empty] | Detector Uniformity
10 |  | Compression Thickness Indicator | [11 empty] | Target/Filter: | W/Rh | W/Ag | W/Al | W/Cu | W/Rh | W/Ag |  |  |  |  | SNR: |  | #DIV/0! | #DIV/0! |  |  | Detector Uniformity deviation
11 |  |  | BR-12 (cm) | Indicated | Difference |  |  |  | Paddle flex is less than 1 cm: |  |  |  |  | kV: | 28 | 31 | 30 | 49 | 28 | 31 |  |  |  |  | CNR: |  | #DIV/0! | #DIV/0! |  |  | W/Rh | W/Ag | W/Al | W/Cu | W/Rh | W/Ag
12 |  |  | 1 | [5 empty] | Overall unit assembly: |  |  |  |  | mAs: | [15 empty] | ROI 1:
13 |  |  | 2 | [5 empty] | Compression thickness indicator: |  |  |  |  | ROI 1: | [10 empty] | Comment Page 1 |  |  |  |  | ROI 2:
14 |  |  | 4 |  |  | Mean: | [7 empty] | ROI 2: | [15 empty] | ROI 3:
15 |  |  | 4 |  |  | StDev: | [7 empty] | ROI 3: | [15 empty] | ROI 4:
16 |  |  | 4 |  |  | Criteria: | Difference between indicated and actual thickness is ± 5mm | [6 empty] | ROI 4: | [15 empty] | ROI 5:
17 |  |  | 4 | [10 empty] | ROI 5:
18 |  |  | 6 | [10 empty] | Average:
19 |  |  | 8 | [10 empty] | Max deviation:
20 | [13 empty] | Acceptable:
21 |  | Artifact Evaluation
22 |  |  | Receptor: | W | W | W | W | W | W |  |  |  |  | Criteria: | Maximum deviation from average is <= 7%
23 |  |  | Target: | W | W | W | W | W | W
24 |  |  | Filter: | Rh | Ag | Al | Cu | Rh | Ag
25 |  |  | Focal spot: | L | L | L | L | S | S
26 |  |  | kV: | 28 | 31 | 30 | 49 | 28 | 31
27 |  |  | mAs: | [10 empty] | Comments:
28 |  |  | EI: | [10 empty] | Do not exceed 390 chars
29 |  |  | Artifact Evaluation Acceptable: | [10 empty] | New:
30 |  |  | Detector Uniformity Deviation:
31 |  |  | Detector Uniformity Acceptable: | [9 empty] | Light Field Brightness |  |  | 1 | Enter 1 for YES, 2 for NO, 3 for NA
32 |  |  | Criteria: | No visible artifacts in the images | [9 empty] | Measurements (lux) |  |  |  |  | Average Illumination
33 |  |  |  | Maximum detector uniformity deviation is <= 7% | [9 empty] | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux)
34 | [21 empty] | Acceptable:
35 |  | Light Field Brightness
36 |  |  | Measurements (lux) |  |  |  |  | Average Illumination |  |  |  |  | Light Field/X-ray Field Alignment
37 |  |  | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux) |  |  |  |  | SID (cm): | 65
38 | [10 empty] | Acceptable: |  |  | 24x30 | 18x24 |  |  | 24x29 | Small breast
39 |  |  | Criteria: | Average light field illumination >= 160 lux. | [10 empty] | Left | Center | Right | Tomo
40 | [12 empty] | Left:
41 |  | Light Field/X-ray Field Alignment |  |  |  | X-ray Field/Image Receptor Alignment | [6 empty] | Right:
42 |  | SID (cm): | [10 empty] | Anterior:
43 |  | 65 | Light Field/X-ray Field % Deviation | [9 empty] | Chest Wall:
44 |  |  | 24x30 | 18x24 |  |  | 24x29 | Small breast | [6 empty] | Enter deviation in cm
45 |  |  | 24x30 | Left | Center | Right | Tomo | Tomo
46 |  | Left: | [11 empty] | % Deviation
47 |  | Right: | [11 empty] | 24x30 | 18x24 |  |  | 24x29 | Small breast
48 |  | Anterior: | [12 empty] | Left | Center | Right | Tomo
49 |  | Chest Wall: | [10 empty] | Left:
50 |  | Left-Right: | [10 empty] | Right:
51 |  | Ant-Chest: | [10 empty] | Anterior:
52 |  | Acceptable: | [10 empty] | Chest Wall:
53 |  |  | X-ray Field/Image Receptor % Deviation | [9 empty] | Left-Right
54 |  | Left: | [10 empty] | Ant-Chest
55 |  | Right: | [11 empty] | Criteria: | Sum of left and right edge deviation is less than 2% SID
56 |  | Anterior: | [12 empty] | Sum of anterior and chest wall edge deviation is less than 2% SID
57 |  | Chest Wall: | [12 empty] | Enter deviation in cm. Positive deviation indicates x-ray field extends beyond light field
58 |  | Paddle: | [10 empty] | X-ray Field/Image Receptor Alignment
59 |  | Acceptable:
60 |  | Criteria: | Sum of left/right and chest wall/anterior deviation is less than 2% SID | [10 empty] | 24x30 | 18x24 |  |  | 24x29 | Small breast
61 |  |  | X-ray field does not extend beyond image receptor by more than 2% SID | [11 empty] | Left | Center | Right | Tomo
62 |  |  | X-ray field extends beyond the chest wall edge of image receptor. | [9 empty] | Left:
63 |  |  | X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side | [9 empty] | Right:
64 |  |  | X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side | [9 empty] | Anterior:
65 |  |  | Difference between chest wall edge of the receptor and compression paddle is less than 1% SID | [9 empty] | Chest Wall:
66 | [12 empty] | Paddle:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  |  | Enter deviation in cm. Positive deviation indicates x-ray field extends beyond image receptor
68 |  | Room Number | [8 empty] | Survey ID:
1 | [11 empty] | Medical University of South Carolina |  | % Deviation
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | 24x30 | 18x24 |  |  | 24x29 | Small breast
3 |  | AEC Thickness Tracking – 2D | [5 empty] | CNR Table: | 4 | [5 empty] | Left | Center | Right | Tomo
4 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | Mag CNR Table: | 4 |  |  |  | Left:
5 |  |  | Thickness |  |  |  |  | Reported |  |  |  |  | Right:
6 |  |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  | Anterior:
7 |  |  | 2 | Rh | 25 | [7 empty] | Chest Wall:
8 |  |  | 4 | Rh | 28 | [7 empty] | Paddle:
9 |  |  | 4 | Rh | 28 | [7 empty] | 1 | X-ray field extends beyond the chest wall edge of the receptor
10 |  |  | 4 | Rh | 28 | [7 empty] | 1 | Chest wall edge of the compression paddle is not visible in the image
11 |  |  | 4 | Rh | 28
12 |  |  | 6 | Rh | 31 | [8 empty] | Criteria: | X-ray field does not extend beyond image receptor by more than 2% SID
13 |  |  | 8 | Ag | 32 | [9 empty] | X-ray field extends beyond the chest wall edge of image receptor.
14 |  |  | 4 | Rh | 29 | [9 empty] | X-ray field does not fall within the image receptor by more than 2% SID on any side
15 | [5 empty] | Average: | [8 empty] | Difference between chest wall edge of the receptor and compression paddle is less than 1% SID
16 |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
17
18 |  | AEC Thickness Tracking – 3D | [10 empty] | System Resolution |  | Large | Small | Large | Small
19 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0 |  |  |  |  | Mode (2D/3D): | 2D |  | 3D
20 |  |  | Thickness |  |  |  |  | Reported | [5 empty] | Target: | W |  | W
21 |  |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | kV: | 28 |  | 30
22 |  |  | 2 | Al | 26 | [8 empty] | mAs: | 50 |  | 50
23 |  |  | 4 | Al | 29 | [8 empty] | lp/mm:
24 |  |  | 4 | Al | 29 | [8 empty] | Acceptable:
25 |  |  | 4 | Al | 29 | [8 empty] | Previous:
26 |  |  | 4 | Al | 29
27 |  |  | 6 | Al | 33 | [8 empty] | Criteria: | Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode
28 |  |  | 8 | Al | 38 | [9 empty] | Limiting system resolution must be 3 lp/mm or higher in 3D mode
29 | [5 empty] | Average: | [6 empty] | AEC Thickness Tracking – 2D | [5 empty] | CNR Table: | 4 | Enter -1 if not applicable
30 |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [8 empty] | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | Mag CNR Table: | 4 | Enter -1 if not applicable
31 | [13 empty] | Thickness | [5 empty] | Reported
32 |  | AEC Density Control | [11 empty] | (cm) | Filter | kV | mAs | EI | EI* | MGD (mGy) | Deviation
33 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | kV: | 28 |  |  |  |  | 2 | Rh | 25
34 | [5 empty] | Reported | [7 empty] | 4 | Rh | 28
35 |  |  | Density | mAs | EI | MGD (mGy) | Change | Acceptable | [5 empty] | 4 | Rh | 28
36 |  |  | 4 | [10 empty] | 4 | Rh | 28
37 |  |  | 3 | [10 empty] | 4 | Rh | 28
38 |  |  | 2 | [10 empty] | 6 | Rh | 31
39 |  |  | 1 | [10 empty] | 8 | Ag | 32
40 |  |  | 0 | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  |  | Mag: | 4 | Rh | 29
41 |  |  | -1 | [11 empty] | EI* = EI - DC offset (50) |  |  | Average:
42 |  |  | -2 | [10 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
43 |  |  | -3
44 |  |  | Criteria: | Change in pixel value at each density step is within specified ranges | [8 empty] | AEC Thickness Tracking – 3D
45 | [12 empty] | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0
46 |  | System Resolution |  | Large | Small | Large | Small | [6 empty] | Thickness | [5 empty] | Reported
47 |  |  | Mode (2D/3D): | 2D |  | 3D | [7 empty] | (cm) | Filter | kV | mAs | EI | EI* | MGD (mGy) | Deviation
48 |  |  | Target: | W |  | W | [7 empty] | 2 | Al | 26
49 |  |  | kV: | 28 |  | 30 |  |  | Acceptable: |  |  |  |  | 4 | Al | 29
50 |  |  | mAs: | 50 |  | 50 | [7 empty] | 4 | Al | 29
51 |  |  | lp/mm: | [10 empty] | 4 | Al | 29
52 |  |  | Criteria: | Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode | [9 empty] | 4 | Al | 29
53 | [13 empty] | 6 | Al | 33
54 |  | Image Quality Evaluation | [11 empty] | 8 | Al | 38
55 |  |  | Printer |  |  | Acq Workstation |  |  |  | Review Workstation |  |  |  |  | EI* = EI - DC offset (50) |  |  | Average:
56 |  |  | kV: | 28 |  | Normal | Mag | Tomo |  | Left | Right |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
57 |  |  | mA: | 0 | kV: | 28 |  | 29 |  | Workstation:
58 |  |  | Background Density: |  | mA: |  |  |  | Fibers:
59 |  |  | Inside disk: |  | EI: |  |  |  | Specks: |  |  |  | AEC Density Control | [10 empty] | Selenia |  | Selenia Dimensions
60 |  |  | Outside disk: |  | Fibers: |  |  |  | Masses: |  |  |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 | Reported |  |  |  | Selenia |  | Low | High | Low | High
61 |  |  | Difference: |  | Specks: |  |  |  |  | Workstation: |  |  |  | Density | kV | mAs | EI | MGD (mGy) | EI* | Change |  | Low | High | 1.44 | 1.76 | 1.61 | 1.89
62 |  |  | Fibers: |  | Masses: |  |  |  | Fibers: |  |  |  |  | 4 | 28 | [6 empty] | 1.44 | 1.76 | 1.31 | 1.6 | 1.4 | 1.64
63 |  |  | Specks: |  | Criteria: |  |  |  | Specks: |  |  |  |  | 3 | [7 empty] | 1.31 | 1.6 | 1.17 | 1.43 | 1.22 | 1.43
64 |  |  | Masses: |  | 2D – Must be able to see 5 fibers, 4 speck groups, 4 masses |  |  |  | Masses: |  |  |  |  | 2 | [7 empty] | 1.17 | 1.43 | 1.04 | 1.27 | 1.06 | 1.24
65 |  |  |  |  | 3D – Must be able to see 4 fibers, 3 speck groups, 3 masses | [8 empty] | 1 | [7 empty] | 0 | 0 | [5 empty] | 1.44 | 1.76 | 1.61 | 1.89
66 | [13 empty] | 0 | [7 empty] | 0.77 | 0.94 | 0.77 | 0.94 | 0.78 | 0.92 |  | 1.31 | 1.6 | 1.4 | 1.64
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | -1 | [7 empty] | 0.63 | 0.77 | 0.63 | 0.77 | 0.66 | 0.78 |  | 1.17 | 1.43 | 1.22 | 1.43
68 |  | Room Number | [8 empty] | Survey ID: |  |  | -2 | [7 empty] | 0.5 | 0.61 | 0.5 | 0.61 | 0.56 | 0.66 |  | 1.04 | 1.27 | 1.06 | 1.24
1 | [11 empty] | Medical University of South Carolina |  | -3
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  |  | EI* = EI - DC offset (50) | [13 empty] | 0.77 | 0.94 | 0.78 | 0.92
3 | Stereotactic Breast Biopsy | [12 empty] | Criteria: | Change in pixel value at each density step is within specified ranges | [13 empty] | 0.63 | 0.77 | 0.66 | 0.78
4 | [28 empty] | 0.5 | 0.61 | 0.56 | 0.66
5 |  | AEC Thickness Tracking - Stereotactic Breast Biopsy Add-on
6 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0 |  |  |  | AEC Thickness Tracking - Stereotactic Breast Biopsy Add-on
7 |  |  | Thickness |  |  |  |  | Reported |  |  |  |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0
8 |  |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | Thickness | [5 empty] | Reported
9 |  |  | 2 | [10 empty] | (cm) | Filter | kV | mAs | EI | EI* | MGD (mGy) | Deviation
10 |  |  | 4 | [10 empty] | 2
11 |  |  | 4 | 0 | 0 | [8 empty] | 4
12 |  |  | 4 | 0 | 0 | [8 empty] | 4 | 0 | 0
13 |  |  | 4 | 0 | 0 | [8 empty] | 4 | 0 | 0
14 |  |  | 6 | [10 empty] | 4 | 0 | 0
15 |  |  | 8 | [10 empty] | 6
16 | [5 empty] | Average: | [7 empty] | 8
17 |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [10 empty] | EI* = EI - DC offset (50) |  |  | Average:
18 | [13 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
19 |  | Mean Glandular Dose – Stereo biopsy | [10 empty] | Mean Glandular Dose – 2D
20 |  |  | Phantom: | 0 |  | kV: | [7 empty] | Phantom: |  |  | kV: | 28
21 |  |  | Target/Filter: |  |  | mA: | [7 empty] | Target/Filter: | W/Rh |  | mA: | 100
22 | [5 empty] | Reported |  | Air KERMA | [8 empty] | Reported |  | Air KERMA | [5 empty] | Bkg Mean | SD | Disk Mean | SD | SNR | CNR
23 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #DIV/0! | [6 empty] | #DIV/0! | #DIV/0!
24 | [9 empty] | ESE (mGy): | [10 empty] | ESE (mGy): | [7 empty] | #DIV/0! | #DIV/0!
25 | [9 empty] | DGN (mrad/R): | [10 empty] | DGN (mrad/R): | TBD | [6 empty] | #DIV/0! | #DIV/0!
26 | [9 empty] | MGD (mGy): | [10 empty] | MGD (mGy): | [7 empty] | #DIV/0! | #DIV/0!
27 | [9 empty] | Prev MGD (mGy): | [10 empty] | Prev MGD (mGy):
28 |  |  | Average: | [6 empty] | % Change: |  |  |  | Average: | [6 empty] | % Change:
29 |  |  | Coeff of Var: | [10 empty] | Coeff of Var:
30 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
31 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
32 | [14 empty] | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
33 |  | Mean Glandular Dose – Tomo biopsy
34 |  |  | Phantom: | 0 |  | kV: | [6 empty] | Mean Glandular Dose – 3D
35 |  |  | Target/Filter: |  |  | mA: | [7 empty] | Phantom: | 0 |  | kV: | 29
36 | [5 empty] | Reported |  | Air KERMA | [5 empty] | Target/Filter: | W/Al |  | mA: | 160
37 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | [6 empty] | Reported |  | Air KERMA
38 | [9 empty] | ESE (mGy): |  |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | 0
39 | [9 empty] | DGN (mrad/R): | [10 empty] | ESE (mGy):
40 | [9 empty] | MGD (mGy): | [10 empty] | DGN (mrad/R): | #N/A
41 | [9 empty] | Prev MGD (mGy): | [10 empty] | MGD (mGy): | #N/A
42 |  |  | Average: | [6 empty] | % Change: | [10 empty] | Prev MGD (mGy):
43 |  |  | Coeff of Var: | [10 empty] | Average: | [6 empty] | % Change: | #N/A
44 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: |  |  |  | Coeff of Var:
45 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
46 | [13 empty] | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A
47 |  | Phantom Image Quality | [12 empty] | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | #N/A
48 |  |  | Acq Workstation | [11 empty] | MGD for 2D + 3D acquisitions is less than 3 mGy
49 |  |  | SBB | SBB Tomo
50 |  | kV:
51 |  | mA: | [10 empty] | Mean Glandular Dose – Combo
52 |  | EI: | [11 empty] | Phantom: | 0
53 |  | Fibers:
54 |  | Specks: | [10 empty] | Combo Mode 2D
55 |  | Masses: | [11 empty] | Target/Filter: | W/Rh |  | kV: | 28
56 |  | Criteria: | SBB Stereo - Must be able to see 5 fibers, 4 speck groups, 3.5 masses | [13 empty] | mA: | 100
57 |  |  | SBB Tomo - Must be able to see 3 fibers, 3 speck groups, 2.5 masses | [13 empty] | Reported |  | Air KERMA
58 | [14 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #DIV/0!
59 | [20 empty] | ESE (mGy):
60 | [20 empty] | DGN (mrad/R): | TBD
61 | [20 empty] | MGD (mGy):
62 | [20 empty] | Prev MGD (mGy):
63 | [13 empty] | Average: | [6 empty] | % Change:
64 | [13 empty] | Coeff of Var:
65 | [20 empty] | Deviation from Reported MGD:
66 | [12 empty] | Combo Mode 3D
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | Target/Filter: | W/Al |  | kV: | 29
68 |  | Room Number | [8 empty] | Survey ID: | [5 empty] | mA: | 160
1 | [11 empty] | Medical University of South Carolina |  |  |  |  | Reported |  | Air KERMA
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | 0
3 |  | Mean Glandular Dose – 2D | [18 empty] | ESE (mGy):
4 |  |  | Phantom: |  |  | kV: | 28 | [13 empty] | DGN (mrad/R): | #N/A
5 |  |  | Target/Filter: | W/Rh |  | mA: | 100 | [13 empty] | MGD (mGy): | #N/A
6 | [5 empty] | Reported |  | Air KERMA | [12 empty] | Prev MGD (mGy):
7 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #DIV/0! |  |  | Average: | [6 empty] | % Change: | #N/A
8 | [9 empty] | ESE (mGy): |  |  |  | Coeff of Var:
9 | [9 empty] | DGN (mrad/R): | TBD | [9 empty] | Deviation from Reported MGD: | #N/A
10 | [9 empty] | MGD (mGy): | [10 empty] | Combo MGD (2D + 3D): | #N/A
11 | [9 empty] | Prev MGD (mGy): | [10 empty] | Prev Combo MGD:
12 |  |  | Average: | [6 empty] | % Change: |  |  | Mean Glandular Dose – CEDM | [10 empty] | Phantom: | 800004-2203184 |  | kV: | 28
13 |  |  | Coeff of Var: | [10 empty] | Phantom: | 0 | [8 empty] | Target/Filter: | W/Rh |  | mA: | 100
14 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
15 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: |  |  | CEDM - Low energy |  |  |  | kV: | 29
16 | [13 empty] | Target/Filter: | W/Rh |  | mA: | 160
17 |  | Mean Glandular Dose – 3D | [14 empty] | Reported |  | Air KERMA |  | HVL: | #DIV/0!
18 |  |  | Phantom: | 0 |  | kV: | 29 | [7 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | ESE (mGy):
19 |  |  | Target/Filter: | W/Al |  | mA: | 160 | [13 empty] | DGN (mrad/R): | TBD
20 | [5 empty] | Reported |  | Air KERMA | [12 empty] | MGD (mGy):
21 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | 0 | [9 empty] | Prev MGD (mGy):
22 | [9 empty] | ESE (mGy): | [10 empty] | % Change:
23 | [9 empty] | DGN (mrad/R): | #N/A |  |  | Average:
24 | [9 empty] | MGD (mGy): | #N/A |  |  | Coeff of Var:
25 | [9 empty] | Prev MGD (mGy): | [10 empty] | Deviation from Reported MGD:
26 |  |  | Average: | [6 empty] | % Change: | #N/A | [9 empty] | mAs at 3 mGy:
27 |  |  | Coeff of Var: | [9 empty] | CEDM - High energy
28 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A |  |  | Target/Filter: | W/Cu |  | kV: | 45
29 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | #N/A | [5 empty] | mA: | 100
30 | [16 empty] | Reported |  | Air KERMA
31 |  | Mean Glandular Dose – Combo | [12 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | 0
32 |  |  | Phantom: | 0 | [16 empty] | ESE (mGy):
33 |  |  |  | Combo Mode 2D Target/Filter: W/Rh | [16 empty] | DGN (mrad/R): | #N/A
34 | [5 empty] | Reported |  | Air KERMA |  | kV: | 28 | [9 empty] | MGD (mGy): | #N/A
35 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #DIV/0! | [9 empty] | Prev MGD (mGy):
36 | [9 empty] | ESE (mGy): |  |  |  | Average: | [6 empty] | % Change: | #N/A
37 | [9 empty] | DGN (mrad/R): | TBD |  |  | Coeff of Var:
38 | [9 empty] | MGD (mGy):
39 | [9 empty] | Prev MGD (mGy): |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A
40 |  |  | Average: | [6 empty] | % Change: |  |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | CEDM MGD (Low + High) (mGy): | #N/A
41 |  |  | Coeff of Var: | [11 empty] | MGD for Low + High acquisitions is less than 3 mGy | [5 empty] | Prev CEDM MGD:
42 | [9 empty] | Deviation from Reported MGD: |  |  | Mean Glandular Dose – Stereo biopsy
43 | [13 empty] | Phantom: | 0 |  | kV:
44 |  |  |  | Combo Mode 3D Target/Filter: W/Al | [9 empty] | Target/Filter: |  |  | mA:
45 | [5 empty] | Reported |  | Air KERMA |  | kV: | 29 | [5 empty] | Reported |  | Air KERMA
46 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | 0 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
47 | [9 empty] | ESE (mGy): | [10 empty] | ESE (mGy):
48 | [9 empty] | DGN (mrad/R): | #N/A | [9 empty] | DGN (mrad/R):
49 | [9 empty] | MGD (mGy): | #N/A | [9 empty] | MGD (mGy):
50 | [9 empty] | Prev MGD (mGy): | [10 empty] | Prev MGD (mGy):
51 |  |  | Average: | [6 empty] | % Change: | #N/A |  |  | Average: | [6 empty] | % Change:
52 |  |  | Coeff of Var: | [10 empty] | Coeff of Var:
53 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A
54 |  |  |  | Deviation from reported MGD is less than15% | [9 empty] | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
55 | [9 empty] | Combo MGD: | #N/A |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
56 | [9 empty] | Prev Combo MGD:
57 | [12 empty] | Mean Glandular Dose – Tomo biopsy
58 | [13 empty] | Phantom: | 0 |  | kV:
59 |  | Signal to Noise Ratio/Contrast to Noise Ratio | [11 empty] | Target/Filter: |  |  | mA:
60 |  |  | kV: | 28 |  |  | Current | Previous | Change | Acceptable | [6 empty] | Reported |  | Air KERMA
61 |  |  | mAs: | 0 |  | SNR: | #DIV/0! |  | #DIV/0! | #DIV/0! |  |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
62 |  |  | Target: | W |  | CNR: | #DIV/0! |  | #DIV/0! | NA | [10 empty] | ESE (mGy):
63 |  |  | Filter: | Rh | [16 empty] | DGN (mrad/R):
64 |  |  | Criteria: | SNR must be >= 40 | [16 empty] | MGD (mGy):
65 |  |  |  | CNR does not change by more than 15% | [16 empty] | Prev MGD (mGy):
66 | [13 empty] | Average: | [6 empty] | % Change:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | Coeff of Var:
68 |  | Room Number | [8 empty] | Survey ID:
1 | [11 empty] | Medical University of South Carolina |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
3 |  | Meter: |  |  |  |  | Calibration Date: | [5 empty] | kVp Accuracy and Output
4 |  | SN: |  |  |  |  | Calibration Due: | [5 empty] | Meter: |  |  |  | Calibration Date:
5 |  | kVp Accuracy | [10 empty] | Serial #: |  |  |  | Calibration Due:
6 |  | Target: | W |  |  | Target: | W |  |  | Target: | W
7 |  | Filter: | Rh |  |  | Filter: | Ag |  |  | Filter: | Al |  |  | Criteria: | kV accuracy is <= 5% |  |  | Measured Values
8 |  | mAs: | 50 |  |  | mAs: | 50 |  |  | mAs: | 50 |  |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
9 |  | Indicated | Measured |  |  | Indicated | Measured |  |  | Indicated | Measured |  |  | W | Rh | 24 | 50
10 |  | kVp | kVp | Error (%) |  | kVp | kVp | Error (%) |  | kVp | kVp | Error (%) |  | W | Rh | 25 | 50
11 |  | 24 |  |  |  | 28 |  |  |  | 28 |  |  |  | W | Rh | 26 | 50
12 |  | 25 |  |  |  | 30 |  |  |  | 30 |  |  |  | W | Rh | 28 | 50
13 |  | 26 |  |  |  | 32 |  |  |  | 32 |  |  |  | W | Rh | 30 | 50
14 |  | 28 |  |  |  | 34 |  |  |  | 34 |  |  |  | W | Rh | 32 | 50
15 |  | 30 |  |  |  | 36 |  |  |  | 36 |  |  |  | W | Rh | 34 | 50
16 |  | 32 |  |  |  | 38 |  |  |  | 38 |  |  |  | W | Rh | 36 | 50
17 |  | 34 | [11 empty] | W | Rh | 38 | 50
18 |  | 36 |  |  |  |  | Acceptable: |  |  |  | Acceptable:
19 |  | 38 | [15 empty] | Measured Values
20 |  |  | Acceptable: |  | Criteria: | kVp accurate to within 5% of indicated | [7 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
21 | [13 empty] | W | Ag | 28 | 50
22 |  | kVp and Output Reproducibility | [11 empty] | W | Ag | 30 | 50
23 |  | Target: | W | Filter: | Rh | [8 empty] | W | Ag | 32 | 50
24 |  | mAs: | 50 | [10 empty] | W | Ag | 34 | 50
25 |  | Indicated | Measured | [10 empty] | W | Ag | 36 | 50
26 |  | kVp | kVp | mGy | mGy/mAs | mGy/s | [7 empty] | W | Ag | 38 | 50
27 |  | 28
28 | [17 empty] | Measured Values
29 | [13 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
30 | [13 empty] | W | Al | 28 | 50
31 |  | Mean: | [11 empty] | W | Al | 30 | 50
32 |  | Std Dev: | [11 empty] | W | Al | 32 | 50
33 |  | Coeff of Var: | [11 empty] | W | Al | 34 | 50
34 |  | Acceptable: |  |  |  | #VALUE! | [7 empty] | W | Al | 36 | 50
35 |  |  | Criteria: | Coefficient of variation <= 2% | [9 empty] | W | Al | 38 | 50
36 |  |  |  | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
37 |  |  |  | Minimum acceptable output is 2 mGy/s at 28 kVp for W target | [13 empty] | Measured Values
38 | [13 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
39 |  | Exposure Linearity | [11 empty] | W | Cu | 40 | 50
40 |  | Target: | W | Filter: | Rh | [8 empty] | W | Cu | 42 | 50
41 |  | Set mAs | kVp | mGy | mGy/mAs | mGy/s | [7 empty] | W | Cu | 44 | 50
42 |  | 20 | [11 empty] | W | Cu | 45 | 50
43 |  | 50 | [11 empty] | W | Cu | 46 | 50
44 |  | 100 | [11 empty] | W | Cu | 48 | 50
45 |  | 300 | [11 empty] | W | Cu | 49 | 50
46 |  | 400 | [11 empty] | Criteria: | kV accuracy is <= 5%
47 |  |  |  | Coeff of Linearity: | [8 empty] | kV, Exposure, and Output Reproducibility
48 |  |  | Criteria: | Coefficient of linearity is <= 10% | [13 empty] | Measured Values
49 | [13 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
50 |  | Half Value Layer | [11 empty] | W | Rh | 28 | 50
51 |  | Target/Filter | W/Rh | W/Rh | W/Rh | W/Rh | W/Ag | W/Ag | W/Ag | W/Ag |  |  |  | W | Rh | 28 | 50
52 |  | kV | 24 | 25 | 28 | 32 | 28 | 30 | 32 | 34 |  |  |  | W | Rh | 28 | 50
53 |  | HVL | [11 empty] | W | Rh | 28 | 50
54 |  | 0 mm | [14 empty] | Mean:
55 |  | Min | 0.27 | 0.28 | 0.31 | 0.35 | 0.31 | 0.33 | 0.35 | 0.37 | [6 empty] | Std Dev:
56 |  | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.47 | 0.49 | 0.51 | 0.53 | [6 empty] | Coeff of Var:
57 |  | Acceptable: | [14 empty] | Previous:
58 | [13 empty] | Criteria: | Coefficient of variation <= 2%
59 |  | Target/Filter | W/Al | W/Al | W/Al | W/Al | W/Al | [7 empty] | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
60 |  | kV | 28 | 30 | 32 | 34 | 36 | [7 empty] | Minimum acceptable output is 2 mGy/s at 28 kVp for W target
61 |  | HVL | [10 empty] | Exposure Linearity
62 |  | Min | 0.31 | 0.33 | 0.35 | 0.37 | 0.39 | [10 empty] | Measured Values
63 |  | Max | 0.4 | 0.42 | 0.44 | 0.46 | 0.55 | [6 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
64 |  | Acceptable: | [11 empty] | W | Rh | 28 | 20
65 |  |  | Criteria: | W targets – HVL is > (kV/100)+0.03 | [9 empty] | W | Rh | 28 | 50
66 | [13 empty] | W | Rh | 28 | 100
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | W | Rh | 28 | 300
68 |  | Room Number | [8 empty] | Survey ID: |  |  | W | Rh | 28 | 400
1 | [11 empty] | Medical University of South Carolina |  | Criteria: | Coefficient of linearity is <= 10% |  |  |  |  | Coeff of Linearity:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Unit Can Maintain Output for 3 sec at 28 kVp
3 | Contrast Enhanced Mammography
4 |  | Target: | W |  |  |  | Meter: |  |  | Calibration Date: |  |  | Half Value Layer – 2D |  | 1 | Compression paddle used?
5 |  | Filter: | Cu |  |  |  | SN: |  |  | Calibration Due: |  |  |  | Target/Filter | W/Rh | W/Rh | W/Rh | W/Rh | W/Ag | W/Ag | W/Ag | W/Ag
6 |  | kVp Accuracy |  |  |  | Half Value Layer | [7 empty] | kV | 24 | 25 | 28 | 32 | 28 | 30 | 32 | 34
7 |  | Indicated | Measured | [10 empty] | Al (mm) | Air Kerma (mGy)
8 |  | kVp | kVp | Error (%) |  | kVp | HVL (mm Al) | Acceptable | [5 empty] | 0
9 |  | 40 |  |  |  | 40 | [7 empty] | HVL
10 |  | 42 |  |  |  | 42 | [7 empty] | Min | 0.27 | 0.28 | 0.31 | 0.35 | 0.31 | 0.33 | 0.35 | 0.37
11 |  | 44 |  |  |  | 44 | [7 empty] | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.47 | 0.49 | 0.51 | 0.53
12 |  | 45 |  |  |  | 45
13 |  | 46 |  |  |  | 46
14 |  | 48 |  |  |  | 48
15 |  | 49 |  |  |  | 49
16 |  |  | Acceptable: | [9 empty] | Half Value Layer – 3D
17 | [13 empty] | Target/Filter | W/Al | W/Al | W/Al | W/Al | W/Al | W/Al
18 |  | Mean Glandular Dose - Low Energy | [11 empty] | kV | 28 | 30 | 32 | 34 | 36 | 38
19 |  | Phantom: | 0 |  | kV: | 29 | [7 empty] | Al (mm) | Air Kerma (mGy)
20 |  | Target/Filter: | W/Rh |  | mA: | 160 | [7 empty] | 0
21 |  |  |  |  | Reported |  | Air KERMA | [6 empty] | HVL
22 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | #DIV/0! |  |  |  | Min | 0.31 | 0.33 | 0.35 | 0.37 | 0.39 | 0.41
23 | [8 empty] | ESE (mGy): |  |  |  |  | Max | 0.4 | 0.42 | 0.44 | 0.46 | 0.55 | 0.57
24 | [8 empty] | DGN (mrad/R): | TBD
25 | [8 empty] | MGD (mGy):
26 | [8 empty] | Prev MGD (mGy): |  |  |  | Half Value Layer – CEDM
27 |  | Average: | [6 empty] | % Change: |  |  |  |  | Target/Filter | W/Cu | W/Cu | W/Cu | W/Cu | W/Cu | W/Cu | W/Cu
28 |  | Coeff of Var: | [11 empty] | kV | 40 | 42 | 44 | 45 | 46 | 48 | 49
29 | [8 empty] | Deviation from Reported MGD: |  |  |  |  | Al (mm) | Air Kerma (mGy)
30 | [8 empty] | mAs at 3 mGy: |  |  |  |  | 0
31 |  | Mean Glandular Dose - High Energy | [11 empty] | HVL
32 |  | Target/Filter: | W/Cu |  | kV: | 45 | [7 empty] | Min | 0.43 | 0.45 | 0.47 | 0.48 | 0.49 | 0.51 | 0.52
33 |  |  |  |  | mA: | 100 | [7 empty] | Max | 0.52 | 0.54 | 0.56 | 0.57 | 0.58 | 0.67 | 0.68
34 |  |  |  |  | Reported |  | Air KERMA
35 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | 0
36 | [8 empty] | ESE (mGy): |  |  |  |  | Criteria: | Mo/Rh targets – HVL is between the minimum and maximum limits
37 | [8 empty] | DGN (mrad/R): | #N/A |  |  |  |  | W targets – HVL is > (kV/100)+0.03
38 | [8 empty] | MGD (mGy): | #N/A
39 | [8 empty] | Prev MGD (mGy): |  |  |  | Image Quality Evaluation
40 |  | Average: | [6 empty] | % Change: | #N/A |  |  | Printer | Model: |  | SN: |  |  | QC Manual Version:
41 |  | Coeff of Var:
42 |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #N/A | [5 empty] | Previous |  |  |  | Previous
43 |  |  | Deviation from reported MGD is less than15% | [10 empty] | kV: | 28 |  |  | Fibers:
44 |  |  | MGD for Low + High acquisitions is less than 3 mGy | [5 empty] | CEDM MGD (Low + High) (mGy): | #N/A |  |  |  | mAs: | 0 |  |  | Specks:
45 | [8 empty] | Prev CEDM MGD (Low + High) (mGy): |  |  |  |  | Background Density: |  |  |  | Masses:
46 | [13 empty] | Inside disk: |  |  |  | Acceptable:
47 | [13 empty] | Outside disk:
48 | [13 empty] | Difference:
49
50 | [12 empty] | Workstation
51 | [14 empty] | Workstation: |  |  |  | Workstation:
52 | [14 empty] | Left | Right |  |  | Left | Right
53 | [13 empty] | Fibers: |  |  |  | Fibers:
54 | [13 empty] | Specks: |  |  |  | Specks:
55 | [13 empty] | Masses: |  |  |  | Masses:
56
57 | [12 empty] | Acquisition Unit
58 | [13 empty] | 2D | Previous | Mag | Previous | 3D | Previous | SBB | Previous | SBB Tomo | Previous
59 | [12 empty] | kV: | 28 |  |  |  | 29
60 | [12 empty] | mAs:
61 | [12 empty] | EI:
62 | [12 empty] | Fibers:
63 | [12 empty] | Specks:
64 | [12 empty] | Masses:
65 | [13 empty] | Criteria: | 2D – Must be able to see 5 fibers, 4 speck groups, 4 masses
66 | [14 empty] | 3D – Must be able to see 4 fibers, 3 speck groups, 3 masses
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  |  | SBB Stereo - Must be able to see 5 fibers, 4 speck groups, 3.5 masses
68 |  | Room Number | [8 empty] | Survey ID: |  |  |  | SBB Tomo - Must be able to see 3 fibers, 3 speck groups, 2.5 masses
Signal to Noise
[18 empty] | Mean | SD
[14 empty] | kV: | 28 |  | In acrylic: | #DIV/0! | #DIV/0!
[14 empty] | mAs: | 0 |  | Background: | #DIV/0! | #DIV/0!
[14 empty] | Target: | W
[14 empty] | Filter: | Rh |  |  | Current | Previous | Change | Acceptable
[17 empty] | SNR: | #DIV/0! |  | #DIV/0! | #DIV/0!
[17 empty] | CNR: | #DIV/0! |  | #DIV/0! | NA
[14 empty] | Criteria: | SNR must be >= 40
CNR does not change by more than 15%
Previous Year Comments
Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0

## Sheet2
Measurements | [10 empty] | Measurements
Set kV (kV) | Set mAs (mAs) | Target | Filter | Calibration | Tube voltage (kV) | Exposure time (ms) | Exposure (mGy) | Exposure rate (mGy/s) | HVL (mm Al) | Calc mA | Light (lx)
24 | 50 | W | Rh | W/50 µm Rh
25 | 50 | W | Rh | W/50 µm Rh
26 | 50 | W | Rh | W/50 µm Rh
28 | 20 | W | Rh | W/50 µm Rh
28 | 50 | W | Rh | W/50 µm Rh
28 | 50 | W | Rh | W/50 µm Rh
28 | 50 | W | Rh | W/50 µm Rh
28 | 50 | W | Rh | W/50 µm Rh
28 | 100 | W | Rh | W/50 µm Rh
28 | 300 | W | Rh | W/50 µm Rh
28 | 400 | W | Rh | W/50 µm Rh
30 | 50 | W | Rh | W/50 µm Rh
32 | 50 | W | Rh | W/50 µm Rh
34 | 50 | W | Rh | W/50 µm Rh
36 | 50 | W | Rh | W/50 µm Rh
38 | 50 | W | Rh | W/50 µm Rh
28 | 50 | W | Ag | W/50 µm Ag
30 | 50 | W | Ag | W/50 µm Ag
32 | 50 | W | Ag | W/50 µm Ag
34 | 50 | W | Ag | W/50 µm Ag
36 | 50 | W | Ag | W/50 µm Ag
38 | 50 | W | Ag | W/50 µm Ag
28 | 50 | W | Al | W/0.70 mm Al
30 | 50 | W | Al | W/0.70 mm Al
32 | 50 | W | Al | W/0.70 mm Al
34 | 50 | W | Al | W/0.70 mm Al
36 | 50 | W | Al | W/0.70 mm Al
38 | 50 | W | Al | W/0.70 mm Al
40 | 50 | W | Cu | W/0.3 mm Cu
42 | 50 | W | Cu | W/0.3 mm Cu
44 | 50 | W | Cu | W/0.3 mm Cu
45 | 50 | W | Cu | W/0.3 mm Cu
46 | 50 | W | Cu | W/0.3 mm Cu
48 | 50 | W | Cu | W/0.3 mm Cu
49 | 50 | W | Cu | W/0.3 mm Cu
Waveforms

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
| kVp | [9 empty] | kVp | [9 empty] | kVp
HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
0.23 | 116 | [8 empty] | 0.23 | [9 empty] | 0.23
0.24 | 121 | 124 | [7 empty] | 0.24 | [9 empty] | 0.24
0.25 | 126 | 129 | 131 | [6 empty] | 0.25 | [9 empty] | 0.25
0.26 | 130 | 133 | 135 | 138 | [5 empty] | 0.26 | [9 empty] | 0.26
0.27 | 135 | 138 | 140 | 142 | 143 |  |  |  |  | 0.27 | [9 empty] | 0.27
0.28 | 140 | 142 | 144 | 146 | 147 | 149 |  |  |  | 0.28 |  |  | 149 | 151 | 154 |  |  |  |  | 0.28 |  |  | 150 | 155 | 159
0.29 | 144 | 146 | 148 | 150 | 151 | 153 | 154 |  |  | 0.29 |  |  | 154 | 156 | 158 | 159 |  |  |  | 0.29 |  |  | 155 | 160 | 164 | 168
0.3 | 149 | 151 | 153 | 155 | 156 | 157 | 158 | 159 |  | 0.3 |  |  | 158 | 160 | 162 | 162 | 163 |  |  | 0.3 |  |  | 160 | 164 | 168 | 172 | 176
0.31 | 154 | 156 | 157 | 159 | 160 | 161 | 162 | 163 | 164 | 0.31 |  |  | 163 | 164 | 166 | 166 | 167 | 167 |  | 0.31 |  |  | 165 | 168 | 172 | 174 | 180 | 182
0.32 | 158 | 160 | 162 | 163 | 164 | 166 | 167 | 168 | 168 | 0.32 |  |  | 167 | 169 | 171 | 171 | 171 | 172 | 172 | 0.32 |  |  | 169 | 173 | 177 | 181 | 184 | 186 | 188
0.33 | 163 | 165 | 166 | 168 | 169 | 170 | 171 | 173 | 173 | 0.33 |  |  | 171 | 173 | 175 | 176 | 176 | 176 | 176 | 0.33 |  |  | 174 | 178 | 181 | 185 | 188 | 190 | 192
0.34 | 168 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 0.34 |  |  | 176 | 178 | 179 | 179 | 180 | 180 | 180 | 0.34 |  |  | 179 | 183 | 186 | 190 | 193 | 195 | 196
0.35 |  | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 0.35 |  |  | 180 | 181 | 183 | 183 | 184 | 185 | 185 | 0.35 |  |  | 184 | 187 | 190 | 194 | 197 | 199 | 201
0.36 |  |  | 179 | 181 | 182 | 183 | 184 | 185 | 185 | 0.36 |  |  | 185 | 186 | 187 | 187 | 188 | 188 | 189 | 0.36 |  |  | 189 | 192 | 195 | 198 | 201 | 204 | 205
0.37 |  |  |  | 185 | 186 | 187 | 188 | 189 | 190 | 0.37 |  |  | 189 | 190 | 191 | 191 | 192 | 193 | 193 | 0.37 |  |  | 193 | 196 | 199 | 202 | 205 | 207 | 209
0.38 |  |  |  |  | 190 | 191 | 192 | 193 | 194 | 0.38 |  |  | 193 | 194 | 196 | 196 | 197 | 197 | 197 | 0.38 |  |  | 198 | 201 | 204 | 207 | 209 | 211 | 213
0.39 |  | 174 178: 196; 175 179: 197; 176 180: 198; 177 181: 198; 0.34 0.35: 0.39 |  |  | 198 | 199 | 200 | 200 | 201 | 201 | 202 | 0.39 |  |  | 203 | 206 | 208 | 211 | 214 | 216 | 217
0.4 |  | 175 179: 201; 176 180: 202; 177 181: 203; 0.34 0.35: 0.4 |  |  | 202 | 203 | 204 | 204 | 205 | 205 | 206 | 0.4 |  |  | 208 | 211 | 213 | 216 | 218 | 220 | 221
0.41 |  | 176 180: 206; 177 181: 207; 0.34 0.35: 0.41 |  |  | 206 | 207 | 208 | 208 | 209 | 209 | 210 | 0.41 |  |  | 213 | 215 | 217 | 220 | 222 | 224 | 225
0.42 |  | 177 181: 211; 0.34 0.35: 0.42 |  |  | 211 | 211 | 212 | 212 | 213 | 213 | 214 | 0.42 |  |  | 218 | 220 | 222 | 224 | 226 | 228 | 229
Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast | [12 empty] | Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
| kVp | [12 empty] | kVp
HVL | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | HVL | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39
0.3 | 152 | 157 | 163 | 166 | 170 | 173 | 175 | 177 | 179 | 182 | 184 | 187 | 0.4 | 222 | 226 | 229 | 231 | 234 | 236 | 239 | 241 | 244 | 246 | 248 | 250 | 252
0.325 | 163 | 169 | 174 | 177 | 181 | 183 | 186 | 188 | 190 | 192 | 195 | 197 | 0.425 | 233 | 236 | 239 | 242 | 244 | 246 | 248 | 251 | 253 | 256 | 258 | 260 | 262
0.35 | 175 | 180 | 185 | 188 | 191 | 194 | 196 | 198 | 200 | 202 | 205 | 207 | 0.45 | 244 | 247 | 249 | 252 | 254 | 256 | 258 | 260 | 263 | 265 | 267 | 269 | 271
0.375 | 186 | 191 | 196 | 199 | 202 | 205 | 207 | 209 | 211 | 213 | 215 | 218 | 0.475 | 254 | 257 | 260 | 262 | 264 | 266 | 268 | 270 | 273 | 275 | 277 | 279 | 281
0.4 | 198 | 203 | 207 | 210 | 213 | 215 | 217 | 219 | 221 | 223 | 226 | 228 | 0.5 | 265 | 267 | 270 | 272 | 274 | 276 | 278 | 280 | 282 | 284 | 286 | 288 | 290
0.425 | 209 | 214 | 218 | 221 | 224 | 226 | 228 | 230 | 232 | 234 | 236 | 238 | 0.525 | 275 | 278 | 280 | 282 | 284 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 300
0.45 | 221 | 226 | 230 | 232 | 235 | 237 | 238 | 240 | 242 | 244 | 246 | 248 | 0.55 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 299 | 301 | 303 | 305 | 307 | 309
0.475 | 233 | 237 | 241 | 243 | 245 | 247 | 249 | 251 | 253 | 254 | 256 | 258 | 0.575 | 296 | 298 | 300 | 302 | 304 | 305 | 307 | 309 | 311 | 313 | 315 | 317 | 318
0.5 | 244 | 248 | 252 | 254 | 256 | 258 | 260 | 261 | 263 | 265 | 267 | 269 | 0.6 | 306 | 308 | 310 | 312 | 313 | 315 | 317 | 319 | 320 | 322 | 324 | 326 | 328
0.525 | 256 | 260 | 263 | 265 | 267 | 269 | 270 | 272 | 273 | 275 | 277 | 279 | 0.625 | 316 | 318 | 320 | 322 | 323 | 325 | 326 | 328 | 330 | 332 | 333 | 335 | 337
0.55 | 267 | 271 | 274 | 276 | 278 | 279 | 281 | 282 | 284 | 285 | 287 | 289 | 0.65 | 326 | 328 | 330 | 331 | 333 | 334 | 336 | 338 | 339 | 341 | 343 | 344 | 346
0.575 | 279 | 282 | 285 | 287 | 288 | 290 | 291 | 292 | 294 | 296 | 297 | 299 | 0.675 | 336 | 338 | 339 | 341 | 342 | 344 | 345 | 347 | 349 | 350 | 352 | 354 | 355
0.6 | 290 | 293 | 296 | 297 | 299 | 300 | 301 | 303 | 304 | 306 | 308 | 310 | 0.7 | 346 | 348 | 349 | 350 | 352 | 353 | 355 | 356 | 358 | 359 | 361 | 363 | 364
0.625 | 301 | 304 | 306 | 308 | 309 | 310 | 312 | 313 | 315 | 316 | 318 | 320 | 0.725 | 356 | 357 | 358 | 360 | 361 | 362 | 364 | 365 | 367 | 368 | 370 | 372 | 373
0.65 | 312 | 314 | 317 | 318 | 320 | 321 | 322 | 323 | 325 | 326 | 328 | 330 | 0.75 | 365 | 367 | 368 | 369 | 370 | 372 | 373 | 375 | 376 | 378 | 379 | 381 | 382
0.675 | 322 | 325 | 327 | 328 | 330 | 331 | 333 | 333 | 335 | 336 | 338 | 340 | 0.775 | 374 | 376 | 377 | 378 | 379 | 381 | 382 | 383 | 385 | 386 | 388 | 390 | 391
0.7 | 333 | 335 | 337 | 339 | 340 | 341 | 342 | 343 | 345 | 346 | 348 | 350 | 0.8 | 384 | 385 | 386 | 387 | 388 | 390 | 391 | 392 | 394 | 395 | 397 | 398 | 400
0.725 | 342 | 345 | 347 | 348 | 349 | 351 | 352 | 353 | 354 | 356 | 358 | 360 | 0.825 | 393 | 394 | 395 | 396 | 397 | 399 | 400 | 401 | 403 | 404 | 406 | 407 | 408
0.75 | 352 | 355 | 357 | 358 | 359 | 360 | 361 | 363 | 364 | 366 | 368 | 369 | 0.85 | 402 | 403 | 404 | 405 | 406 | 407 | 409 | 410 | 411 | 413 | 414 | 416 | 417
0.775 | 361 | 365 | 367 | 368 | 369 | 370 | 371 | 372 | 374 | 375 | 377 | 379 | 0.875 | 410 | 411 | 412 | 413 | 415 | 416 | 417 | 418 | 420 | 421 | 423 | 424 | 425
0.8 | 369 | 374 | 376 | 377 | 378 | 379 | 380 | 382 | 383 | 385 | 387 | 389 | 0.9 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 427 | 428 | 429 | 431 | 432 | 434
Exposure to Dose Conversion Factors W/Al for 4.2cm 50/50 Breast
kVp
HVL | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49
0.2 | 120 | 125 | 133 | 138 | 143 | 148 | 156 | 160 | 165 | 169 | 177 | 181 | 185 | 188 | 195 | 198 | 200 | 203 | 208 | 216 | 219 | 222 | 225 | 231 | 234
0.25 | 143 | 148 | 155 | 160 | 165 | 169 | 176 | 181 | 185 | 189 | 196 | 200 | 204 | 207 | 213 | 216 | 219 | 221 | 226 | 233 | 236 | 239 | 241 | 247 | 250
0.3 | 166 | 171 | 177 | 182 | 187 | 191 | 197 | 202 | 206 | 210 | 216 | 220 | 223 | 226 | 232 | 235 | 237 | 239 | 244 | 250 | 253 | 256 | 258 | 264 | 266
0.35 | 189 | 195 | 200 | 204 | 209 | 213 | 218 | 222 | 226 | 230 | 236 | 239 | 243 | 246 | 250 | 253 | 255 | 257 | 262 | 267 | 270 | 273 | 275 | 280 | 282
0.4 | 212 | 217 | 222 | 226 | 231 | 235 | 239 | 243 | 247 | 251 | 255 | 258 | 262 | 265 | 269 | 271 | 273 | 275 | 280 | 285 | 287 | 290 | 292 | 296 | 298
0.45 | 234 | 239 | 244 | 248 | 252 | 256 | 260 | 264 | 268 | 271 | 275 | 278 | 281 | 284 | 287 | 290 | 292 | 294 | 298 | 302 | 304 | 307 | 309 | 313 | 315
0.5 | 256 | 261 | 265 | 269 | 273 | 277 | 280 | 284 | 288 | 291 | 294 | 297 | 300 | 303 | 306 | 308 | 310 | 312 | 316 | 319 | 321 | 324 | 326 | 329 | 331
0.55 | 278 | 282 | 286 | 290 | 293 | 297 | 300 | 304 | 307 | 310 | 313 | 316 | 319 | 322 | 324 | 326 | 328 | 330 | 334 | 336 | 339 | 341 | 343 | 345 | 347
0.6 | 300 | 303 | 307 | 310 | 313 | 317 | 320 | 323 | 326 | 329 | 332 | 335 | 337 | 340 | 342 | 344 | 346 | 348 | 351 | 354 | 356 | 358 | 360 | 362 | 363
0.65 | 321 | 324 | 327 | 330 | 333 | 336 | 339 | 342 | 345 | 347 | 350 | 353 | 355 | 358 | 360 | 362 | 363 | 365 | 368 | 370 | 372 | 374 | 376 | 378 | 380
0.7 | 341 | 344 | 347 | 350 | 352 | 355 | 358 | 361 | 363 | 366 | 368 | 370 | 373 | 375 | 377 | 379 | 380 | 382 | 385 | 387 | 389 | 390 | 392 | 394 | 395
0.75 | 360 | 363 | 366 | 369 | 370 | 374 | 376 | 379 | 381 | 383 | 386 | 388 | 390 | 392 | 394 | 395 | 397 | 398 | 401 | 403 | 404 | 406 | 408 | 409 | 411
0.8 | 379 | 382 | 385 | 387 | 389 | 392 | 394 | 396 | 398 | 401 | 403 | 405 | 407 | 408 | 410 | 412 | 413 | 414 | 417 | 418 | 420 | 421 | 423 | 424 | 426
0.85 | 398 | 400 | 403 | 405 | 407 | 409 | 411 | 413 | 415 | 417 | 419 | 421 | 423 | 425 | 426 | 427 | 429 | 430 | 432 | 434 | 435 | 436 | 438 | 439 | 440
0.9 | 415 | 418 | 420 | 422 | 424 | 426 | 428 | 430 | 432 | 434 | 435 | 437 | 439 | 440 | 442 | 443 | 444 | 445 | 447 | 449 | 450 | 451 | 452 | 454 | 455
0.95 | 432 | 435 | 437 | 438 | 440 | 442 | 444 | 446 | 448 | 449 | 451 | 453 | 454 | 455 | 457 | 458 | 459 | 460 | 462 | 463 | 464 | 465 | 467 | 468 | 469
1 | 448 | 451 | 453 | 454 | 456 | 458 | 460 | 461 | 463 | 465 | 466 | 467 | 469 | 470 | 471 | 472 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482
Exposure to Dose Conversion Factors W/Cu for 4.2cm 50/50 Breast
kVp
HVL | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49
2 | 677 | 675 | 674 | 672 | 671 | 669 | 668 | 667 | 666 | 665
2.1 | 694 | 692 | 690 | 689 | 687 | 686 | 684 | 683 | 681 | 680
2.2 | 711 | 709 | 707 | 705 | 704 | 702 | 700 | 699 | 697 | 696
2.3 | 726 | 724 | 722 | 721 | 719 | 717 | 715 | 714 | 712 | 711
2.4 | 740 | 739 | 737 | 736 | 734 | 732 | 731 | 729 | 727 | 725
2.5 | 753 | 752 | 751 | 750 | 748 | 747 | 745 | 743 | 741 | 739
2.6 | 766 | 765 | 764 | 763 | 762 | 760 | 758 | 757 | 755 | 753
2.7 | 777 | 777 | 776 | 775 | 774 | 773 | 771 | 770 | 768 | 766
2.8 | 788 | 788 | 787 | 787 | 786 | 785 | 783 | 782 | 780 | 779
2.9 | 799 | 799 | 798 | 798 | 797 | 796 | 795 | 794 | 792 | 791
3 | 809 | 808 | 808 | 807 | 807 | 806 | 805 | 804 | 803 | 802
3.1 | 818 | 818 | 818 | 817 | 817 | 816 | 815 | 814 | 813 | 812
3.2 | 827 | 827 | 827 | 826 | 826 | 825 | 824 | 824 | 823 | 822
3.3 | 836 | 836 | 835 | 835 | 834 | 834 | 833 | 833 | 832 | 831
3.4 | 843 | 844 | 844 | 843 | 843 | 842 | 842 | 841 | 840 | 840
3.5 | 849 | 851 | 852 | 851 | 851 | 850 | 850 | 849 | 848 | 848
3.6 | 854 | 858 | 859 | 859 | 859 | 858 | 857 | 857 | 856 | 856
3.7 | 858 | 863 | 865 | 866 | 866 | 865 | 865 | 864 | 864 | 863
3.8 | 862 | 867 | 871 | 873 | 873 | 872 | 872 | 871 | 871 | 870
3.9 | 865 | 871 | 876 | 878 | 879 | 879 | 879 | 878 | 877 | 877
4 | 868 | 875 | 880 | 883 | 885 | 885 | 885 | 885 | 884 | 883
kV Accuracy/Output |  | 170 174: Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV | [8 empty] | Radcal kV correction for Hologic Selenia/Dimension
Target | kV | mGy/mAs | mGy/s | Target | kV | mGy/mAs | mGy/s | Target | kV | mGy/mAs | mGy/s | Target | kV | mGy/mAs | mGy/s |  |  |  |  | Selenia |  | Dimension
W/Rh | 24 | 0 | 0 | W/Ag | 28 | 0 | 0 | W/Al | 28 | 0 | 0 | W/Cu | 40 | 0 | 0 |  |  |  | Measured | W/Rh | W/Ag | W/Rh | W/Ag | W/Al
Filter | 25 | 0 | 0 | Filter | 30 | 0 | 0 | Filter | 30 | 0 | 0 | Filter | 42 | 0 | 0
Rh | 26 | 0 | 0 | Ag | 32 | 0 | 0 | Al | 32 | 0 | 0 | Cu | 44 | 0 | 0
| 28 | #DIV/0! | #DIV/0! |  | 34 | 0 | 0 |  | 34 | 0 | 0 |  | 45 | 0 | 0
| 30 | 0 | 0 |  | 36 | 0 | 0 |  | 36 | 0 | 0 |  | 46 | 0 | 0
| 32 | 0 | 0 |  | 38 | 0 | 0 |  | 38 | 0 | 0 |  | 48 | 0 | 0
| 34 | 170 174: 0; 171 175: 0 |  |  | 49 | 0 | 0
| 36 | 170 174: 0; 171 175: 0
| 38 | 170 174: 0; 171 175: 0
| Slope: | #DIV/0! | #DIV/0! |  | Slope: | 0 | 0 |  | Slope: | 0 | 0 |  | Slope: | 0 | 0
| Intercept: | #DIV/0! | #DIV/0! |  | Intercept: | 0 | 0 |  | Intercept: | 0 | 0 |  | Intercept: | 0 | 0
| R^2: | #DIV/0! | #DIV/0! |  | R^2: | #DIV/0! | #DIV/0! |  | R^2: | #DIV/0! | #DIV/0! |  | R^2: | #DIV/0! | #DIV/0!
Half Value Layer
Target | kV | HVL | Target | kV | HVL | Target | kV | HVL | Target | kV | HVL
W/Rh | 24 | 0 | W/Ag | 28 | 0 | W/Al | 28 | 0 | W/Cu | 40 | 0
Filter | 25 | 0 | Filter | 30 | 0 | Filter | 30 | 0 | Filter | 42 | 0
Rh | 26 | 0 | Ag | 32 | 0 | Al | 32 | 0 | Cu | 44 | 0
| 28 | #DIV/0! |  | 34 | 0 |  | 34 | 0 |  | 45 | 0
| 30 | 0 |  | 36 | 0 |  | 36 | 0 |  | 46 | 0
| 32 | 0 |  | 38 | 0 |  | 38 | 0 |  | 48 | 0
| 34 | 170 174: 0; 0.34 0.35: 49 | 0
| 36 | 170 174: 0
| 38 | 170 174: 0
| Slope: | #DIV/0! |  | Slope: | 0 |  | Slope: | 0 |  | Slope: | 0
| Intercept: | #DIV/0! |  | Intercept: | 0 |  | Intercept: | 0 |  | Intercept: | 0
| R^2: | #DIV/0! |  | R^2: | #DIV/0! |  | R^2: | #DIV/0! |  | R^2: | #DIV/0!
Target | Filter | 170 174: C; 172 176: Critiera:; 173 177: Min HVL > kV/100 (no paddle)
Mo | Mo | 170 174: 0.12; 173 177: Min HVL > kV/100 + 0.03 (with paddle)
Mo | Rh | 170 174: 0.19; 173 177: Max HVL < kV/100 + C
Rh | Rh | 170 174: 0.22
DGN | #DIV/0! | 174 178: Combo
Mo/Mo | Mo/Rh | Rh/Rh | W/Rh | W/Ag | W/Al | W/Rh | W/Ag | W/Al | W/Cu
#DIV/0! | #DIV/0! | #DIV/0! | TBD | TBD | #N/A | TBD | TBD | #N/A | #N/A
Radcal kV Correction for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Rh | < | 27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| >= | 27.586 | -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Ag | < | 30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| >= | 30.1 | -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
DGN Mo/Mo: Radcal kV Correction for Hologic Dimension | Barracuda/MPD kVp correction table for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | W/Rh (RQ M6) |  |  | W/Ag (RQ M10)
W/Rh | < | 26.9 | 138.8867 | -10.7264 | 0.26216 | -0.00082 |  |  |  | kV | LFS | SFS | kV | LFS | SFS
| >= | 26.9 | -5009.7752 | 605.732 | -27.3019 | 0.54671 | -0.0040986 |  |  | 22 | 0.2 | 0.1 | 22 | 0.2 | -0.2
[10 empty] | 23 | 0.2 | -0.1 | 23 | 0.4 | -0.1
W/Ag | kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | 24 | 0.1 | -0.4 | 24 | 0.4 | 0
| < | 28.7 | 296.3419 | -31.6292 | 1.1803 | -0.01417 |  |  |  | 25 | 0.1 | -0.3 | 25 | 0.5 | -0.1
#DIV/0! Mo/Rh: >=; #DIV/0! Rh/Rh: 28.7; #DIV/0! W/Rh: 4.8345; #DIV/0! W/Ag: 0.91924 | 26 | 0 | -0.2 | 26 | 0.5 | -0.2
[10 empty] | 27 | 0.1 | -0.3 | 27 | 0.7 | -0.2
W/Al |  |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | 28 | 0.2 | -0.5 | 28 | 0.9 | -0.1
#DIV/0! Mo/Rh: <; #DIV/0! Rh/Rh: 28.7; #DIV/0! W/Rh: 49.3111; #DIV/0! W/Ag: -2.9302; #DIV/0! W/Al: 0.07379 | 29 | 0.4 | -0.2 | 29 | 0.8 | -0.3
#DIV/0! Rh/Rh: 28.7 < kV < 30.1; #DIV/0! W/Rh: -24.875; #DIV/0! W/Ag: 1.8032 | 30 | 0.6 | 0 | 30 | 0.8 | -0.4
#DIV/0! Mo/Rh: >=; #DIV/0! Rh/Rh: 30.1; #DIV/0! W/Rh: -4.8346; #DIV/0! W/Ag: 1.1571 | 31 | 0.8 | 0.1 | 31 | 0.8 | -0.3
[10 empty] | 32 | 1 | 0.2 | 32 | 0.8 | -0.3
DGN Mo/Mo: CNR Correction | 33 | 1 | 0.4 | 33 | 0.9 | -0.3
DGN Mo/Mo: Conventional (XX8xxxxx series detectors) | 34 | 1.1 | 0.6 | 34 | 1.1 | -0.3
Thickness (cm) | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 |  | 35 | 1.2 | 0.6 | 35 | 1.1 | -0.2
2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1.25 |  | 36 | 1.4 | 0.6 | 36 | 1.2 | -0.2
4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1.2 |  | 37 | 1.5 | 0.7 | 37 | 1.1 | -0.1
6 | 1.04 | 1.02 | 1.01 | 1.08 | 1.35 | 1.35 | 1.31 | 1.35 |  | 38 | 1.6 | 0.9 | 38 | 1.1 | 0
8 | 1.23 | 1.2 | 1.18 | 1.1 | 1.82 | 1.85 | 1.82 | 1.82 |  | 39 | 1.7 | 0.8 | 39 | 1.3 | 0.3
Magnification (XX8xxxxx series detectors)
Thickness (cm) | 0 | 1 | 2 | 3 | 4
2 | 1.15 | 1.15 | 1.15 | 1.15 | 1.5
4 | 1.15 | 1.15 | 1.15 | 1.15 | 1.5
6 | 1.19 | 1.18 | 1.18 | 1.18 | 1.55
8 | 1.28 | 1.24 | 1.22 | 1.29 | 1.67
Stereotactic biopsy (XX8xxxxx series detectors) |  |  |  | CE2D (XX8xxxxx series detectors)
Thickness (cm) | 0 |  |  | Thickness | 0 Low | 0 High
2 | 0.91 |  |  | 2 | 1 | 0.72
4 | 1 |  |  | 4 | 1 | 0.9
6 | 1.32 |  |  | 6 | 1.91 | 1.71
8 | 1.88 |  |  | 8 | 1.81 | 2.22
For software versions Dimensions 1.8.2 - 1.9.x and 3Dimensions 2.0.x and detector SN XX8xxxxx
Tomo 15 Proj STD (XX8xxxxx series detectors) |  |  |  |  | Tomo 15 Proj ENH (XX8xxxxx series detectors)
Thickness (cm) | 0 | 1 | 2 |  | Thickness | 0
2 | 0.7 | 0.7 | 0.88 |  | 2 | 0.57
4 | 0.91 | 0.91 | 0.89 |  | 4 | 0.91
6 | 1.55 | 1.55 | 1.59 |  | 6 | 1.68
8 | 2.78 | 2.78 | 2.68 |  | 8 | 1.93
For software versions Dimensions 1.10.x+ and 3Dimensions 2.1.x+ and detector SN XX8xxxxx
Tomo 15 Proj STD (XX8xxxxx series detectors)
Thickness (cm) | 0 | 1 | 2
2 | 0.64 | 0.7 | 0.88
4 | 0.91 | 0.91 | 0.89
6 | 1.55 | 1.55 | 1.59
8 | 2.78 | 2.78 | 2.68
CNR Correction (Selenia Dimension W/Rh)
Thickness (cm) | 0 |  | 1 |  | 2 |  | 2
| LFS | SFS | LFS | SFS | LFS | SFS | LFS | SFS
0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
3 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
5 | 1 | 1 | 1 | 1 | 1.15 | 1 | 1 | 1
6 | 1 | 1 | 1 | 1.15 | 1.15 | 1 | 1 | 1
7 | 1.1 | 1 | 1.1 | 1.15 | 1.15 | 1 | 1 | 1
8 | 1.15 | 1 | 1.15 | 1.15 | 1.15 | 1 | 1 | 1
9 | 1.1 | 1 | 1.1 | 1.15 | 1.15 | 1 | 1 | 1
10 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
11 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
12 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
13 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
14 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
CNR Correction (Affirm Prone Biopsy)
| Standard biopsy |  |  | Tomo biopsy
Thickness (cm) | 0 | 1 | Thickness (cm) | 0 | 1
2 | 0.71 | 0.69 | 2 | 0.88 | 0.65
4 | 0.94 | 0.98 | 4 | 0.94 | 0.96
6 | 1.3 | 1.09 | 6 | 1.53 | 1.61
8 | 1.71 | 1.53 | 8 | 2.32 | 2.91
Note on A71: Mah, Eugene: Selenia 3Dimensions QC Manual MAN-03706 Rev 010
Note on A160: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.2, MAN-03706 Rev 10
Note on A166: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.2, MAN-03706 Rev 10
Note on A172: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.3, MAN-03706 Rev 10
Note on E172: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.4, MAN-03706 Rev 10
Note on A180: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.5.1, MAN-03706 Rev 10
Note on F180: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.6, MAN-03706 Rev 10
Note on A188: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.5.2, MAN-03706 Rev 10

## DataPage
Sheet | MAMMO_HOL
SurveyID
Light field
Resolution | 2D | 0 | 3D | 0
| W | 0 | W | 0
| 0 | 0 | 0 | 0
2D MGD
3D MGD | #N/A
Combo MGD | #N/A
kV Accuracy | W/Rh | kV | Output | W/Ag | kV | Output | W/Al | kV | Output
Output | 24 |  |  | 28 |  |  | 28
| 25 |  |  | 30 |  |  | 30
| 26 |  |  | 32 |  |  | 32
| 28 |  |  | 34 |  |  | 34
| 30 |  |  | 36 |  |  | 36
| 32 |  |  | 38
34
Linearity | W/Rh
20
50
100
300
HVL | W/Rh | 24
| W/Rh | 25
| W/Rh | 28
| W/Rh | 32
| W/Ag | 28
| W/Ag | 30
| W/Ag | 32
| W/Ag | 34
| W/Al | 28
| W/Al | 30
| W/Al | 32
| W/Al | 34
| W/Al | 36
Acq | 2D | Mag | 3D | Film
Fibers | 0 | 0 | 0 | 0
Specks | 0 | 0 | 0 | 0
Masses | 0 | 0 | 0 | 0
SNR target/filter | W/Rh
SNR | #DIV/0!
CNR | #DIV/0!
Note on B4: Eugene Mah: Target/Filter
Note on B5: Eugene Mah: kV
Note on B6: Eugene Mah: Line pairs/mm

## Corrected kV
GENERAL
Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
Read the true corrected kV in the adjacent cell to the right.
If that cell is light green, the corrected kV is accurate.
If the cell is red, the entered kV is out of range and the correction should not be used.
If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.
Tungsten-Rhodium (W / Rh)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| Radcal kV | True kV
| 29 | 30.2821
Tungsten-Silver (W / Ag)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
| Radcal kV | True kV
| 37.1 | 38.9675
27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
Selenia
Selenia Dimensions
Affirm Prone SBB
3Dimensions