## Barco
1					Revision 1.1-201500513					Print Area
2					Medical University of South Carolina					Medical University of South Carolina
3					Charleston, South Carolina					Charleston, South Carolina
4					All:
5					Mammography Monitor Compliance Inspection					Mammography Monitor Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10			Facility:					Site Number:					Facility:				Site Number:					Facility:
11			Department:					Date of Installation:					Department:				Date of Installation:					Department:
12			Area/Division:					Workstation ID:					Area/Division:				Workstation ID:					Area/Division:
13			Survey ID:					Left monitor SN:	123456789					Survey ID:				Left monitor SN:	123456789	123456789				Survey ID:
14			Room Number:					Right monitor SN:	987654321					Room Number:				Right monitor SN:	987654321	987654321				Room Number:
15			QC Manual version:					QC Manual version:					QC Manual version:
16					Site Number:
17					Medical University of South Carolina					Inspection Results					Date of Installation:
18					Mammography Monitor Compliance Inspection		Enter 1 for YES, 2 for NO, 3 for NA					Workstation ID:
19					Corrective Actions & Comments					Comments					Left monitor SN:		Change	123456789
20			Monitor surface is clean	TBD					Monitor surface is clean					Right monitor SN:		Change	987654321
21			Viewing conditions	TBD					Viewing conditions
22			TG18-QC test pattern	TBD					TG18-QC test pattern					Ambient Light:		Change	1.3
23			Rapid Frame acceptable	TBD					Rapid Frame acceptable
24			Display Test	TBD
25					Room Conditions
26		Room Conditions					Ambient Light:	1.3	lux
27		Ambient light:	1.3	lux		Criteria:	Ambient light measured at the surface of the monitor(s) must be less than 20 lux					Previous:		lux	Criteria:	Ambient light measured at the surface of the monitor(s) must be less than 20 lux
28		Comments:					Comments:
29					Do not exceed 390 chars		0
30					New:
31
32		Auto Calibration					Auto Calibration
33		Date:			Lmin	Lmax	Luminance Ratio					Date:			Left (6789)	Right (4321)
34				Left (6789)	0.59	600.2	1017.3					Lmin (cd/m^2)	0.59	0.89
35				Right (4321)	0.89	600.1	674.3					Lmax (cd/m^2	600.2	600.1
36		Comments:					Luminance Ratio	1017.3	674.3
37
38					Comments:
39					Do not exceed 390 chars		0
40		Compliance Test					New:
41		Date:			Left (6789)	Right (4321)
42				DICOM GSDF curve acceptable:	YES	YES					Compliance Test
43		Comments:					Date:
44					Enter 1 for YES, 2 for NO, 3 for NA			Left (6789)	Right (4321)
45					DICOM GSDF curve acceptable:	1	1
46					Comments:
47		TG-18 Test Patterns					Do not exceed 390 chars		0
48		TG-18 LN01					TG-18 LN18					New:
49		Left (6789)	0.8	Right (4321)	0.88			Left (6789)	461.6	Right (4321)	484.3
50		Comments:					Display Test
51					Date:
52					Comments:
53					Do not exceed 390 chars		0
54		TG-18 QC					New:
55					Left (6789)	Right (4321)
56				5%/95% patches visible:	YES	YES					TG-18 Test Patterns
57				Resolution acceptable:	YES	YES					TG-18 Black/White		Left (6789)	Right (4321)
58				Display is free of artifacts/noise:	YES	YES					LN-01 (cd/m^2)	0.8	0.88
59		Comments:					LN-18 (cd/m^2)	461.6	484.3
60					Comments:
61					Do not exceed 390 chars		0
62					New:
63
64					TG-18 QC
65		Date:					Inspector:	Eugene Mah		Enter 1 for YES, 2 for NO, 3 for NA			Left (6789)	Right (4321)
66		Room Number					Survey ID:					5%/95% patches visible:	1	1
1					Medical University of South Carolina				Resolution acceptable:	1	1
2					Mammography Monitor Compliance Inspection				Display is free of artifacts/noise:	1	1
3					Comments:
4		TG-18 Uniformity					Do not exceed 390 chars		0
5			UNL-10 (cd/m^2)		UNL-80 (cd/m^2)			Visual Assessment					New:
6			Left (6789)	Right (4321)	Left (6789)	Right (4321)		YES	Free from gross non-uniformities
7		Average:	3.0	3.0	197.7	197.9		YES	Free from non-symmetric non-uniformities					TG-18 Uniformity					Visual assessment
8		SD:	0.0	0.1	4.7	4.8		YES	Free from significant luminance variation					UNL-10 (cd/m^2)		UNL-80 (cd/m^2)			Enter 1 for YES, 2 for NO, 3 for NA
9		Deviation (%)	1.7%	2.9%	3.3%	3.1%					Left (6789)	Right (4321)	Left (6789)	Right (4321)		1	Free from gross non-uniformities
10			PASS	PASS	PASS	PASS		Criteria:	Luminance deviation is < 30%					Top left	3.02	3.05	196.6	203.6		1	Free from non-symmetric non-uniformities
11		Comments:					Top right	2.94	2.91	198.1	197.6		1	Free from significant luminance variation
12					Center	2.96	3.05	191.1	195.5
13					Bottom left	3.01	2.88	198.4	191.5
14					Bottom right	3.04	2.93	204.3	201.3
15		Comments:					Average (cd/m^2):	2.99	2.96	197.70	197.90
16					SD (cd/m^2):	0.04	0.08	4.72	4.77
17					Luminance Deviation (%):	1.7%	2.9%	3.3%	3.1%	Criteria:	Luminance deviation is < 30%
18					Comments:
19					Do not exceed 390 chars		0
20					New:
21
22
23					Comments Page 1
24					Previous Year Comments
25					Comments:
26					Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
27					Additional Comments:
28					Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
29					Additional Comments:
30					Do not exceed 390 chars		0
31					Additional Comments:
32					Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
33					Additional Comments:
34					Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
35					Additional Comments:
36					Do not exceed 390 chars		0
37					Additional Comments:
38					Do not exceed 390 chars		0
39					Additional Comments:
40					Do not exceed 390 chars		0
41					Additional Comments:
42					Do not exceed 390 chars		0
43					Additional Comments:
44					Do not exceed 390 chars		0
45					Additional Comments:
46					Do not exceed 390 chars		0
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID: