## Sheet1
1					Revision 1.0-20150602					Print Area
2					Medical University of South Carolina					Medical University of South Carolina
3					Charleston, South Carolina					Charleston, South Carolina
4					All:
5					Bone Mineral Density Compliance Inspection					Bone Mineral Density Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Number of X-ray Tubes:					Survey ID:					Survey ID:
14					Room Number:					Room Number:
15			X-Ray Generator					Site Number:
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18				Max kVp:					Max mA:					Model:				Serial Number:
19					Max kVp:				Max mA:
20			X-Ray Tube 1					Focal Spot Sizes (mm)					X-Ray Generator
21				Tube Designation/Use:					Large:					X-Ray Tube 1					Focal Spot Sizes (mm)					Manufacturer:
22				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Model:
23			Insert					Micro:					Manufacture Date:				Small:					Max kVp:
24				Manufacturer:					Insert					Micro:					Manufacture Date:
25				Model:					Filtration					Manufacturer:					Serial Number:
26				Serial Number:					Inherent:					Model:				Filtration					Max mA:
27			Housing					Added:					Serial Number:				Inherent:					X-Ray Tube 1
28				Manufacturer:					Housing					Added:					Tube Designation/Use:
29				Model:					Manufacturer:					Manufacture Date:
30				Serial Number:					Model:					Manufacturer:
31					Serial Number:					Model:
32					Serial Number:
33					Inspection Results					Bone Mineral Density Compliance Inspection					Manufacturer:
34		Rule Number			Labels, Notices, Postings					Compliance					Inspection Results					Model:
35		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Serial Number:
36		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD					Labels, Notices, Postings					Large:
37		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD				DHEC Registration sticker is present, clearly visible and legible					Small:
38				Operator manuals are available.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Micro:
39				Monthly radiation monitoring reports are posted.					TBD				Radiation warning label posted on the generator control panel					Inherent:
40					Operator manuals are available.					Added:
41					Radiation Safety					Monthly radiation monitoring reports are posted.
42				Operator has full visibility of the exposure factors and patient					TBD
43				Operator position is at least 1m away from patient and unit					TBD					Radiation Safety
44					Operator has full visibility of the exposure factors and patient
45					Physical Inspection					Operator position is at least 1m away from patient and unit
46				System clean and free of debris					TBD
47				All cables free from excessive wear or restraints					TBD					Physical Inspection
48				System free of missing or damaged components					TBD				System clean and free of debris
49				All other movements and locks functional					TBD				All cables free from excessive wear or restraints
50					System free of missing or damaged components
51					All other movements and locks functional
52
53
54					Scan Techniques					Measurement Locations
55					Field size
56					Mode	kV	mA	Time (s)	Trans (cm)	Long (cm)		Location	Distance (cm)
57					Array	140/100	2.5	82	11.43	20.32		Tableside	50
58					High Def	140/100	2.5	163	11.43	20.32		Operator	135
59					Fast Array	140/100	2.5	41	11.43	20.32		Door	240
60					Turbo	140/100	2.5	15	11.43	20.32
61
62					Background
63					Time (s):	82	Reading (nGy):	1.2	1.2	1.2		Avg Bkg: (nGy):	1.2
64
65		Date:					Inspector:	Eugene Mah		Scatter Measurements					Average
66		Room Number					Survey ID:			Array	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
1					Medical University of South Carolina			Tableside	50	369.8	369.8	369.8	368.6	16.18
2					Inspection Results					Bone Mineral Density Compliance Inspection			Operator	135	54	54	54	52.8	2.32
3		Scan Techniques					Door	240	32	32	32	30.8	1.35
4					Field Size			Measurement Locations
5			Mode	kV	mA	Time (s)	Trans (cm)	Long (cm)		Location	Distance (cm)
6			Array	140/100	2.5	82	11.43	20.32		Tableside	50			High Def	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
7			High Def	140/100	2.5	163	11.43	20.32		Operator	135				Tableside	50	351.3			348.91	7.71
8			Fast Array	140/100	2.5	41	11.43	20.32		Door	240				Operator	135	53.5			51.11	1.13
9			Turbo	140/100	2.5	15	11.43	20.32					Door	240	31.4			29.01	0.64
10
11
12					Fast Array	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
13					Tableside	50	194.3			193.70	17.01
14					Operator	135	27.3			26.70	2.34
15		Scatter Measurements					Door	240	17.2			16.60	1.46
16
17			Array	Location	Distance (cm)	nGy/scan	uGy/h
18				Tableside	50	368.6	16.18					Turbo	Location	Distance (cm)	nGy/scan			nGy/scan	uGy/h
19				Operator	135	52.8	2.32					Tableside	50	136.7			136.48	32.76
20				Door	240	30.8	1.35					Operator	135	19.6			19.38	4.65
21					Door	240	11.3			11.08	2.66
22
23			High Def	Location	Distance (cm)	nGy/scan	uGy/h
24				Tableside	50	348.91	7.71
25				Operator	135	51.11	1.13
26				Door	240	29.01	0.64
27
28
29			Fast Array	Location	Distance (cm)	nGy/scan	uGy/h
30				Tableside	50	193.70	17.01
31				Operator	135	26.70	2.34
32				Door	240	16.60	1.46
33
34
35			Turbo	Location	Distance (cm)	nGy/scan	uGy/h
36				Tableside	50	136.48	32.76
37				Operator	135	19.38	4.65
38				Door	240	11.08	2.66
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID: