## Table
[6 empty] | Shielding Calculations | [8 empty] | Reference List
Location: | [13 empty] | 1. | Rossi RP, Ritenour R, Christodoulou E; Broad Beam Transmission Properties ..., Health Physics 61(5) 601-608 (1991)
[14 empty] | 2. | Simpkin D; Shielding Requirements for Mammography, Health Physics 53(3) 267-279 (1987). | [16 empty] | Min.
Date: | [5 empty] | Performed by: | Eugene Mah, M.Sc. | [6 empty] | 3. | NCRP49 | [14 empty] | mm | inch | Pb
[14 empty] | 4. | NCRP50 | [14 empty] | -25.4 | -1 | 0"
[14 empty] | 5. | Archer BR; Radiation Design Criteria for Diagnostic X-Ray Facilities, Health Physics, 44(5), 507-517, 1983 | [14 empty] | 0 | 0 | <=1/256"
|  | Ave. number of patients/wk: | [11 empty] | 6. | Simpkin D; A General Solution Shielding of Medical X- and Gamma-Rays by NCRP49 Methods; Health Physics 52(?) 431-436 (1987). | [14 empty] | 0.099219 | 0.0039062 | 0.0078125
|  | Ave. mAs/exposure: |  |  |  | Continuous Current: | [7 empty] | 7. | Archer BR, Fewell TR, et al; Attenuation Properties of Diagnostic X-Ray Shielding Materials. Med Phy 21(9) 1499-1507 (1994). | [14 empty] | 0.19844 | 0.0078125 | 0.015625
|  | Number of exposures/patient: |  |  |  | Field Size (cm^2): | [7 empty] | 8. | Simpkin DJ, Transmission Data for Shielding Diagnostic X-Ray Faciliites; Health Physics68(5), 704-709, 1995. | [14 empty] | 0.39688 | 0.015625 | 0.03125
|  | Weekly Workload (mA*min): | [27 empty] | 0.59531 | 0.023438 | 0.03125
[23 empty] | Scatter to Incident Ratio (a) | [6 empty] | 0.79375 | 0.03125 | 0.03125
|  | kVp: |  | (acceptable values are 25-150 kVp in steps of 5 kVp) | [18 empty] | angle in degrees | [6 empty] | 0.99219 | 0.039062 | 0.0625
[21 empty] | kVp | 30 | 45 | 60 | 90 | 120 | 135 |  |  | 1.1906 | 0.046875 | 0.0625
Fitting Parameters for the above conditions (mm^-1) | [29 empty] | 1.3891 | 0.054688 | 0.0625
Attenuator | a | b | g | Ko' | HVL | TVL |  |  |  | Conversion Factors: | 0.3048 | m/ft | [8 empty] | 30 | [8 empty] | 1.5875 | 0.0625 | 0.0625
Lead | [9 empty] | 1/16 = | 0.0625 | [9 empty] | 35 | [8 empty] | 1.7859 | 0.070312 | 0.09375
Concrete | [9 empty] | 1/8 = | 0.125 | [9 empty] | 50 | 0.0005 | 0.0002 | 0.00025 | 0.00035 | 0.0008 | 0.001 |  |  | 1.9844 | 0.078125 | 0.09375
Gypsum | [7 empty] | Conversion Assistance | [12 empty] | 55 | 0.0005375 | 0.0002375 | 0.000275 | 0.0003875 | 0.00085 | 0.001075 |  |  | 2.1828 | 0.085938 | 0.09375
Steel | [7 empty] | Feet | Meter | Meter | Feet | [9 empty] | 60 | 0.000575 | 0.000275 | 0.0003 | 0.000425 | 0.0009 | 0.00115 |  |  | 2.3813 | 0.09375 | 0.09375
Glass | [8 empty] | 0 |  | 0 | [9 empty] | 65 | 0.0006125 | 0.0003125 | 0.000325 | 0.0004625 | 0.00095 | 0.001225 |  |  | 2.5797 | 0.10156 | 0.125
Wood | [7 empty] | inch | mm | mm | inch | [9 empty] | 70 | 0.00065 | 0.00035 | 0.00035 | 0.0005 | 0.001 | 0.0013 |  |  | 2.7781 | 0.10938 | 0.125
[9 empty] | 0 |  | 0 |  |  |  | Primary | [5 empty] | 75 | 0.00079167 | 0.00049167 | 0.00049167 | 0.00063333 | 0.0011667 | 0.00145 |  |  | 2.9766 | 0.11719 | 0.125
| All distances are in meters | [6 empty] | Scatter |  | P | Primary Kux | Scatter Kux |  | # of HVL | # of HVL | ratio |  |  |  |  | 80 | 0.00093333 | 0.00063333 | 0.00063333 | 0.00076667 | 0.0013333 | 0.0016 |  |  | 3.175 | 0.125 | 0.125
Barrier | D(pri) | D(sec) | D(sca) | D(leak) | U** | T | Area | Angle | ratio (a) | (mSv/wk) | (mSv) | (mSv) | Leakage Blx | N | Np | column |  |  |  |  | 85 | 0.001075 | 0.000775 | 0.000775 | 0.0009 | 0.0015 | 0.00175 |  |  | 3.3734 | 0.13281 | 0.15625
[21 empty] | 90 | 0.0012167 | 0.00091667 | 0.00091667 | 0.0010333 | 0.0016667 | 0.0019 |  |  | 3.5719 | 0.14062 | 0.15625
[21 empty] | 95 | 0.0013583 | 0.0010583 | 0.0010583 | 0.0011667 | 0.0018333 | 0.00205 |  |  | 3.7703 | 0.14844 | 0.15625
[21 empty] | 100 | 0.0015 | 0.0012 | 0.0012 | 0.0013 | 0.002 | 0.0022 |  |  | 3.9688 | 0.15625 | 0.15625
[21 empty] | 105 | 0.00156 | 0.00126 | 0.00126 | 0.00134 | 0.00206 | 0.00226 |  |  | 4.1672 | 0.16406 | 0.1875
[21 empty] | 110 | 0.00162 | 0.00132 | 0.00132 | 0.00138 | 0.00212 | 0.00232 |  |  | 4.3656 | 0.17188 | 0.1875
[21 empty] | 115 | 0.00168 | 0.00138 | 0.00138 | 0.00142 | 0.00218 | 0.00238 |  |  | 4.5641 | 0.17969 | 0.1875
[21 empty] | 120 | 0.00174 | 0.00144 | 0.00144 | 0.00146 | 0.00224 | 0.00244 |  |  | 4.7625 | 0.1875 | 0.1875
[21 empty] | 125 | 0.0018 | 0.0015 | 0.0015 | 0.0015 | 0.0023 | 0.0025 |  |  | 4.9609 | 0.19531 | 0.21875
[21 empty] | 130 | 0.00184 | 0.00152 | 0.00152 | 0.00152 | 0.00232 | 0.00252 |  |  | 5.1594 | 0.20312 | 0.21875
[21 empty] | 135 | 0.00188 | 0.00154 | 0.00154 | 0.00154 | 0.00234 | 0.00254 |  |  | 5.3578 | 0.21094 | 0.21875
[21 empty] | 140 | 0.00192 | 0.00156 | 0.00156 | 0.00156 | 0.00236 | 0.00256 |  |  | 5.5563 | 0.21875 | 0.21875
[21 empty] | 145 | 0.00196 | 0.00158 | 0.00158 | 0.00158 | 0.00238 | 0.00258 |  |  | 5.7547 | 0.22656 | 0.25
[21 empty] | 150 | 0.002 | 0.0016 | 0.0016 | 0.0016 | 0.0024 | 0.0026 |  |  | 5.9531 | 0.23438 | 0.25
[23 empty] | gray cells are interpolated data | [6 empty] | 6.1516 | 0.24219 | 0.25
** Leave blank if the barrier is secondary only. |  |  |  | Total: | [25 empty] | 6.35 | 0.25 | >=1/4"
25.4
Shielding Requirements (mm) | [12 empty] | Min. Lead | Fitting Parameters for Diagnostic Shielding Calculations* | [15 empty] | Fitting Parameters for Diagnostic Shielding Calculations*
| Primary |  |  | Scatter |  |  | Leakage |  |  | Scatter and Leakage Combined |  |  | Shielding
Barrier | Lead | Concrete | Gypsum | Lead | Concrete | Gypsum | Lead | Concrete | Gypsum | Lead | Concrete | Gypsum | Required |  | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
[14 empty] | kVp | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | Ko**
[14 empty] | 25 | 49.52 | 194 | 0.3037 | 0.013997 | 0.046498 | 0.3904 | 1.645 | 0.2757 | 1.7755 | 5.898 | 0.1576 | 0.7175 | 0.3048 | 4.3981 | 14.6103 | 9.364 | 41.25 | 0.3202 | 0.074023 | 0.2459 | 0.3804 | 1.543 | 0.2869 | 1.8222 | 6.0531 | 0.0223 | 0.0434 | 0.1937 | 31.0828 | 103.2549 | 1.4835
[14 empty] | 30 | 38.8 | 178 | 0.3473 | 0.017865 | 0.059345 | 0.3173 | 1.698 | 0.3593 | 2.1845 | 7.2568 | 0.1208 | 0.7043 | 0.3613 | 5.738 | 19.0611 | 7.406 | 41.93 | 0.3959 | 0.093593 | 0.31091 | 0.3061 | 1.599 | 0.3693 | 2.2644 | 7.5223 | 0.02166 | 0.03966 | 0.2843 | 32.0013 | 106.3059 | 2.466
[14 empty] | 35 | 29.55 | 164.7 | 0.3948 | 0.023457 | 0.077922 | 0.2528 | 1.807 | 0.4648 | 2.7419 | 9.1083 | 0.08878 | 0.6988 | 0.4245 | 7.8075 | 25.9359 | 5.716 | 43.41 | 0.4857 | 0.12126 | 0.40283 | 0.2396 | 1.694 | 0.4683 | 2.8929 | 9.6101 | 0.01901 | 0.03873 | 0.3732 | 36.4622 | 121.1249 | 3.4485
[14 empty] | kVp: 40; Concrete (mm) a: 0.1297; Concrete (mm) b: 0.178; Concrete (mm) g: 0.2189; Concrete (mm) HVL: 5.3442; Concrete (mm) TVL: 17.7532 | 0.7189
[14 empty] | kVp: 45; Concrete (mm) a: 0.1095; Concrete (mm) b: 0.1741; Concrete (mm) g: 0.2269; Concrete (mm) HVL: 6.3301; Concrete (mm) TVL: 21.0282 | 0.87211
[14 empty] | 50 | 8.801 | 27.28 | 0.2957 | 0.078758 | 0.26163 | 0.0932 | 0.1712 | 0.2324 | 7.4372 | 24.7058 | 0.03883 | 0.0873 | 0.5105 | 17.8508 | 59.2991 | 1.817 | 4.84 | 0.4021 | 0.38148 | 1.2672 | 0.09721 | 0.1799 | 0.4912 | 7.1304 | 23.6867 | 0.01076 | 0.001862 | 1.17 | 64.4189 | 213.9949 | 1.0655
[14 empty] | 55 | 7.839 | 25.92 | 0.3499 | 0.088423 | 0.29373 | 0.07422 | 0.1697 | 0.2454 | 9.3391 | 31.0238 | 0.03419 | 0.08315 | 0.5606 | 20.2734 | 67.3467 | 1.493 | 4.515 | 0.4293 | 0.46426 | 1.5423 | 0.08552 | 0.1661 | 0.5112 | 8.1051 | 26.9245 | 0.01012 | 0.001404 | 1.269 | 68.4928 | 227.5282 | 1.2967
[14 empty] | 60 | 6.951 | 24.89 | 0.4198 | 0.099719 | 0.33126 | 0.06251 | 0.1692 | 0.2733 | 11.0886 | 36.8355 | 0.02985 | 0.07961 | 0.6169 | 23.221 | 77.1385 | 1.183 | 4.219 | 0.4571 | 0.58592 | 1.9464 | 0.07452 | 0.1539 | 0.5304 | 9.3015 | 30.8989 | 0.009512 | 0.0009672 | 1.333 | 72.8708 | 242.0716 | 1.5634
[14 empty] | 65 | 6.13 | 24.09 | 0.5019 | 0.11307 | 0.37563 | 0.05528 | 0.1696 | 0.3217 | 12.5388 | 41.6531 | 0.02609 | 0.07597 | 0.6756 | 26.5675 | 88.2555 | 0.9172 | 3.982 | 0.4922 | 0.75572 | 2.5105 | 0.06514 | 0.1443 | 0.5582 | 10.6409 | 35.3483 | 0.00899 | 0.000647 | 1.353 | 77.102 | 256.1274 | 1.8633
[14 empty] | 70 | 5.369 | 23.49 | 0.5881 | 0.1291 | 0.42887 | 0.05087 | 0.1696 | 0.3847 | 13.6259 | 45.2641 | 0.02302 | 0.07163 | 0.7299 | 30.1107 | 100.0254 | 0.7149 | 3.798 | 0.5378 | 0.96957 | 3.2208 | 0.05791 | 0.1357 | 0.5967 | 11.9694 | 39.7614 | 0.00855 | 0.000539 | 1.194 | 81.0698 | 269.3082 | 2.1939
[14 empty] | 75 | 4.666 | 22.69 | 0.6618 | 0.14855 | 0.49348 | 0.04797 | 0.1663 | 0.4492 | 14.4496 | 48.0005 | 0.02066 | 0.06649 | 0.775 | 33.5502 | 111.4514 | 0.5793 | 3.629 | 0.5908 | 1.1965 | 3.9748 | 0.05291 | 0.128 | 0.6478 | 13.1005 | 43.5189 | 0.005203 | 0.0006421 | 1.062 | 133.2207 | 442.5495 | 2.5529
[14 empty] | 80 | 4.04 | 21.69 | 0.7187 | 0.17157 | 0.56995 | 0.04583 | 0.1549 | 0.4926 | 15.1243 | 50.2419 | 0.01886 | 0.06093 | 0.8103 | 36.7522 | 122.0883 | 0.4921 | 3.428 | 0.6427 | 1.4085 | 4.6791 | 0.04955 | 0.1208 | 0.7097 | 13.9888 | 46.4699 | 0.007903 | 0.000864 | 0.9703 | 87.7068 | 291.3558 | 2.938
[14 empty] | 85 | 3.504 | 20.37 | 0.755 | 0.19782 | 0.65713 | 0.04398 | 0.1348 | 0.4943 | 15.7605 | 52.3553 | 0.01746 | 0.05558 | 0.8392 | 39.6992 | 131.8777 | 0.4355 | 3.178 | 0.6861 | 1.5916 | 5.2872 | 0.04721 | 0.114 | 0.7786 | 14.6822 | 48.7732 | 0.007686 | 0.001056 | 1.015 | 90.1831 | 299.5817 | 3.3468
[14 empty] | 90 | 3.067 | 18.83 | 0.7726 | 0.226 | 0.75076 | 0.04228 | 0.1137 | 0.469 | 16.3942 | 54.4604 | 0.01663 | 0.05039 | 0.8585 | 41.6805 | 138.4597 | 0.3971 | 2.913 | 0.7204 | 1.7455 | 5.7985 | 0.0455 | 0.1077 | 0.8522 | 15.234 | 50.6063 | 0.007511 | 0.001159 | 1.081 | 92.2843 | 306.5617 | 3.777
[14 empty] | 95 | 2.731 | 17.07 | 0.7714 | 0.25381 | 0.84313 | 0.04068 | 0.09705 | 0.4406 | 17.039 | 56.6024 | 0.01543 | 0.04571 | 0.8763 | 44.922 | 149.2278 | 0.3681 | 2.654 | 0.7461 | 1.883 | 6.2553 | 0.0441 | 0.1013 | 0.9222 | 15.7176 | 52.2128 | 0.007345 | 0.001133 | 1.116 | 94.3699 | 313.4901 | 4.2261
| Primary |  |  | Scatter |  |  | Leakage |  |  | Scatter and Leakage Combined |  |  |  | 100 | 2.5 | 15.28 | 0.7557 | 0.27726 | 0.92103 | 0.03925 | 0.08567 | 0.4273 | 17.6598 | 58.6646 | 0.01466 | 0.04171 | 0.8939 | 47.2815 | 157.0658 | 0.3415 | 2.42 | 0.7645 | 2.0297 | 6.7426 | 0.04278 | 0.09466 | 0.9791 | 16.2026 | 53.8239 | 0.00723 | 0.0009343 | 1.309 | 95.871 | 318.4765 | 4.692
Barrier | Steel | Glass | Wood | Steel | Glass | Wood | Steel | Glass | Wood | Steel | Glass | Wood |  | 105 | 2.364 | 13.41 | 0.7239 | 0.29321 | 0.97402 | 0.03808 | 0.07862 | 0.4394 | 18.2024 | 60.467 | 0.01397 | 0.03815 | 0.908 | 49.6168 | 164.8236 | 0.3135 | 2.227 | 0.7788 | 2.211 | 7.3448 | 0.04143 | 0.08751 | 1.014 | 16.7306 | 55.5777 | 0.00705 | 0.0006199 | 1.365 | 98.3187 | 326.6078 | 5.1722
[14 empty] | 110 | 2.296 | 11.7 | 0.6827 | 0.30189 | 1.0029 | 0.03715 | 0.07436 | 0.4752 | 18.6581 | 61.9808 | 0.01336 | 0.03521 | 0.9244 | 51.8823 | 172.3492 | 0.2849 | 2.061 | 0.7897 | 2.4329 | 8.0821 | 0.04008 | 0.08047 | 1.03 | 17.2941 | 57.4497 | 0.006921 | 0.0001976 | 3.309 | 100.1513 | 332.6954 | 5.6643
[14 empty] | 115 | 2.265 | 10.21 | 0.6363 | 0.30603 | 1.0166 | 0.03636 | 0.07201 | 0.5319 | 19.0635 | 63.3274 | 0.01283 | 0.06271 | 0.9423 | 54.0255 | 179.4688 | 0.2579 | 1.922 | 0.8008 | 2.6877 | 8.9282 | 0.03878 | 0.07394 | 1.033 | 17.8738 | 59.3756 | 0.006864 | -0.0003908 | 0.6469 | 100.983 | 335.4582 | 6.166
[14 empty] | 120 | 2.246 | 8.95 | 0.5873 | 0.30861 | 1.0252 | 0.03566 | 0.07109 | 0.6073 | 19.4377 | 64.5705 | 0.01235 | 0.03047 | 0.9566 | 56.1253 | 186.4441 | 0.2336 | 1.797 | 0.8116 | 2.9672 | 9.857 | 0.03758 | 0.06808 | 1.031 | 18.4446 | 61.2716 | 0.006726 | -0.0008308 | 1.006 | 103.0549 | 342.3409 | 6.675
[14 empty] | 125 | 2.219 | 7.923 | 0.5386 | 0.31237 | 1.0377 | 0.03502 | 0.07113 | 0.6974 | 19.7929 | 65.7506 | 0.01192 | 0.02863 | 0.9684 | 58.1499 | 193.1699 | 0.213 | 1.677 | 0.8217 | 3.2542 | 10.8103 | 0.03652 | 0.06304 | 1.031 | 18.9799 | 63.05 | 0.006584 | -0.001214 | 1.192 | 105.2775 | 349.7243 | 7.1889
[14 empty] | 130 | 2.17 | 7.094 | 0.4909 | 0.31942 | 1.0611 | 0.03445 | 0.0716 | 0.7969 | 20.1204 | 66.8385 | 0.01155 | 0.02702 | 0.9802 | 60.0127 | 199.358 | 0.1969 | 1.557 | 0.8309 | 3.5203 | 11.6942 | 0.03561 | 0.05874 | 1.037 | 19.465 | 64.6612 | 0.006472 | -0.001539 | 1.285 | 107.0994 | 355.7764 | 7.7053
[14 empty] | 135 | 2.102 | 6.45 | 0.4469 | 0.32976 | 1.0954 | 0.03394 | 0.07263 | 0.9099 | 20.4227 | 67.8428 | 0.01122 | 0.02561 | 0.9901 | 61.7778 | 205.2215 | 0.1838 | 1.44 | 0.8391 | 3.7712 | 12.5277 | 0.03481 | 0.05519 | 1.049 | 19.9123 | 66.1472 | 0.006306 | -0.001731 | 1.465 | 109.9187 | 365.1419 | 8.2219
[14 empty] | 140 | 2.009 | 5.916 | 0.4018 | 0.34502 | 1.1461 | 0.03345 | 0.07476 | 1.047 | 20.7219 | 68.8366 | 0.01088 | 0.02436 | 0.9964 | 63.7084 | 211.6347 | 0.1724 | 1.328 | 0.8458 | 4.0206 | 13.3561 | 0.03407 | 0.05145 | 1.057 | 20.3448 | 67.5839 | 0.006191 | -0.001849 | 1.53 | 111.9605 | 371.9246 | 8.7364
[14 empty] | 145 | 1.895 | 5.498 | 0.358 | 0.36578 | 1.2151 | 0.03296 | 0.07875 | 1.224 | 21.03 | 69.86 | 0.01056 | 0.02313 | 0.9987 | 65.6389 | 218.0478 | 0.1616 | 1.225 | 0.8519 | 4.2893 | 14.2487 | 0.03336 | 0.04795 | 1.063 | 20.7778 | 69.0223 | 0.006115 | -0.001869 | 1.498 | 113.352 | 376.547 | 9.2464
[14 empty] | 150 | 1.757 | 5.177 | 0.3156 | 0.39451 | 1.3105 | 0.03243 | 0.08599 | 1.467 | 21.3736 | 71.0017 | 0.0103 | 0.02198 | 1.013 | 67.2958 | 223.552 | 0.1501 | 1.132 | 0.8566 | 4.6179 | 15.3403 | 0.03266 | 0.04491 | 1.073 | 21.2231 | 70.5017 | 0.00602 | -0.001752 | 1.483 | 115.1407 | 382.4892 | 9.7495
* 25-35 kVp data is for Mo/Mo x-ray tubes, all other data is for W x-ray tubes (ref 8).
** Ko is the exposure at the point of interest with no attenuator.
The data in this column is based on Simpkin's output analysis.
Shielding Requirements (inches) | [12 empty] | Min. Lead
| Primary |  |  | Scatter |  |  | Leakage |  |  | Scatter and Leakage Combined |  |  | Shielding
Barrier | Lead | Concrete | Gypsum | Lead | Concrete | Gypsum | Lead | Concrete | Gypsum | Lead | Concrete | Gypsum | Required
| Primary |  |  | Scatter |  |  | Leakage |  |  | Scatter and Leakage Combined
Barrier | Steel | Glass | Wood | Steel | Glass | Wood | Steel | Glass | Wood | Steel | Glass | Wood
Note on K24: Eugene Mah: Dose limits per NCRP 116 recommendations.
Note on F25: Eugene Mah: Barrier use factor. Leave blank for secondary barriers.
Note on G25: Eugene Mah: Occupancy factor
Note on H25: Eugene Mah: Type of area U - Uncontrolled C - Controlled C assumed if blank
Note on K43: Eugene Mah: Max of either scatter or leakage. If Scatter-Leakage is less than 1 TVL, then an additional HVL is added to the min required thickness.

## RadFluoro
Location: | [13 empty] | Wall A/Door A | Leakage, side scatter RF Rad
Date: | [14 empty] | Leakage, side scatter RF Fluoro
Performed by: | Eugene Mah, M.Sc. | [13 empty] | Leakage, side scatter Rad Wall
[14 empty] | Wall B | Primary Rad Wall
| Primary |  | Secondary AK (mGy/pt) | [5 empty] | Primary RF (rad) | [5 empty] | Primary RF Rad xtable
| Wnorm | Kp | Wnorm | Leakage | Side Scatter | Forward/Back |  |  | a | b | g | b/a | ag |  | Leakage, side scatter RF Rad
Table | 1.8508 | 5.2 |  | 0.00014 | 0.023 | 0.033 |  | Pb | 2.295 | 13 | 0.5573 | 5.6645 | 1.279 |  | Leakage, side scatter RF Fluoro
Wall | 0.60091 | 2.3 |  | 0.00039 | 0.0049 | 0.0069 |  | Concrete | 0.03549 | 0.1164 | 0.5774 | 3.2798 | 0.020492 |  | Leakage, forward/back scatter RF Fluoro
All | 2.4525 |  |  | 0.00053 | 0.034 | 0.048 |  |  | Secondary RF (rad) |  |  |  |  | Wall C | Primary RF Rad xtable
RF Fluoro | 12.8934 |  | 12.8934 | 0.012 | 0.31 | 0.44 |  |  | a | b | g | b/a | ag |  | Leakage, side scatter RF Rad
RF Rad | 1.51 | 5.9 | 1.51 | 0.00094 | 0.028 | 0.039 |  | Pb | 2.272 | 13.6 | 0.7184 | 5.9859 | 1.6322 |  | Leakage, side scatter RF Fluoro
[8 empty] | Concrete | 0.0356 | 0.1114 | 0.662 | 3.1292 | 0.023567 |  | Leakage, side scatter Rad Wall
Fl Pts/wk | [8 empty] | Primary RF (fluoro) |  |  |  |  | Wall D/Door D | Leakage, side scatter RF Rad
Rad Pts/wk | [8 empty] | a | b | g | b/a | ag |  | Leakage, side scatter RF Fluoro
Mean DAP/pt |  | uGy-m^2 | [5 empty] | Pb | 2.347 | 12.67 | 0.6149 | 5.3984 | 1.4432 |  | Leakage, forward/back scatter RF Fluoro
[8 empty] | Concrete | 0.03616 | 0.09721 | 0.5186 | 2.6883 | 0.018753 |  | Leakage, forward/back scatter Rad Wall
[9 empty] | Secondary RF (fluoro) |  |  |  |  | Wall E/Door E | Leakage, side scatter RF Rad
[9 empty] | a | b | g | b/a | ag |  | Leakage, side scatter RF Fluoro
[8 empty] | Pb | 2.322 | 12.91 | 0.7575 | 5.5599 | 1.7589 |  | Leakage, forward/back scatter RF Fluoro
[8 empty] | Concrete | 0.0363 | 0.0936 | 0.5955 | 2.5785 | 0.021617 |  | Leakage, forward/back scatter Rad Wall
[14 empty] | Control | Leakage, side scatter RF Rad
Leakage, side scatter RF Fluoro
| Dtable | Dleak | U | P | T | Kp(0) | Bpri | x (mm) | Ksec(0) |  | Ksec(tot) | Bsec | x (mm) |  | Leakage, side scatter Rad Wall
Wall A | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | Floor | Primary RF Rad
Door A | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! |  | Leakage, side scatter Rad Wall
Wall B | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! |  | Leakage, forward/back scatter RF Fluoro
Wall C | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! |  | Leakage, side scatter RF Fluoro
Wall D | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | Ceiling | Leakage, forward/back scatter RF Rad
Door D | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! |  | Leakage, side scatter Rad Wall
Wall E | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! |  | Leakage, forward/back scatter RF Fluoro
Door E1 | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
Door E2 | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
Control Wall | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
Control Window | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
Floor | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
Ceiling | [8 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
| Dwall | Dleak | U | P | T | Kp(0) | Bpri | x (mm) | Ksec(0)
Wall A | [8 empty] | #DIV/0!
Door A | [8 empty] | #DIV/0!
Wall B | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Wall C | [8 empty] | #DIV/0!
Wall D | [8 empty] | #DIV/0!
Door D | [8 empty] | #DIV/0!
Wall E | [8 empty] | #DIV/0!
Door E1 | [8 empty] | #DIV/0!
Door E2 | [8 empty] | #DIV/0!
Control Wall | [8 empty] | #DIV/0!
Control Window | [8 empty] | #DIV/0!
Floor | [8 empty] | #DIV/0!
Ceiling | [8 empty] | #DIV/0!
| Dfluoro | Dleak | U | P | T | Kp(0) | Bpri | x (mm) | Ksec(0)
Wall A | [8 empty] | #DIV/0!
Door A | [8 empty] | #DIV/0!
Wall B | [8 empty] | #DIV/0!
Wall C | [8 empty] | #DIV/0!
Wall D | [8 empty] | #DIV/0!
Door D | [8 empty] | #DIV/0!
Wall E | [8 empty] | #DIV/0!
Door E1 | [8 empty] | #DIV/0!
Door E2 | [8 empty] | #DIV/0!
Control Wall | [8 empty] | #DIV/0!
Control Window | [8 empty] | #DIV/0!
Floor | [8 empty] | #DIV/0!
Ceiling | [8 empty] | #DIV/0!
[6 empty] | Lead | Concrete
| Rel AK | D (m) | AK (uGy) | Ksec(0) | Bsec | x (mm) | x (mm)
Wall A |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Door A |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Wall B |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Wall C |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Wall D |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Door D |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Wall E |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Door E1 |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Door E2 |  |  | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! |  | #VALUE!
Control Wall |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE!
Control Window |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Floor |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Ceiling |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Note on E6: Eugene Mah: Leakage (mGy/pt)
Note on F6: Eugene Mah: Side scatter (mGy/pt)
Note on G6: Eugene Mah: Forward/Back scatter (mGy/pt)

## CT
MUSC Shawn Jenkins Children's Hospital | [7 empty] | Design Method
Siemens template (attached) gives the scatter air kerma (uGy) at various distances per mAs rotation.
Charleston, SC 29425 | [7 empty] | Using the number of roations per year one can calculate the exposure at the isodose distance.
This value is inverse square law corrected to the barrier distance
Barrier transmission factor (B) is determined from the design limit corrected for occupancy
Shielding Design | [7 empty] | Required barrier thickness is determined using Eq A.2 from NCRP 147
Location
Interlocks and Warnings
RHB 4.11.2.4 Requires door interlocks to prevent the initiation of a scan when the CT room door is open.
Doors A will be interlocked
Uses | Percent Workload
RHB 4.11.2.3 Requires booth areas in the first scatter to be marked
There do not seem to be areas affected by 4.11.2.3 in this design
This will be verified on installation
Workload | [7 empty] | Design Limits | [6 empty] | Fit Parameters
| Patients per day |  |  | Rotations per week | 0 |  |  | Occupational |  | 5 | mGy/year |  |  |  | Lead (mm) |  |  | Concrete (mm)
| Rotations per patient |  |  | Rotations per year | 0 |  |  | Non Occupational |  | 1 | mGy/year |  |  | kVp | a | b | g | a | b | g
| Rotation collimation (cm) |  |  | Rotation-mAs/yr | 0 | [8 empty] | 100 | 2.507 | 15.33 | 0.9124 | 0.0395 | 0.0844 | 0.5191
| mAs per rotation |  |  | Workload (mA-min/wk) | 0 |  |  | Half Value Layer | [5 empty] | 120 | 2.37 | 11.609 | 0.82095 | 0.0373 | 0.0752 | 0.65115
| kV | [6 empty] | HVL (mm-lead) |  | 0.27 |  |  |  | 125 | 2.233 | 7.888 | 0.7295 | 0.0351 | 0.066 | 0.7832
[8 empty] | HVL (inches-concrete) |  | 0.7973 |  |  |  |  | Fit Parameters
Barrier Design | [14 empty] | Lead (mm) |  |  | Concrete (mm)
|  |  |  | Design |  |  | Calculated |  |  | Exposure |  |  |  | kVp | a | b | g | a | b | g
| Distance |  |  | Limit | Isocontour | Isocontour | Lead | Inv Sq to barrier | Design | at Barrier |  |  |  | 120 | 2.2486 | 8.61 | 0.77103 | 0.0351 | 0.066 | 0.7832
Designation | (m) | Area | Occupancy | mGy/yr | mGy/mAs | Dist (m) | (mm) |  | mGy/mAs | mGy/yr | Transmission |  | mm Pb | mm Conc
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[7 empty] | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Eugene Mah, M.Sc.
Certified Radiological Physicist
Note on Q18: Eugene Mah: Fit parameters from Fig A.2 NCRP 147. Linear interpolation is used to obtain fit parameters for 130 kVp
Note on Q24: Eugene Mah: Fit parameters for selected kV (C23)
Note on J27: Eugene Mah: Inverse square correction from isocontour distance to barrier distance
Note on K27: Eugene Mah: Isocontour exposure at barrier distance (mGy/mAs)
Note on C28: Eugene Mah: Type of area U - Uncontrolled C - Controlled C assumed if blank
Note on L28: Eugene Mah: Annual isocontour exposure at barrier
Note on M28: Eugene Mah: Barrier transmission factor (B) required to reduce barrier exposure to design limit.
Note on O28: Eugene Mah: Barrier thickness from Eq A2, NCRP 147

## ShieldEvaluation
|  |  | Shielding Evaluation |  |  |  |  | Revision 1.0.2.25032002 |  | Reference List
0 | [8 empty] | 1. | Rossi RP, Ritenour R, Christodoulou E; Broad Beam Transmission Properties ..., Health Physics 61(5) 601-608 (1991)
[9 empty] | 2. | Simpkin D; Shielding Requirements for Mammography, Health Physics 53(3) 267-279 (1987).
[9 empty] | 3. | NCRP49
Date |  |  | Survey ID: | [5 empty] | 4. | NCRP50
[9 empty] | 5. | Archer BR; Radiation Design Criteria for Diagnostic X-Ray Facilities, Health Physics, 44(5), 507-517, 1983
Equivalent Lead Measurements | [8 empty] | 6. | Simpkin D; A General Solution Shielding of Medical X- and Gamma-Rays by NCRP49 Methods; Health Physics 52(?) 431-436 (1987).
[9 empty] | 7. | Archer BR, Fewell TR, et al; Attenuation Properties of Diagnostic X-Ray Shielding Materials. Med Phy 21(9) 1499-1507 (1994).
Radiographic Room | [8 empty] | 8. | Simpkin DJ, Transmission Data for Shielding Diagnostic X-Ray Faciliites; Health Physics68(5), 704-709, 1995.
Min.
[9 empty] | Fitting Parameters for the above conditions (mm^-1) | [6 empty] | mm | inch | Pb
Test Exposure | Unit |  |  |  | SN: |  |  |  | Attenuator | a | b | g | HVL | TVL |  | -25.4 | -1 | 0"
[9 empty] | Lead | [6 empty] | 0 | 0 | <=1/256"
| Time |  |  |  | Distance | Exposure |  |  | Concrete | [6 empty] | 0.099219 | 0.0039062 | 0.0078125
| (msec) | kV | mA | mAs | (m) | (mGy) | (mGy/mAs) |  | Gypsum | [6 empty] | 0.19844 | 0.0078125 | 0.015625
[9 empty] | Steel | [6 empty] | 0.39688 | 0.015625 | 0.03125
[9 empty] | Glass | [6 empty] | 0.59531 | 0.023438 | 0.03125
|  |  | Incident | Transmitted | Transmission factor |  |  |  | Wood | [6 empty] | 0.79375 | 0.03125 | 0.03125
|  | Distance | Exposure | Exposure |  | Lead | Lead | [8 empty] | 0.99219 | 0.039062 | 0.0625
Barrier | mAs | (m) | (mGy) | (mGy) |  | (mm) | (in) |  | Barrier thicknesses | [6 empty] | 1.1906 | 0.046875 | 0.0625
[9 empty] | Trans | Lead | Concrete | Gypsum | Steel | Glass | Wood | 1.3891 | 0.054688 | 0.0625
[16 empty] | 1.5875 | 0.0625 | 0.0625
[16 empty] | 1.7859 | 0.070312 | 0.09375
[16 empty] | 1.9844 | 0.078125 | 0.09375
[16 empty] | 2.1828 | 0.085938 | 0.09375
[16 empty] | 2.3813 | 0.09375 | 0.09375
[16 empty] | 2.5797 | 0.10156 | 0.125
[16 empty] | 2.7781 | 0.10938 | 0.125
[16 empty] | 2.9766 | 0.11719 | 0.125
[16 empty] | 3.175 | 0.125 | 0.125
[16 empty] | 3.3734 | 0.13281 | 0.15625
[16 empty] | 3.5719 | 0.14062 | 0.15625
Eugene Mah, M.Sc. | [15 empty] | 3.7703 | 0.14844 | 0.15625
Medical Physicist | [15 empty] | 3.9688 | 0.15625 | 0.15625
[16 empty] | 4.1672 | 0.16406 | 0.1875
[16 empty] | 4.3656 | 0.17188 | 0.1875
[16 empty] | 4.5641 | 0.17969 | 0.1875
[16 empty] | 4.7625 | 0.1875 | 0.1875
[16 empty] | 4.9609 | 0.19531 | 0.21875
[16 empty] | 5.1594 | 0.20312 | 0.21875
[16 empty] | 5.3578 | 0.21094 | 0.21875
[16 empty] | 5.5563 | 0.21875 | 0.21875
[16 empty] | 5.7547 | 0.22656 | 0.25
[16 empty] | 5.9531 | 0.23438 | 0.25
[16 empty] | 6.1516 | 0.24219 | 0.25
[16 empty] | 6.35 | 0.25 | >=1/4"
25.4
Note on L11: Fit parameters are for barrier thickness equation from Ref 8 x = ln((B^-g+b/a)/(1+b/a))/ag
Note on F14: Source to detector distance for reference exposure
Note on G14: Reference exposure
Note on C15: kVp used for reference exposure
Note on D15: mA used for reference exposure
Note on E15: mAs uesd for reference exposure
Note on F18: Transmission = Transmitted exposure/Incident exposure
Note on C19: Source to barrier distance
Note on D19: Incident exposure at barrier
Note on E19: Transmitted exposure at barrier
Note on G19: Barrier lead equivalent (mm)
Note on L20: Barrier thickness is based on calculated transmission (Tran) and Equation 3 from Ref 8

## FitParameters
Fitting Parameters for Diagnostic Shielding Calculations* | [15 empty] | Fitting Parameters for Diagnostic Shielding Calculations*
| Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
kVp | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | a | b | g | HVL | TVL | Ko**
25 | 49.52 | 194 | 0.3037 | 0.013997 | 0.046498 | 0.3904 | 1.645 | 0.2757 | 1.7755 | 5.898 | 0.1576 | 0.7175 | 0.3048 | 4.3981 | 14.6103 | 9.364 | 41.25 | 0.3202 | 0.074023 | 0.2459 | 0.3804 | 1.543 | 0.2869 | 1.8222 | 6.0531 | 0.0223 | 0.0434 | 0.1937 | 31.0828 | 103.2549 | 1.4835
30 | 38.8 | 178 | 0.3473 | 0.017865 | 0.059345 | 0.3173 | 1.698 | 0.3593 | 2.1845 | 7.2568 | 0.1208 | 0.7043 | 0.3613 | 5.738 | 19.0611 | 7.406 | 41.93 | 0.3959 | 0.093593 | 0.31091 | 0.3061 | 1.599 | 0.3693 | 2.2644 | 7.5223 | 0.02166 | 0.03966 | 0.2843 | 32.0013 | 106.3059 | 2.466
35 | 29.55 | 164.7 | 0.3948 | 0.023457 | 0.077922 | 0.2528 | 1.807 | 0.4648 | 2.7419 | 9.1083 | 0.08878 | 0.6988 | 0.4245 | 7.8075 | 25.9359 | 5.716 | 43.41 | 0.4857 | 0.12126 | 0.40283 | 0.2396 | 1.694 | 0.4683 | 2.8929 | 9.6101 | 0.01901 | 0.03873 | 0.3732 | 36.4622 | 121.1249 | 3.4485
kVp: 40; Concrete (mm) a: 0.1297; Concrete (mm) b: 0.178; Concrete (mm) g: 0.2189; Concrete (mm) HVL: 5.3442; Concrete (mm) TVL: 17.7532 | 0.7189
kVp: 45; Concrete (mm) a: 0.1095; Concrete (mm) b: 0.1741; Concrete (mm) g: 0.2269; Concrete (mm) HVL: 6.3301; Concrete (mm) TVL: 21.0282 | 0.87211
50 | 8.801 | 27.28 | 0.2957 | 0.078758 | 0.26163 | 0.0932 | 0.1712 | 0.2324 | 7.4372 | 24.7058 | 0.03883 | 0.0873 | 0.5105 | 17.8508 | 59.2991 | 1.817 | 4.84 | 0.4021 | 0.38148 | 1.2672 | 0.09721 | 0.1799 | 0.4912 | 7.1304 | 23.6867 | 0.01076 | 0.001862 | 1.17 | 64.4189 | 213.9949 | 1.0655
55 | 7.839 | 25.92 | 0.3499 | 0.088423 | 0.29373 | 0.07422 | 0.1697 | 0.2454 | 9.3391 | 31.0238 | 0.03419 | 0.08315 | 0.5606 | 20.2734 | 67.3467 | 1.493 | 4.515 | 0.4293 | 0.46426 | 1.5423 | 0.08552 | 0.1661 | 0.5112 | 8.1051 | 26.9245 | 0.01012 | 0.001404 | 1.269 | 68.4928 | 227.5282 | 1.2967
60 | 6.951 | 24.89 | 0.4198 | 0.099719 | 0.33126 | 0.06251 | 0.1692 | 0.2733 | 11.0886 | 36.8355 | 0.02985 | 0.07961 | 0.6169 | 23.221 | 77.1385 | 1.183 | 4.219 | 0.4571 | 0.58592 | 1.9464 | 0.07452 | 0.1539 | 0.5304 | 9.3015 | 30.8989 | 0.009512 | 0.0009672 | 1.333 | 72.8708 | 242.0716 | 1.5634
65 | 6.13 | 24.09 | 0.5019 | 0.11307 | 0.37563 | 0.05528 | 0.1696 | 0.3217 | 12.5388 | 41.6531 | 0.02609 | 0.07597 | 0.6756 | 26.5675 | 88.2555 | 0.9172 | 3.982 | 0.4922 | 0.75572 | 2.5105 | 0.06514 | 0.1443 | 0.5582 | 10.6409 | 35.3483 | 0.00899 | 0.000647 | 1.353 | 77.102 | 256.1274 | 1.8633
70 | 5.369 | 23.49 | 0.5881 | 0.1291 | 0.42887 | 0.05087 | 0.1696 | 0.3847 | 13.6259 | 45.2641 | 0.02302 | 0.07163 | 0.7299 | 30.1107 | 100.0254 | 0.7149 | 3.798 | 0.5378 | 0.96957 | 3.2208 | 0.05791 | 0.1357 | 0.5967 | 11.9694 | 39.7614 | 0.00855 | 0.000539 | 1.194 | 81.0698 | 269.3082 | 2.1939
75 | 4.666 | 22.69 | 0.6618 | 0.14855 | 0.49348 | 0.04797 | 0.1663 | 0.4492 | 14.4496 | 48.0005 | 0.02066 | 0.06649 | 0.775 | 33.5502 | 111.4514 | 0.5793 | 3.629 | 0.5908 | 1.1965 | 3.9748 | 0.05291 | 0.128 | 0.6478 | 13.1005 | 43.5189 | 0.005203 | 0.0006421 | 1.062 | 133.2207 | 442.5495 | 2.5529
80 | 4.04 | 21.69 | 0.7187 | 0.17157 | 0.56995 | 0.04583 | 0.1549 | 0.4926 | 15.1243 | 50.2419 | 0.01886 | 0.06093 | 0.8103 | 36.7522 | 122.0883 | 0.4921 | 3.428 | 0.6427 | 1.4085 | 4.6791 | 0.04955 | 0.1208 | 0.7097 | 13.9888 | 46.4699 | 0.007903 | 0.000864 | 0.9703 | 87.7068 | 291.3558 | 2.938
85 | 3.504 | 20.37 | 0.755 | 0.19782 | 0.65713 | 0.04398 | 0.1348 | 0.4943 | 15.7605 | 52.3553 | 0.01746 | 0.05558 | 0.8392 | 39.6992 | 131.8777 | 0.4355 | 3.178 | 0.6861 | 1.5916 | 5.2872 | 0.04721 | 0.114 | 0.7786 | 14.6822 | 48.7732 | 0.007686 | 0.001056 | 1.015 | 90.1831 | 299.5817 | 3.3468
90 | 3.067 | 18.83 | 0.7726 | 0.226 | 0.75076 | 0.04228 | 0.1137 | 0.469 | 16.3942 | 54.4604 | 0.01663 | 0.05039 | 0.8585 | 41.6805 | 138.4597 | 0.3971 | 2.913 | 0.7204 | 1.7455 | 5.7985 | 0.0455 | 0.1077 | 0.8522 | 15.234 | 50.6063 | 0.007511 | 0.001159 | 1.081 | 92.2843 | 306.5617 | 3.777
95 | 2.731 | 17.07 | 0.7714 | 0.25381 | 0.84313 | 0.04068 | 0.09705 | 0.4406 | 17.039 | 56.6024 | 0.01543 | 0.04571 | 0.8763 | 44.922 | 149.2278 | 0.3681 | 2.654 | 0.7461 | 1.883 | 6.2553 | 0.0441 | 0.1013 | 0.9222 | 15.7176 | 52.2128 | 0.007345 | 0.001133 | 1.116 | 94.3699 | 313.4901 | 4.2261
100 | 2.5 | 15.28 | 0.7557 | 0.27726 | 0.92103 | 0.03925 | 0.08567 | 0.4273 | 17.6598 | 58.6646 | 0.01466 | 0.04171 | 0.8939 | 47.2815 | 157.0658 | 0.3415 | 2.42 | 0.7645 | 2.0297 | 6.7426 | 0.04278 | 0.09466 | 0.9791 | 16.2026 | 53.8239 | 0.00723 | 0.0009343 | 1.309 | 95.871 | 318.4765 | 4.692
105 | 2.364 | 13.41 | 0.7239 | 0.29321 | 0.97402 | 0.03808 | 0.07862 | 0.4394 | 18.2024 | 60.467 | 0.01397 | 0.03815 | 0.908 | 49.6168 | 164.8236 | 0.3135 | 2.227 | 0.7788 | 2.211 | 7.3448 | 0.04143 | 0.08751 | 1.014 | 16.7306 | 55.5777 | 0.00705 | 0.0006199 | 1.365 | 98.3187 | 326.6078 | 5.1722
110 | 2.296 | 11.7 | 0.6827 | 0.30189 | 1.0029 | 0.03715 | 0.07436 | 0.4752 | 18.6581 | 61.9808 | 0.01336 | 0.03521 | 0.9244 | 51.8823 | 172.3492 | 0.2849 | 2.061 | 0.7897 | 2.4329 | 8.0821 | 0.04008 | 0.08047 | 1.03 | 17.2941 | 57.4497 | 0.006921 | 0.0001976 | 3.309 | 100.1513 | 332.6954 | 5.6643
115 | 2.265 | 10.21 | 0.6363 | 0.30603 | 1.0166 | 0.03636 | 0.07201 | 0.5319 | 19.0635 | 63.3274 | 0.01283 | 0.06271 | 0.9423 | 54.0255 | 179.4688 | 0.2579 | 1.922 | 0.8008 | 2.6877 | 8.9282 | 0.03878 | 0.07394 | 1.033 | 17.8738 | 59.3756 | 0.006864 | -0.0003908 | 0.6469 | 100.983 | 335.4582 | 6.166
120 | 2.246 | 8.95 | 0.5873 | 0.30861 | 1.0252 | 0.03566 | 0.07109 | 0.6073 | 19.4377 | 64.5705 | 0.01235 | 0.03047 | 0.9566 | 56.1253 | 186.4441 | 0.2336 | 1.797 | 0.8116 | 2.9672 | 9.857 | 0.03758 | 0.06808 | 1.031 | 18.4446 | 61.2716 | 0.006726 | -0.0008308 | 1.006 | 103.0549 | 342.3409 | 6.675
125 | 2.219 | 7.923 | 0.5386 | 0.31237 | 1.0377 | 0.03502 | 0.07113 | 0.6974 | 19.7929 | 65.7506 | 0.01192 | 0.02863 | 0.9684 | 58.1499 | 193.1699 | 0.213 | 1.677 | 0.8217 | 3.2542 | 10.8103 | 0.03652 | 0.06304 | 1.031 | 18.9799 | 63.05 | 0.006584 | -0.001214 | 1.192 | 105.2775 | 349.7243 | 7.1889
130 | 2.17 | 7.094 | 0.4909 | 0.31942 | 1.0611 | 0.03445 | 0.0716 | 0.7969 | 20.1204 | 66.8385 | 0.01155 | 0.02702 | 0.9802 | 60.0127 | 199.358 | 0.1969 | 1.557 | 0.8309 | 3.5203 | 11.6942 | 0.03561 | 0.05874 | 1.037 | 19.465 | 64.6612 | 0.006472 | -0.001539 | 1.285 | 107.0994 | 355.7764 | 7.7053
135 | 2.102 | 6.45 | 0.4469 | 0.32976 | 1.0954 | 0.03394 | 0.07263 | 0.9099 | 20.4227 | 67.8428 | 0.01122 | 0.02561 | 0.9901 | 61.7778 | 205.2215 | 0.1838 | 1.44 | 0.8391 | 3.7712 | 12.5277 | 0.03481 | 0.05519 | 1.049 | 19.9123 | 66.1472 | 0.006306 | -0.001731 | 1.465 | 109.9187 | 365.1419 | 8.2219
140 | 2.009 | 5.916 | 0.4018 | 0.34502 | 1.1461 | 0.03345 | 0.07476 | 1.047 | 20.7219 | 68.8366 | 0.01088 | 0.02436 | 0.9964 | 63.7084 | 211.6347 | 0.1724 | 1.328 | 0.8458 | 4.0206 | 13.3561 | 0.03407 | 0.05145 | 1.057 | 20.3448 | 67.5839 | 0.006191 | -0.001849 | 1.53 | 111.9605 | 371.9246 | 8.7364
145 | 1.895 | 5.498 | 0.358 | 0.36578 | 1.2151 | 0.03296 | 0.07875 | 1.224 | 21.03 | 69.86 | 0.01056 | 0.02313 | 0.9987 | 65.6389 | 218.0478 | 0.1616 | 1.225 | 0.8519 | 4.2893 | 14.2487 | 0.03336 | 0.04795 | 1.063 | 20.7778 | 69.0223 | 0.006115 | -0.001869 | 1.498 | 113.352 | 376.547 | 9.2464
150 | 1.757 | 5.177 | 0.3156 | 0.39451 | 1.3105 | 0.03243 | 0.08599 | 1.467 | 21.3736 | 71.0017 | 0.0103 | 0.02198 | 1.013 | 67.2958 | 223.552 | 0.1501 | 1.132 | 0.8566 | 4.6179 | 15.3403 | 0.03266 | 0.04491 | 1.073 | 21.2231 | 70.5017 | 0.00602 | -0.001752 | 1.483 | 115.1407 | 382.4892 | 9.7495

## NCRP147_4.1
Suggested occupancy factors (for use as a guide in planning shielding where other occupancy data are not available)
Location | Occupancy (T)
Administrative or clerical offices; laboratories, pharmacies and other work areas fully occupied by an individual; receptionist areas, attended waiting rooms, children's indoor play areas, adjacent x-ray rooms, film reading areas, nurse's stations, x-ray control rooms | 1
Rooms used for patient examinations and treatments | 0.5
Corridors, patient rooms, employee lounges, staff rest rooms | 0.2
Corridor doors | 0.125
Public toilets, unattended vending areas, storage rooms, outdoor areas with seating, unattended waiting rooms, patient holding areas | 0.05
Outdoor areas with only transient pedestrian or vehicular traffic, unattended parking lots, vehicular drop off areas (unattended), attics, stairways, unattended elevators, janitor's closets | 0.025

## NCRP147_4.2
Operating potential (kVp) distribution of workload (mA-min) normalized per patient, from survey conducted by AAPM TG9 (Simpkin, 1996a)
| Radiography room |  |  | Rad/Fluoro
kVp | Rad room (all barriers) | Rad room (chest bucky) | Rad room (floor or other) | Fluoro tube (RF) | Rad tube (RF) | Chest room | Mammo room | Cardiac | Angio
25 | 0 | 0 | 0 | 0 | 0 | 0 | 0.925 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 4.67 | 0 | 0
35 | 0 | 0 | 0 | 0 | 0 | 0 | 1.1 | 0 | 0
40 | 0.000138 | 0 | 0.000138 | 0 | 0 | 0 | 0 | 0 | 0
45 | 0.00071 | 0 | 0.00071 | 0 | 0.000578 | 0 | 0 | 0 | 0
50 | 0.00848 | 0.00678 | 0.0017 | 0 | 0.000765 | 0 | 0 | 0.34 | 0.0894
55 | 0.0109 | 0.000456 | 0.0104 | 0.0702 | 0.000726 | 0 | 0 | 0.42 | 0.0398
60 | 0.0981 | 0.00896 | 0.0891 | 0.113 | 0.0152 | 0 | 0 | 1.96 | 0.699
65 | 0.104 | 0.0342 | 0.07 | 0.187 | 0.0252 | 0 | 0 | 4.55 | 15
70 | 0.458 | 0.0725 | 0.385 | 0.145 | 0.0889 | 0.0202 | 0 | 6.03 | 12.2
75 | 0.501 | 0.0953 | 0.405 | 0.194 | 0.224 | 0.00236 | 0 | 8.02 | 15.3
80 | 0.56 | 0.14 | 0.42 | 1.72 | 0.428 | 0 | 0 | 25.4 | 11
85 | 0.315 | 0.0662 | 0.249 | 2.19 | 0.218 | 0.000783 | 0 | 40.3 | 4.09
90 | 0.176 | 0.0141 | 0.162 | 1.46 | 0.0533 | 0 | 0 | 21 | 3.43
95 | 0.0218 | 0.00351 | 0.0182 | 1.15 | 0.0489 | 0 | 0 | 10.6 | 0.673
100 | 0.0155 | 0.000884 | 0.0146 | 1.12 | 0.0587 | 0.0301 | 0 | 7.4 | 1.53
105 | 0.00348 | 0.00197 | 0.00151 | 0.964 | 0.0105 | 0 | 0 | 7.02 | 0.0927
110 | 0.0105 | 0.00991 | 0.000551 | 0.747 | 0.0646 | 0.0214 | 0 | 6.59 | 0.0305
115 | 0.041 | 0.0374 | 0.00369 | 1.44 | 0.029 | 0.0936 | 0 | 13.8 | 0
120 | 0.0699 | 0.0512 | 0.0187 | 0.937 | 0.104 | 0.0474 | 0 | 3.35 | 0
125 | 0.0484 | 0.0481 | 0.000347 | 0.138 | 0.0813 | 0 | 0 | 2.75 | 0
130 | 0.00184 | 0.00171 | 0.000125 | 0.153 | 0.0446 | 0 | 0 | 0.031 | 0
135 | 0.00773 | 0.00773 | 0 | 0.146 | 0.00947 | 0 | 0 | 0 | 0
140 | 0 | 0 | 0 | 0.0192 | 0.00426 | 0 | 0 | 0 | 0
Tot Workload | 2.4525 | 0.60091 | 1.8508 | 12.8934 | 1.51 | 0.21584 | 6.695 | 159.561 | 64.1744
Pt/week | 110 |  |  | 18 | 23 | 210 | 47 | 19 | 21

## NCRP147_4.3
Estimated total workloads in various medical x-ray imaging installations in clinics and hospitals.
The total workload values are for general guidance and are to be used only if the actual workloads are not available.
|  | Typical # of patients (N) per 40h week |  | W_tot (mA-min/week)
Room Type | W_norm (mA-min/pt) | Average | Busy | Average | Busy
Rad room (Chest bucky) | 0.60091 | 120 | 160 | 72.1092 | 96.1456
Rad room (floor or other) | 1.8508 | 120 | 160 | 72.1092 | 96.1456
Chest room | 0.21584 | 200 | 400 | 120.182 | 240.364
Fluoro tube (RF) | 12.8934 | 20 | 30 | 12.0182 | 18.0273
Rad tube (RF) | 1.51 | 25 | 40 | 15.0228 | 24.0364
Mammo | 6.695 | 80 | 160 | 48.0728 | 96.1456
Cardiac | 159.561 | 20 | 30 | 12.0182 | 18.0273
Angio | 64.1744 | 20 | 30 | 12.0182 | 18.0273

## NCRP147_4.4
Primary beam use factors (U) for a general radiographic room determined from the survey of clinical sites (Simpkin, 1996a)
Barrier | Use factor (U)
Floor | 0.89
Cross-table wall | 0.09
Wall #3 | 0.02
Chest image reecptor | 1

## NCRP147_4.5
Unshielded primary air kerma per patient [K1_p (in mGy/pt) for the indicated workload [Wnorm (mA-min/pt)] and workload distribution, normalized to primary beam distance d_p = 1m
Workload Distribution | Wnorm (mA-min/pt) | K1_p (mGy/pt)
Rad room (Chest bucky) | 0.6 | 2.3
Rad room (floor or other) | 1.9 | 5.2
Rad tube (RF) | 1.5 | 5.9
Chest room | 0.22 | 1.2

## NCRP147_4.7
Unshielded leakage, scattered and total secondary air kermas (in mGy/pt) for the indicated workload distribution at d_s = d_L = 1m
The workload distributions and total workloads per patient (Wnorm) for the indicated clinical sites are the average per patient surveyed by AAPM TG9 (Simpkin, 1996a)
The primary field size F (in cm^2) is known at primary distance d_F. Side scattered radiation is calculated for 90 degree scatter.
Forward and backscattered radiations are calculated for 135 degree scatter.
Leakage radiation technique factors are 150 kVp at 3.3 mA to achieve 0.876 mGy/h for all tubes except mammography, which assumes leakage radiation technique factors of 50 kVp at 5 mA
Unshielded Air Kerma (mGy/pt) at 1m
Workload Distribution | Wnorm (mAmin/pt) | F (cm^2) | d_F (m) | Leakage | Side scatter | Leakage and side scatter | Forward/Backscatter | Lekage and forward/back scatter)
Rad room (all barriers) | 2.4525 | 1000 | 1 | 0.00053 | 0.034 | 0.03453 | 0.048 | 0.04853
Rad room (chest bucky) | 0.60091 | 1535 | 1.83 | 0.00039 | 0.0049 | 0.00529 | 0.0069 | 0.00729
Rad room (floor or other) | 1.8508 | 1000 | 1 | 0.00014 | 0.023 | 0.02314 | 0.033 | 0.03314
Fluoro tube (RF) | 12.8934 | 730 | 0.8 | 0.012 | 0.31 | 0.322 | 0.44 | 0.452
Rad tube (RF) | 1.51 | 1000 | 1 | 0.00094 | 0.028 | 0.02894 | 0.039 | 0.03994
Chest room | 0.21584 | 1535 | 2 | 0.00038 | 0.0023 | 0.00268 | 0.0032 | 0.00358
Mammo room | 6.695 | 720 | 0.58 | 1.1e-05 | 0.011 | 0.011011 | 0.049 | 0.049011
Cardiac | 159.561 | 730 | 0.9 | 0.088 | 2.6 | 2.688 | 3.7 | 3.788
Angio | 64.1744 | 730 | 0.9 | 0.0034 | 0.66 | 0.6634 | 0.95 | 0.9534

## NCRP147_5.2
Currently suggested defalut DLP values per procedure.
For use as a guide in planning shielding in cases where facility-specific DLP values are not available
Procedure | CTDIvol (mGy) | Scan length (cm) | DLP (mGy-cm)
Head | 60 | 20 | 1200
Chest | 15 | 35 | 525
Abdomen | 25 | 25 | 625
Pelvis | 25 | 20 | 500
Body avg |  |  | 550

## NCRP147_A1
Fits of transmission for broad primary x-ray beams (for lead, concrete, gypsum wall board, steel, plate glass, and wood) to Equation A.2 (thickness x is input in mm)
| Pb (mm) | [6 empty] | Concrete (mm) | [6 empty] | Gypsum | [6 empty] | Steel | [6 empty] | Plate Glass | [6 empty] | Wood
kVp | a | b | g | b/a | ag | HVL | TVL | a | b | g | b/a | ag | HVL | TVL | a | b | g | b/a | ag | HVL | TVL | a | b | g | b/a | ag | HVL | TVL | a | b | g | b/a | ag | HVL | TVL | a | b | g | b/a | ag | HVL | TVL
25 | 49.52 | 194 | 0.3037 | 3.9176 | 15.0392 | 0.013997 | 0.046498 | 0.3904 | 1.645 | 0.2757 | 4.2136 | 0.10763 | 1.7755 | 5.898 | 0.1576 | 0.7175 | 0.3048 | 4.5527 | 0.048036 | 4.3981 | 14.6103 | 9.364 | 41.25 | 0.3202 | 4.4052 | 2.9984 | 0.074023 | 0.2459 | 0.3804 | 1.543 | 0.2869 | 4.0563 | 0.10914 | 1.8222 | 6.0531 | 0.0223 | 0.0434 | 0.1937 | 1.9462 | 0.0043195 | 31.0828 | 103.2549
30 | 38.8 | 178 | 0.3473 | 4.5876 | 13.4752 | 0.017865 | 0.059345 | 0.3173 | 1.698 | 0.3593 | 5.3514 | 0.11401 | 2.1845 | 7.2568 | 0.1208 | 0.7043 | 0.3613 | 5.8303 | 0.043645 | 5.738 | 19.0611 | 7.406 | 41.93 | 0.3959 | 5.6616 | 2.932 | 0.093593 | 0.31091 | 0.3061 | 1.599 | 0.3693 | 5.2238 | 0.11304 | 2.2644 | 7.5223 | 0.02166 | 0.03966 | 0.2843 | 1.831 | 0.0061579 | 32.0013 | 106.3059
35 | 29.55 | 164.7 | 0.3948 | 5.5736 | 11.6663 | 0.023457 | 0.077922 | 0.2528 | 1.807 | 0.4648 | 7.1479 | 0.1175 | 2.7419 | 9.1083 | 0.08878 | 0.6988 | 0.4245 | 7.8711 | 0.037687 | 7.8075 | 25.9359 | 5.716 | 43.41 | 0.4857 | 7.5945 | 2.7763 | 0.12126 | 0.40283 | 0.2396 | 1.694 | 0.4683 | 7.0701 | 0.1122 | 2.8929 | 9.6101 | 0.01901 | 0.03873 | 0.3732 | 2.0373 | 0.0070945 | 36.4622 | 121.1249
40 | [7 empty] | 0.1297 | 0.178 | 0.2189 | 1.3724 | 0.028391 | 5.3442 | 17.7532
45 | [7 empty] | 0.1095 | 0.1741 | 0.2269 | 1.59 | 0.024846 | 6.3301 | 21.0282
50 | 8.801 | 27.28 | 0.2957 | 3.0996 | 2.6025 | 0.078758 | 0.26163 | 0.0932 | 0.1712 | 0.2324 | 1.8369 | 0.02166 | 7.4372 | 24.7058 | 0.03883 | 0.0873 | 0.5105 | 2.2483 | 0.019823 | 17.8508 | 59.2991 | 1.817 | 4.84 | 0.4021 | 2.6637 | 0.73062 | 0.38148 | 1.2672 | 0.09721 | 0.1799 | 0.4912 | 1.8506 | 0.04775 | 7.1304 | 23.6867 | 0.01076 | 0.001862 | 1.17 | 0.17305 | 0.012589 | 64.4189 | 213.9949
55 | 7.839 | 25.92 | 0.3499 | 3.3065 | 2.7429 | 0.088423 | 0.29373 | 0.07422 | 0.1697 | 0.2454 | 2.2864 | 0.018214 | 9.3391 | 31.0238 | 0.03419 | 0.08315 | 0.5606 | 2.432 | 0.019167 | 20.2734 | 67.3467 | 1.493 | 4.515 | 0.4293 | 3.0241 | 0.64094 | 0.46426 | 1.5423 | 0.08552 | 0.1661 | 0.5112 | 1.9422 | 0.043718 | 8.1051 | 26.9245 | 0.01012 | 0.001404 | 1.269 | 0.13874 | 0.012842 | 68.4928 | 227.5282
60 | 6.951 | 24.89 | 0.4198 | 3.5808 | 2.918 | 0.099719 | 0.33126 | 0.06251 | 0.1692 | 0.2733 | 2.7068 | 0.017084 | 11.0886 | 36.8355 | 0.02985 | 0.07961 | 0.6169 | 2.667 | 0.018414 | 23.221 | 77.1385 | 1.183 | 4.219 | 0.4571 | 3.5664 | 0.54075 | 0.58592 | 1.9464 | 0.07452 | 0.1539 | 0.5304 | 2.0652 | 0.039525 | 9.3015 | 30.8989 | 0.009512 | 0.0009672 | 1.333 | 0.10168 | 0.012679 | 72.8708 | 242.0716
65 | 6.13 | 24.09 | 0.5019 | 3.9299 | 3.0766 | 0.11307 | 0.37563 | 0.05528 | 0.1696 | 0.3217 | 3.068 | 0.017784 | 12.5388 | 41.6531 | 0.02609 | 0.07597 | 0.6756 | 2.9118 | 0.017626 | 26.5675 | 88.2555 | 0.9172 | 3.982 | 0.4922 | 4.3415 | 0.45145 | 0.75572 | 2.5105 | 0.06514 | 0.1443 | 0.5582 | 2.2152 | 0.036361 | 10.6409 | 35.3483 | 0.00899 | 0.000647 | 1.353 | 0.071969 | 0.012163 | 77.102 | 256.1274
70 | 5.369 | 23.49 | 0.5881 | 4.3751 | 3.1575 | 0.1291 | 0.42887 | 0.05087 | 0.1696 | 0.3847 | 3.334 | 0.01957 | 13.6259 | 45.2641 | 0.02302 | 0.07163 | 0.7299 | 3.1116 | 0.016802 | 30.1107 | 100.0254 | 0.7149 | 3.798 | 0.5378 | 5.3126 | 0.38447 | 0.96957 | 3.2208 | 0.05791 | 0.1357 | 0.5967 | 2.3433 | 0.034555 | 11.9694 | 39.7614 | 0.00855 | 0.000539 | 1.194 | 0.063041 | 0.010209 | 81.0698 | 269.3082
75 | 4.666 | 22.69 | 0.6618 | 4.8628 | 3.088 | 0.14855 | 0.49348 | 0.04797 | 0.1663 | 0.4492 | 3.4668 | 0.021548 | 14.4496 | 48.0005 | 0.02066 | 0.06649 | 0.775 | 3.2183 | 0.016012 | 33.5502 | 111.4514 | 0.5793 | 3.629 | 0.5908 | 6.2645 | 0.34225 | 1.1965 | 3.9748 | 0.05291 | 0.128 | 0.6478 | 2.4192 | 0.034275 | 13.1005 | 43.5189 | 0.005203 | 0.0006421 | 1.062 | 0.12341 | 0.0055256 | 133.2207 | 442.5495
80 | 4.04 | 21.69 | 0.7187 | 5.3688 | 2.9035 | 0.17157 | 0.56995 | 0.04583 | 0.1549 | 0.4926 | 3.3799 | 0.022576 | 15.1243 | 50.2419 | 0.01886 | 0.06093 | 0.8103 | 3.2306 | 0.015282 | 36.7522 | 122.0883 | 0.4921 | 3.428 | 0.6427 | 6.9661 | 0.31627 | 1.4085 | 4.6791 | 0.04955 | 0.1208 | 0.7097 | 2.4379 | 0.035166 | 13.9888 | 46.4699 | 0.007903 | 0.000864 | 0.9703 | 0.10933 | 0.0076683 | 87.7068 | 291.3558
85 | 3.504 | 20.37 | 0.755 | 5.8134 | 2.6455 | 0.19782 | 0.65713 | 0.04398 | 0.1348 | 0.4943 | 3.065 | 0.021739 | 15.7605 | 52.3553 | 0.01746 | 0.05558 | 0.8392 | 3.1833 | 0.014652 | 39.6992 | 131.8777 | 0.4355 | 3.178 | 0.6861 | 7.2974 | 0.2988 | 1.5916 | 5.2872 | 0.04721 | 0.114 | 0.7786 | 2.4147 | 0.036758 | 14.6822 | 48.7732 | 0.007686 | 0.001056 | 1.015 | 0.13739 | 0.0078013 | 90.1831 | 299.5817
90 | 3.067 | 18.83 | 0.7726 | 6.1396 | 2.3696 | 0.226 | 0.75076 | 0.04228 | 0.1137 | 0.469 | 2.6892 | 0.019829 | 16.3942 | 54.4604 | 0.01663 | 0.05039 | 0.8585 | 3.0301 | 0.014277 | 41.6805 | 138.4597 | 0.3971 | 2.913 | 0.7204 | 7.3357 | 0.28607 | 1.7455 | 5.7985 | 0.0455 | 0.1077 | 0.8522 | 2.367 | 0.038775 | 15.234 | 50.6063 | 0.007511 | 0.001159 | 1.081 | 0.15431 | 0.0081194 | 92.2843 | 306.5617
95 | 2.731 | 17.07 | 0.7714 | 6.2505 | 2.1067 | 0.25381 | 0.84313 | 0.04068 | 0.09705 | 0.4406 | 2.3857 | 0.017924 | 17.039 | 56.6024 | 0.01543 | 0.04571 | 0.8763 | 2.9624 | 0.013521 | 44.922 | 149.2278 | 0.3681 | 2.654 | 0.7461 | 7.21 | 0.27464 | 1.883 | 6.2553 | 0.0441 | 0.1013 | 0.9222 | 2.2971 | 0.040669 | 15.7176 | 52.2128 | 0.007345 | 0.001133 | 1.116 | 0.15425 | 0.008197 | 94.3699 | 313.4901
100 | 2.5 | 15.28 | 0.7557 | 6.112 | 1.8893 | 0.27726 | 0.92103 | 0.03925 | 0.08567 | 0.4273 | 2.1827 | 0.016772 | 17.6598 | 58.6646 | 0.01466 | 0.04171 | 0.8939 | 2.8452 | 0.013105 | 47.2815 | 157.0658 | 0.3415 | 2.42 | 0.7645 | 7.0864 | 0.26108 | 2.0297 | 6.7426 | 0.04278 | 0.09466 | 0.9791 | 2.2127 | 0.041886 | 16.2026 | 53.8239 | 0.00723 | 0.0009343 | 1.309 | 0.12923 | 0.0094641 | 95.871 | 318.4765
105 | 2.364 | 13.41 | 0.7239 | 5.6726 | 1.7113 | 0.29321 | 0.97402 | 0.03808 | 0.07862 | 0.4394 | 2.0646 | 0.016732 | 18.2024 | 60.467 | 0.01397 | 0.03815 | 0.908 | 2.7309 | 0.012685 | 49.6168 | 164.8236 | 0.3135 | 2.227 | 0.7788 | 7.1037 | 0.24415 | 2.211 | 7.3448 | 0.04143 | 0.08751 | 1.014 | 2.1122 | 0.04201 | 16.7306 | 55.5777 | 0.00705 | 0.0006199 | 1.365 | 0.087929 | 0.0096232 | 98.3187 | 326.6078
110 | 2.296 | 11.7 | 0.6827 | 5.0958 | 1.5675 | 0.30189 | 1.0029 | 0.03715 | 0.07436 | 0.4752 | 2.0016 | 0.017654 | 18.6581 | 61.9808 | 0.01336 | 0.03521 | 0.9244 | 2.6355 | 0.01235 | 51.8823 | 172.3492 | 0.2849 | 2.061 | 0.7897 | 7.2341 | 0.22499 | 2.4329 | 8.0821 | 0.04008 | 0.08047 | 1.03 | 2.0077 | 0.041282 | 17.2941 | 57.4497 | 0.006921 | 0.0001976 | 3.309 | 0.028551 | 0.022902 | 100.1513 | 332.6954
115 | 2.265 | 10.21 | 0.6363 | 4.5077 | 1.4412 | 0.30603 | 1.0166 | 0.03636 | 0.07201 | 0.5319 | 1.9805 | 0.01934 | 19.0635 | 63.3274 | 0.01283 | 0.06271 | 0.9423 | 4.8878 | 0.01209 | 54.0255 | 179.4688 | 0.2579 | 1.922 | 0.8008 | 7.4525 | 0.20653 | 2.6877 | 8.9282 | 0.03878 | 0.07394 | 1.033 | 1.9067 | 0.04006 | 17.8738 | 59.3756 | 0.006864 | -0.0003908 | 0.6469 | -0.056935 | 0.0044403 | 100.983 | 335.4582
120 | 2.246 | 8.95 | 0.5873 | 3.9849 | 1.3191 | 0.30861 | 1.0252 | 0.03566 | 0.07109 | 0.6073 | 1.9936 | 0.021656 | 19.4377 | 64.5705 | 0.01235 | 0.03047 | 0.9566 | 2.4672 | 0.011814 | 56.1253 | 186.4441 | 0.2336 | 1.797 | 0.8116 | 7.6926 | 0.18959 | 2.9672 | 9.857 | 0.03758 | 0.06808 | 1.031 | 1.8116 | 0.038745 | 18.4446 | 61.2716 | 0.006726 | -0.0008308 | 1.006 | -0.12352 | 0.0067664 | 103.0549 | 342.3409
125 | 2.219 | 7.923 | 0.5386 | 3.5705 | 1.1952 | 0.31237 | 1.0377 | 0.03502 | 0.07113 | 0.6974 | 2.0311 | 0.024423 | 19.7929 | 65.7506 | 0.01192 | 0.02863 | 0.9684 | 2.4018 | 0.011543 | 58.1499 | 193.1699 | 0.213 | 1.677 | 0.8217 | 7.8732 | 0.17502 | 3.2542 | 10.8103 | 0.03652 | 0.06304 | 1.031 | 1.7262 | 0.037652 | 18.9799 | 63.05 | 0.006584 | -0.001214 | 1.192 | -0.18439 | 0.0078481 | 105.2775 | 349.7243
130 | 2.17 | 7.094 | 0.4909 | 3.2691 | 1.0653 | 0.31942 | 1.0611 | 0.03445 | 0.0716 | 0.7969 | 2.0784 | 0.027453 | 20.1204 | 66.8385 | 0.01155 | 0.02702 | 0.9802 | 2.3394 | 0.011321 | 60.0127 | 199.358 | 0.1969 | 1.557 | 0.8309 | 7.9076 | 0.1636 | 3.5203 | 11.6942 | 0.03561 | 0.05874 | 1.037 | 1.6495 | 0.036928 | 19.465 | 64.6612 | 0.006472 | -0.001539 | 1.285 | -0.23779 | 0.0083165 | 107.0994 | 355.7764
135 | 2.102 | 6.45 | 0.4469 | 3.0685 | 0.93938 | 0.32976 | 1.0954 | 0.03394 | 0.07263 | 0.9099 | 2.14 | 0.030882 | 20.4227 | 67.8428 | 0.01122 | 0.02561 | 0.9901 | 2.2825 | 0.011109 | 61.7778 | 205.2215 | 0.1838 | 1.44 | 0.8391 | 7.8346 | 0.15423 | 3.7712 | 12.5277 | 0.03481 | 0.05519 | 1.049 | 1.5855 | 0.036516 | 19.9123 | 66.1472 | 0.006306 | -0.001731 | 1.465 | -0.2745 | 0.0092383 | 109.9187 | 365.1419
140 | 2.009 | 5.916 | 0.4018 | 2.9447 | 0.80722 | 0.34502 | 1.1461 | 0.03345 | 0.07476 | 1.047 | 2.235 | 0.035022 | 20.7219 | 68.8366 | 0.01088 | 0.02436 | 0.9964 | 2.239 | 0.010841 | 63.7084 | 211.6347 | 0.1724 | 1.328 | 0.8458 | 7.703 | 0.14582 | 4.0206 | 13.3561 | 0.03407 | 0.05145 | 1.057 | 1.5101 | 0.036012 | 20.3448 | 67.5839 | 0.006191 | -0.001849 | 1.53 | -0.29866 | 0.0094722 | 111.9605 | 371.9246
145 | 1.895 | 5.498 | 0.358 | 2.9013 | 0.67841 | 0.36578 | 1.2151 | 0.03296 | 0.07875 | 1.224 | 2.3893 | 0.040343 | 21.03 | 69.86 | 0.01056 | 0.02313 | 0.9987 | 2.1903 | 0.010546 | 65.6389 | 218.0478 | 0.1616 | 1.225 | 0.8519 | 7.5804 | 0.13767 | 4.2893 | 14.2487 | 0.03336 | 0.04795 | 1.063 | 1.4374 | 0.035462 | 20.7778 | 69.0223 | 0.006115 | -0.001869 | 1.498 | -0.30564 | 0.0091603 | 113.352 | 376.547
150 | 1.757 | 5.177 | 0.3156 | 2.9465 | 0.55451 | 0.39451 | 1.3105 | 0.03243 | 0.08599 | 1.467 | 2.6516 | 0.047575 | 21.3736 | 71.0017 | 0.0103 | 0.02198 | 1.013 | 2.134 | 0.010434 | 67.2958 | 223.552 | 0.1501 | 1.132 | 0.8566 | 7.5416 | 0.12858 | 4.6179 | 15.3403 | 0.03266 | 0.04491 | 1.073 | 1.3751 | 0.035044 | 21.2231 | 70.5017 | 0.00602 | -0.001752 | 1.483 | -0.29103 | 0.0089277 | 115.1407 | 382.4892
Air kerma at 1m per unit workload (mGy/mA-min)
kVp | Ko**
25 | 1.4835
30 | 2.466
35 | 3.4485
40 | 0.7189
45 | 0.87211
50 | 1.0655
55 | 1.2967
60 | 1.5634
65 | 1.8633
70 | 2.1939
75 | 2.5529
80 | 2.938
85 | 3.3468
90 | 3.777
95 | 4.2261
100 | 4.692
105 | 5.1722
110 | 5.6643
115 | 6.166
120 | 6.675
125 | 7.1889
130 | 7.7053
135 | 8.2219
140 | 8.7364
145 | 9.2464
150 | 9.7495

## NCRP147_B1
Fitting parameters for transmission of broad primary x-ray beams to Equation A.2 (thickness is input in mm)
Workload Distribution | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
| a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag
Rad Room (all) | 2.346 | 15.9 | 0.4982 | 6.7775 | 1.1688 | 0.03626 | 0.1429 | 0.4932 | 3.941 | 0.017883 | 0.0142 | 0.05781 | 0.7445 | 4.0711 | 0.010572 | 0.2163 | 3.101 | 0.5745 | 14.3366 | 0.12426 | 0.03907 | 0.1069 | 0.594 | 2.7361 | 0.023208 | 0.007616 | 0.000767 | 1.027 | 0.10071 | 0.0078216
Rad Room (chest bucky) | 2.264 | 13.08 | 0.56 | 5.7774 | 1.2678 | 0.03552 | 0.1177 | 0.6007 | 3.3136 | 0.021337 | 0.01278 | 0.04848 | 0.8609 | 3.7934 | 0.011002 | 0.2179 | 2.677 | 0.7209 | 12.2855 | 0.15708 | 0.03762 | 0.09751 | 0.7867 | 2.592 | 0.029596 | 0.007142 | 0.000308 | 1.617 | 0.043125 | 0.011549
Rad Room (floor or other) | 2.651 | 16.56 | 0.4585 | 6.2467 | 1.2155 | 0.03994 | 0.1448 | 0.4231 | 3.6254 | 0.016899 | 0.01679 | 0.06124 | 0.7356 | 3.6474 | 0.012351 | 2535 | 2.74 | 0.4297 | 0.0010809 | 1089.2895 | 0.04361 | 0.1082 | 0.5463 | 2.4811 | 0.023824 | 0.007915 | 0.00088 | 0.979 | 0.11118 | 0.0077488
Fluoroscopy tube (RF) | 2.347 | 12.67 | 0.6149 | 5.3984 | 1.4432 | 0.03616 | 0.09721 | 0.5186 | 2.6883 | 0.018753 | 0.0134 | 0.04283 | 0.8796 | 3.1963 | 0.011787 | 0.2323 | 2.19 | 0.6509 | 9.4275 | 0.1512 | 0.03901 | 0.08588 | 0.8081 | 2.2015 | 0.031524 | 0.007089 | 0.000474 | 1.58 | 0.066864 | 0.011201
Rad tube (RF) | 2.295 | 13 | 0.5573 | 5.6645 | 1.279 | 0.03549 | 0.1164 | 0.5774 | 3.2798 | 0.020492 | 0.013 | 0.04778 | 0.8485 | 3.6754 | 0.011031 | 0.2126 | 2.568 | 0.6788 | 12.079 | 0.14431 | 0.03778 | 0.09365 | 0.7483 | 2.4788 | 0.028271 | 0.007162 | 0.000411 | 1.541 | 0.057386 | 0.011037
Chest Room | 2.283 | 10.74 | 0.637 | 4.7043 | 1.4543 | 0.03622 | 0.07766 | 0.5404 | 2.1441 | 0.019573 | 0.01286 | 0.03505 | 0.9356 | 2.7255 | 0.012032 | 0.25 | 1.989 | 0.7721 | 7.956 | 0.19303 | 0.03866 | 0.07721 | 0.9843 | 1.9972 | 0.038053 | 0.00765 | -0.00098 | 0.08083 | -0.1281 | 0.00061835
Mammo Room | 30.6 | 177.6 | 0.3308 | 5.8039 | 10.1225 | 0.2577 | 1.765 | 0.3644 | 6.849 | 0.093906 | 0.09148 | 0.709 | 0.3459 | 7.7503 | 0.031643 | 5.998 | 42.91 | 0.3927 | 7.1541 | 2.3554 | 0.2467 | 1.654 | 0.3694 | 6.7045 | 0.091131 | 0.01914 | 0.04166 | 0.2858 | 2.1766 | 0.0054702
Cardiac Angiography | 2.389 | 14.26 | 0.5948 | 5.969 | 1.421 | 0.03717 | 0.1087 | 0.4879 | 2.9244 | 0.018135 | 0.01409 | 0.04814 | 0.8419 | 3.4166 | 0.011862 | 0.2533 | 2.461 | 0.6243 | 9.7158 | 0.15814 | 0.04025 | 0.09482 | 0.7523 | 2.3558 | 0.03028 | 0.007303 | 0.000722 | 1.204 | 0.098863 | 0.0087928
Peripheral Angiography | 2.728 | 18.52 | 0.4614 | 6.7889 | 1.2587 | 0.04292 | 0.1538 | 0.4236 | 3.5834 | 0.018181 | 0.01774 | 0.06449 | 0.7158 | 3.6353 | 0.012698 | 0.367 | 3.26 | 0.5036 | 8.8828 | 0.18482 | 0.04642 | 0.1203 | 0.5763 | 2.5916 | 0.026752 | 0.008103 | 0.000844 | 0.9754 | 0.10416 | 0.0079037

## NCRP147_C1
Fits of transmission for broad-beam secondary transmission to Equation A.2 (thickness x is input in millimeters)
Workload Distribution | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
| a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag
30 | 38.79 | 180 | 0.356 | 4.6404 | 13.8092 | 0.3174 | 1.724 | 0.3705 | 5.4316 | 0.1176 | 0.1198 | 0.7137 | 0.3703 | 5.9574 | 0.044362 | 7.408 | 42.49 | 0.4061 | 5.7357 | 3.0084 | 0.306 | 1.62 | 0.3793 | 5.2941 | 0.11607 | 0.02158 | 0.03971 | 0.2852 | 1.8401 | 0.0061546
50 | 8.801 | 27.28 | 0.2957 | 3.0996 | 2.6025 | 0.0903 | 0.1712 | 0.2324 | 1.8959 | 0.020986 | 0.0388 | 0.0873 | 0.5105 | 2.25 | 0.019807 | 1.817 | 4.84 | 0.4021 | 2.6637 | 0.73062 | 0.09721 | 0.1799 | 0.4912 | 1.8506 | 0.04775 | 0.01076 | 0.001862 | 1.17 | 0.17305 | 0.012589
70 | 5.369 | 23.49 | 0.5883 | 4.3751 | 3.1586 | 0.0509 | 0.1697 | 0.3849 | 3.334 | 0.019591 | 0.023 | 0.0716 | 0.73 | 3.113 | 0.01679 | 0.7149 | 3.789 | 0.5381 | 5.3 | 0.38469 | 0.05791 | 0.1357 | 0.5968 | 2.3433 | 0.034561 | 0.00855 | 0.000539 | 1.194 | 0.063041 | 0.010209
100 | 2.507 | 15.33 | 0.9124 | 6.1149 | 2.2874 | 0.0395 | 0.0844 | 0.5191 | 2.1367 | 0.020504 | 0.0147 | 0.04 | 0.9752 | 2.7211 | 0.014335 | 0.3424 | 2.456 | 0.9388 | 7.1729 | 0.32145 | 0.04279 | 0.08948 | 1.029 | 2.0911 | 0.044031 | 0.00723 | 0.000894 | 1.316 | 0.12365 | 0.0095147
125 | 2.233 | 7.888 | 0.7295 | 3.5325 | 1.629 | 0.0351 | 0.066 | 0.7832 | 1.8803 | 0.02749 | 0.012 | 0.0267 | 10.79 | 2.225 | 0.12948 | 0.2138 | 1.69 | 1.086 | 7.9046 | 0.23219 | 0.03654 | 0.0579 | 1.093 | 1.5846 | 0.039938 | 0.006587 | -0.00114 | 1.172 | -0.17307 | 0.00772
150 | 1.791 | 5.478 | 0.5678 | 3.0586 | 1.0169 | 0.0324 | 0.0775 | 1.566 | 2.392 | 0.050738 | 0.0104 | 0.0202 | 1.135 | 1.9423 | 0.011804 | 0.1511 | 1.124 | 1.151 | 7.4388 | 0.17392 | 0.03267 | 0.04074 | 1.134 | 1.247 | 0.037048 | 0.006027 | -0.00163 | 1.44 | -0.27045 | 0.0086789
Rad Room (all) | 2.298 | 17.38 | 0.6193 | 7.5631 | 1.4232 | 0.0361 | 0.1433 | 0.56 | 3.9695 | 0.020216 | 0.0138 | 0.057 | 0.7937 | 4.1304 | 0.010953 | 0.2191 | 3.49 | 0.7358 | 15.9288 | 0.16121 | 0.03873 | 0.1054 | 0.6397 | 2.7214 | 0.024776 | 0.007552 | 0.000737 | 1.044 | 0.09759 | 0.0078843
Rad Room (chest bucky) | 2.256 | 13.8 | 0.8837 | 6.117 | 1.9936 | 0.0356 | 0.1079 | 0.7705 | 3.0309 | 0.02743 | 0.0127 | 0.0445 | 1.049 | 3.5039 | 0.013322 | 0.2211 | 2.836 | 1.123 | 12.8268 | 0.2483 | 0.03749 | 0.0871 | 0.9086 | 2.3233 | 0.034063 | 0.007058 | 0.000229 | 1.875 | 0.032445 | 0.013234
Rad Room (floor or other) | 2.513 | 17.34 | 0.4994 | 6.9001 | 1.255 | 0.0392 | 0.1464 | 0.4486 | 3.7347 | 0.017585 | 0.0164 | 0.0608 | 0.7472 | 3.7073 | 0.012254 | 0.244 | 3.012 | 0.5019 | 12.3443 | 0.12246 | 0.04299 | 0.107 | 0.5538 | 2.489 | 0.023808 | 0.007887 | 0.000877 | 0.98 | 0.1112 | 0.0077293
Fluoroscopy tube (RF) | 2.322 | 12.91 | 0.7575 | 5.5599 | 1.7589 | 0.0363 | 0.0936 | 0.5955 | 2.5785 | 0.021617 | 0.0133 | 0.041 | 0.9566 | 3.0827 | 0.012723 | 0.2331 | 2.213 | 0.8051 | 9.4938 | 0.18767 | 0.03886 | 0.08091 | 0.852 | 2.0821 | 0.033109 | 0.007057 | 0.000422 | 1.664 | 0.059799 | 0.011743
Rad tube (RF) | 2.272 | 13.6 | 0.7184 | 5.9859 | 1.6322 | 0.0356 | 0.1114 | 0.662 | 3.1292 | 0.023567 | 0.0129 | 0.0457 | 0.9355 | 3.5426 | 0.012068 | 0.2149 | 2.695 | 0.8768 | 12.5407 | 0.18842 | 0.03762 | 0.08857 | 0.8087 | 2.3543 | 0.030423 | 0.007102 | 0.000345 | 1.698 | 0.048578 | 0.012059
Chest Room | 2.288 | 9.848 | 1.054 | 4.3042 | 2.4116 | 0.0364 | 0.0659 | 0.7543 | 1.8104 | 0.027457 | 0.013 | 0.0297 | 1.195 | 2.2846 | 0.015535 | 0.2518 | 1.829 | 1.273 | 7.2637 | 0.32054 | 0.03866 | 0.0627 | 1.128 | 1.6218 | 0.043608 | 0.007485 | -0.00081 | 0.09459 | -0.10822 | 0.00070801
Mammo Room | 29.91 | 184.4 | 0.355 | 6.1652 | 10.6181 | 0.2539 | 1.8411 | 0.3924 | 7.2513 | 0.09963 | 0.0883 | 0.7526 | 0.3786 | 8.5232 | 0.03343 | 5.789 | 44.12 | 0.4124 | 7.6214 | 2.3874 | 0.2404 | 1.709 | 0.3918 | 7.109 | 0.094189 | 0.01888 | 0.04172 | 0.2903 | 2.2097 | 0.0054809
Cardiac Angiography | 2.354 | 14.94 | 0.7481 | 6.3466 | 1.761 | 0.0371 | 0.1067 | 0.5733 | 2.876 | 0.021269 | 0.0139 | 0.0464 | 0.9185 | 3.3381 | 0.012767 | 0.253 | 2.592 | 0.7999 | 10.2451 | 0.20237 | 0.04001 | 0.0903 | 0.8019 | 2.2569 | 0.032084 | 0.007266 | 0.000674 | 1.2335 | 0.092761 | 0.0089626
Peripheral Angiography | 2.661 | 19.54 | 0.5094 | 7.3431 | 1.3555 | 0.04219 | 0.1559 | 0.4472 | 3.6952 | 0.018867 | 0.01747 | 0.06422 | 0.7299 | 3.676 | 0.012751 | 0.3579 | 3.466 | 0.56 | 9.6843 | 0.20042 | 0.04612 | 0.1198 | 0.5907 | 2.5976 | 0.027243 | 0.008079 | 0.000847 | 0.9742 | 0.10484 | 0.0078706
120 | 2.2486 | 8.61 | 0.77103

## W_Rh_Al_Ag
Table 1: Fitted results of a, b, and g for broad tungsten-anode, silver-filter x-ray beams. Units of a and b in mm^-1
W/Ag | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
kVp | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag
25 | 67 | 21.26 | 0.2288 | 0.31731 | 15.3296 | 0.4288 | 0.5526 | 0.4359 | 1.2887 | 0.18691 | 0.2174 | 0.2774 | 0.3707 | 1.276 | 0.08059 | 11.15 | 7.133 | 0.1884 | 0.63973 | 2.1007 | 0.4812 | 0.601 | 0.472 | 1.249 | 0.22713 | 0.02201 | 0.01011 | 0.3083 | 0.45934 | 0.0067857
30 | 46.95 | 33.73 | 0.1923 | 0.71842 | 9.0285 | 0.2891 | 0.4375 | 0.2742 | 1.5133 | 0.079271 | 0.1384 | 0.2394 | 0.2387 | 1.7298 | 0.033036 | 8.295 | 8.655 | 0.2527 | 1.0434 | 2.0961 | 0.3197 | 0.5001 | 0.2888 | 1.5643 | 0.092329 | 0.01756 | 0.008786 | 0.1797 | 0.50034 | 0.0031555
35 | 28.41 | 62.94 | 0.3201 | 2.2154 | 9.094 | 0.1541 | 0.5974 | 0.309 | 3.8767 | 0.047617 | 0.06557 | 0.3031 | 0.2593 | 4.6225 | 0.017002 | 5.402 | 16.81 | 0.5086 | 3.1118 | 2.7475 | 0.1514 | 0.6343 | 0.2725 | 4.1896 | 0.041257 | 0.01186 | 0.01312 | 0.1603 | 1.1062 | 0.0019012
39 | 15.44 | 67.38 | 0.2861 | 4.364 | 4.4174 | 0.1292 | 0.6718 | 0.4475 | 5.1997 | 0.057817 | 0.06427 | 0.3709 | 0.4628 | 5.771 | 0.029744 | 2.956 | 14.31 | 0.3706 | 4.841 | 1.0955 | 0.1443 | 0.7505 | 0.4603 | 5.201 | 0.066421 | 0.01071 | 0.01373 | 0.2387 | 1.282 | 0.0025565
Source: | Li X, Zhang D, Liu B, “Workload and transmission data for the installation of a digital breast tomosynthesis system”, MedPhys 40(6) 063901-1
Table 1: Fitted results of a, b, and g for broad tungsten-anode, aluminum-filter x-ray beams. Units of a and b in mm^-1
W/Al | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
kVp | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag
25 | 63.06 | 27.09 | 0.149 | 0.42959 | 9.3959 | 0.4283 | 0.645 | 0.434 | 1.506 | 0.18588 | 0.218 | 0.3308 | 0.3798 | 1.5174 | 0.082796 | 11.44 | 8.967 | 0.2344 | 0.78383 | 2.6815 | 0.4733 | 0.7575 | 0.4454 | 1.6005 | 0.21081 | 0.0219 | 0.01372 | 0.3181 | 0.62648 | 0.0069664
30 | 25.72 | 49.12 | 0.1217 | 1.9098 | 3.1301 | 0.2545 | 0.4988 | 0.3782 | 1.9599 | 0.096252 | 0.1391 | 0.2945 | 0.4519 | 2.1172 | 0.062859 | 5.831 | 9.782 | 0.2166 | 1.6776 | 1.263 | 0.2919 | 0.6105 | 0.4443 | 2.0915 | 0.12969 | 0.01723 | 0.01196 | 0.4733 | 0.69414 | 0.008155
35 | 22.2 | 42.34 | 0.2152 | 1.9072 | 4.7774 | 0.1826 | 0.4967 | 0.4828 | 2.7202 | 0.088159 | 0.092 | 0.2694 | 0.4893 | 2.9283 | 0.045016 | 4.344 | 9.436 | 0.331 | 2.1722 | 1.4379 | 0.1975 | 0.5284 | 0.4656 | 2.6754 | 0.091956 | 0.01417 | 0.009502 | 0.463 | 0.67057 | 0.0065607
40 | 17.92 | 41.58 | 0.3201 | 2.3203 | 5.7362 | 0.1379 | 0.462 | 0.5482 | 3.3503 | 0.075597 | 0.06727 | 0.251 | 0.5468 | 3.7312 | 0.036783 | 2.957 | 8.541 | 0.3569 | 2.8884 | 1.0554 | 0.1492 | 0.4492 | 0.5028 | 3.0107 | 0.075018 | 0.01244 | 0.00844 | 0.5356 | 0.67846 | 0.0066629
45 | 12.97 | 39.85 | 0.3514 | 3.0725 | 4.5577 | 0.1093 | 0.4065 | 0.5917 | 3.7191 | 0.064673 | 0.05115 | 0.221 | 0.583 | 4.3206 | 0.02982 | 2.297 | 8.032 | 0.4421 | 3.4967 | 1.0155 | 0.1173 | 0.4391 | 0.5791 | 3.7434 | 0.067928 | 0.01142 | 0.008285 | 0.7651 | 0.72548 | 0.0087374
49 | 11.5 | 37.07 | 0.4101 | 3.2235 | 4.7161 | 0.09267 | 0.4031 | 0.6295 | 4.3498 | 0.058336 | 0.04267 | 0.2252 | 0.6361 | 5.2777 | 0.027142 | 1.752 | 8.332 | 0.4704 | 4.7557 | 0.82414 | 0.1007 | 0.4478 | 0.643 | 4.4469 | 0.06475 | 0.01073 | 0.008158 | 0.8549 | 0.7603 | 0.0091731
Source: | Li X, Zhang D, Liu B, “Transmission of broad W/Rh and W/Al (target/filter) x-ray beams operated at 25-29 kVp through common shielding materials”, MedPhys 39(7) 4132
Table 2: Fitted results of a, b, and g for broad tungste-anode, rhodium-filter x-ray beams. Units of a and b in mm^-1
W/Rh | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
kVp | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag
25 | 67.94 | 20.96 | 0.08503 | 0.30851 | 5.7769 | 0.4906 | 0.4668 | 0.3878 | 0.95149 | 0.19025 | 0.2432 | 0.2539 | 0.3153 | 1.044 | 0.076681 | 13.71 | 4.772 | 0.1771 | 0.34807 | 2.428 | 0.5308 | 0.5201 | 0.3626 | 0.97984 | 0.19247 | 0.02429 | 0.009839 | 0.3314 | 0.40506 | 0.0080497
30 | 49.72 | 45.85 | 0.2231 | 0.92216 | 11.0925 | 0.2721 | 0.6599 | 0.2827 | 2.4252 | 0.076923 | 0.1538 | 0.3565 | 0.3216 | 2.3179 | 0.049462 | 7.87 | 12.8 | 0.2533 | 1.6264 | 1.9935 | 0.3101 | 0.7473 | 0.3029 | 2.4099 | 0.093929 | 0.01855 | 0.01307 | 0.2267 | 0.70458 | 0.0042053
35 | 25.61 | 82.42 | 0.3307 | 3.2183 | 8.4692 | 0.158 | 0.8254 | 0.3908 | 5.2241 | 0.061746 | 0.07447 | 0.4409 | 0.3739 | 5.9205 | 0.027844 | 3.199 | 17.11 | 0.2814 | 5.3485 | 0.9002 | 0.1702 | 0.9212 | 0.3908 | 5.4125 | 0.066514 | 0.01378 | 0.02024 | 0.3624 | 1.4688 | 0.0049939
39 | 17.96 | 81.68 | 0.3821 | 4.5479 | 6.8625 | 0.1376 | 0.9611 | 0.5957 | 6.9847 | 0.081968 | 0.06656 | 0.508 | 0.5791 | 7.6322 | 0.038545 | 3.134 | 18.14 | 0.4678 | 5.7881 | 1.4661 | 0.1485 | 1.079 | 0.5887 | 7.266 | 0.087422 | 0.01154 | 0.01996 | 0.4113 | 1.7296 | 0.0047464
Source: | Li X, Zhang D, Liu B, “Transmission of broad W/Rh and W/Al (target/filter) x-ray beams operated at 25-29 kVp through common shielding materials”, MedPhys 39(7) 4132
Table V: Fitting parameters a, b, and g for the transmission of the primary radiation generated by the combined workload distribution [Table IV(c).
|  | Pb (mm) |  |  |  |  | Concrete (mm) |  |  |  |  | Gypsum |  |  |  |  | Steel |  |  |  |  | Plate Glass |  |  |  |  | Wood
|  | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag | a | b | g | b/a | ag
W/Rh, W/Ag, W/Al |  | 13.33 | 65.26 | 0.2119 | 4.8957 | 2.8246 | 0.1338 | 0.6346 | 0.3777 | 4.7429 | 0.050536 | 0.06343 | 0.3415 | 0.3708 | 5.3839 | 0.02352 | 2.56 | 13.66 | 0.2724 | 5.3359 | 0.69734 | 0.1444 | 0.6858 | 0.3714 | 4.7493 | 0.05363 | 0.01285 | 0.0141 | 0.3422 | 1.0973 | 0.0043973
| W/Al | 12.22 | 53.6 | 0.1948 | 4.3863 | 2.3805 | 0.134 | 0.5085 | 0.3928 | 3.7948 | 0.052635 | 0.06354 | 0.2744 | 0.3843 | 4.3185 | 0.024418 | 2.422 | 10.91 | 0.2576 | 4.5045 | 0.62391 | 0.144 | 0.5412 | 0.3799 | 3.7583 | 0.054706 | 0.01297 | 0.01057 | 0.401 | 0.81496 | 0.005201
Source: | Li X, Zhang D, Liu B, “Workload and transmission data for the installation of a digital breast tomosynthesis system”, MedPhys 40(6) 063901-1