## Barco
1 | [11 empty] | Revision 1.4-20200130 | [10 empty] | Print Area |  |  |  |  | Test frequency
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [10 empty] | Standard monitor |  | Tomosynthesis monitor
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina | [10 empty] | No QAWS | QAWS | No QAWS | QAWS
4 | [22 empty] | All: |  |  |  | Autocalibration | Semiannual | Semiannual | Semiannual | Semiannual
5 | [5 empty] | Mammography Monitor Compliance Inspection | [10 empty] | Mammography Monitor Compliance Inspection | [9 empty] | Compliance Test | Weekly | Weekly | Weekly | Monthly
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year | Display Test |  |  | NA | Weekly
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  |  |  | Rapid Frame | NA | NA | Quarterly | Quarterly
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [5 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah
10 |  | Facility: |  |  |  |  | Site Number: | [6 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility:
11 |  | Department: |  |  |  |  | Date of Installation: | [6 empty] | Department: |  |  |  | Date of Installation: |  |  |  |  | Department:
12 |  | Area/Division: |  |  |  |  | Workstation ID: | [6 empty] | Area/Division: |  |  |  | Workstation ID: |  |  |  |  | Area/Division:
13 |  | Survey ID: |  |  |  |  | Left monitor SN: | [6 empty] | Survey ID: |  |  |  | Left monitor SN: |  |  |  |  | Survey ID:
14 |  | Room Number: |  |  |  |  | Right monitor SN: | [6 empty] | Room Number: |  |  |  | Right monitor SN: |  |  |  |  | Room Number:
15 |  | QC Manual version: | IMPAX 4.2 2015 | [10 empty] | QC Manual version: | IMPAX 4.2 2015 | IMPAX 4.2 2015 | [6 empty] | QC Manual version: |  | Change | IMPAX 4.2 2015
16 | [22 empty] | Site Number:
17 | [10 empty] | Medical University of South Carolina | [5 empty] | Inspection Results | [5 empty] | Date of Installation:
18 | [10 empty] | Mammography Monitor Compliance Inspection | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Workstation ID:
19 |  |  |  |  | Comments |  |  |  | Test Equipment | [5 empty] | Comments |  |  |  | Test Equipment |  |  |  | Left monitor SN:
20 |  | Monitor surface is clean: | [6 empty] | Meter: | Radcal 9010 |  |  | Monitor surface is clean | [5 empty] | Meter: | Radcal 9010 | Radcal 9010 |  | Right monitor SN:
21 |  | Viewing conditions: | [6 empty] | Calibration Date: | 2019-07-24 |  |  | Viewing conditions | [5 empty] | Calibration Date: | 2019-07-24 | 2019-07-24
22 |  | TG18-QC test pattern: | [6 empty] | Calibration Due: | 2020-07-24 |  |  | TG18-QC test pattern | [5 empty] | Calibration Due: | 2020-07-24 | 2020-07-24 |  | Test Equipment
23 |  | Rapid Frame acceptable: | [10 empty] | Rapid Frame acceptable | [9 empty] | Meter: |  | Change | Radcal 9010
24 |  | Display Test: | [20 empty] | Calibration Date: |  | Change | 2019-07-24
25 | [11 empty] | Room Conditions | [5 empty] | Frequency: | Monthly |  |  |  | Calibration Due: |  | Change | 2020-07-24
26 | Room Conditions | [10 empty] | Ambient Light: |  | lux
27 | Ambient light: |  | lux |  | Criteria: | Ambient light measured at the surface of the monitor(s) must be less than 20 lux | [5 empty] | Previous: |  | lux | Criteria: | Ambient light measured at the surface of the monitor(s) must be less than 20 lux | [6 empty] | Ambient Light:
28 | Comments: | [11 empty] | Comments: | [9 empty] | Auto Cal Lum Ratio L:
29 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | Auto Cal Lum Ratio R:
30 | [12 empty] | New: | [9 empty] | Uniformity Dev UNL10 L:
31 | [22 empty] | Uniformity Dev UNL10 R:
32 | Auto Calibration | [10 empty] | Auto Calibration | [5 empty] | Frequency: | Semi-annually |  |  |  | Uniformity Dev UNL80 L:
33 | Date: |  |  | Lmin | Lmax | Luminance Ratio | [5 empty] | Date: |  |  | Left () | Right () | [6 empty] | Uniformity Dev UNL80 R:
34 |  |  | Left () | [10 empty] | Lmin (cd/m^2) | [8 empty] | QUALITY CONTROL visibility:
35 |  |  | Right () | [10 empty] | Lmax (cd/m^2) | [8 empty] | Black L:
36 | Comments: | [12 empty] | Luminance Ratio | [8 empty] | Gray L:
37 | [22 empty] | White L:
38 | [12 empty] | Comments: | [9 empty] | Black R:
39 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | Gray R:
40 | Compliance Test | [11 empty] | New: | [9 empty] | White R:
41 | Date: |  |  | Left () | Right ()
42 |  |  | DICOM GSDF curve acceptable: | [8 empty] | Compliance Test | [5 empty] | Frequency: | Monthly (if using QA Web Server) |  |  |  | Comments
43 | Comments: | [10 empty] | Date: | [6 empty] | Weekly (if not using QA Web Server) |  |  |  | Room Conditions:
44 | [11 empty] | Enter 1 for YES, 2 for NO, 3 for NA |  |  | Left () | Right () | [6 empty] | Auto Calibration:
45 | [13 empty] | DICOM GSDF curve acceptable: | [8 empty] | Compliance Test:
46 | [12 empty] | Comments: | [9 empty] | Display Test:
47 | TG-18 Test Patterns | [11 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | TG-18 Black/White:
48 | TG-18 LN01 | [5 empty] | TG-18 LN18 | [5 empty] | New: | [9 empty] | TG-18 QC:
49 | Left () |  | Right () |  |  |  | Left () |  | Right () | [13 empty] | TG-18 Uniformity:
50 | Comments: | [10 empty] | Display Test | [5 empty] | Frequency: | Weekly (if using QA Web Server)
51 | [11 empty] | Date: | [6 empty] | Not available if not using QA Web Server |  |  |  | Comments
52 | [12 empty] | Comments:
53 | [12 empty] | Do not exceed 390 chars |  | 0
54 | TG-18 QC | [5 empty] | QUALITY CONTROL visibility | [5 empty] | New:
55 |  |  |  | Left () | Right () |  | Left () |  | Right ()
56 |  |  | 5%/95% patches visible: |  |  | Black | [5 empty] | TG-18 Test Patterns
57 |  |  | Resolution acceptable: |  |  | Gray | [5 empty] | TG-18 Black/White |  | Left () | Right ()
58 |  |  | Display is free of artifacts/noise: |  |  | White | [6 empty] | LN-01 (cd/m^2)
59 | Comments: | [11 empty] | LN-18 (cd/m^2)
60 | [12 empty] | Comments:
61 | [12 empty] | Do not exceed 390 chars |  | 0
62 | [12 empty] | New:
63
64 | [11 empty] | TG-18 QC | [5 empty] | Frequency: | Annually
65 | Date: | [8 empty] | Inspector: | Eugene Mah | Enter 1 for YES, 2 for NO, 3 for NA |  |  | Left () | Right ()
66 | Room Number | [8 empty] | Survey ID: |  |  |  | 5%/95% patches visible:
1 | [10 empty] | Medical University of South Carolina |  |  | Resolution acceptable:
2 | [10 empty] | Mammography Monitor Compliance Inspection |  |  | Display is free of artifacts/noise:
3 | [13 empty] | QUALITY CONTROL visibility: | Enter as much of “QUALITY CONTROL” that is visible |  |  |  | Previous Year
4 | TG-18 Uniformity | [12 empty] | Black
5 |  | UNL-10 (cd/m^2) |  | UNL-80 (cd/m^2) |  |  | Visual Assessment
6 |  | Left () | Right () | Left () | Right () |  | TBD | Free from gross non-uniformities | [5 empty] | Gray
7 | Average: | [5 empty] | TBD | Free from non-symmetric non-uniformities
8 | SD: | [5 empty] | TBD | Free from significant luminance variation | [5 empty] | White
9 | Deviation (%)
10 | [6 empty] | Criteria: | Luminance deviation is < 30% | [5 empty] | Criteria: | There should be no significant changes to how much of QUALITY CONTROL is visible
11 | Comments:
12 | [12 empty] | Comments:
13 | [12 empty] | Do not exceed 390 chars |  | 0
14 | [12 empty] | New:
15 | Comments:
16 | [11 empty] | TG-18 Uniformity | [6 empty] | Visual assessment
17 | [13 empty] | UNL-10 (cd/m^2) |  | UNL-80 (cd/m^2) |  |  | Enter 1 for YES, 2 for NO, 3 for NA
18 | [13 empty] | Left () | Right () | Left () | Right () |  |  | Free from gross non-uniformities
19 | [12 empty] | Top left | [6 empty] | Free from non-symmetric non-uniformities
20 | [12 empty] | Top right | [6 empty] | Free from significant luminance variation
21 | [12 empty] | Center
22 | [12 empty] | Bottom left
23 | [12 empty] | Bottom right
24 | [12 empty] | Average (cd/m^2):
25 | [12 empty] | SD (cd/m^2):
26 | [12 empty] | Luminance Deviation (%): |  |  |  |  | Criteria: | Luminance deviation is < 30%
27 | [12 empty] | Comments:
28 | [12 empty] | Do not exceed 390 chars |  | 0
29 | [12 empty] | New:
30
31
32 | [16 empty] | Comments Page 1
33 | [15 empty] | Previous Year Comments
34 | [12 empty] | Comments:
35 | [12 empty] | Do not exceed 390 chars |  | 0
36 | [12 empty] | Additional Comments:
37 | [12 empty] | Do not exceed 390 chars |  | 0
38 | [12 empty] | Additional Comments:
39 | [12 empty] | Do not exceed 390 chars |  | 0
40 | [12 empty] | Additional Comments:
41 | [12 empty] | Do not exceed 390 chars |  | 0
42 | [12 empty] | Additional Comments:
43 | [12 empty] | Do not exceed 390 chars |  | 0
44 | [12 empty] | Additional Comments:
45 | [12 empty] | Do not exceed 390 chars |  | 0
46 | [12 empty] | Additional Comments:
47 | [12 empty] | Do not exceed 390 chars |  | 0
48 | [12 empty] | Additional Comments:
49 | [12 empty] | Do not exceed 390 chars |  | 0
50 | [12 empty] | Additional Comments:
51 | [12 empty] | Do not exceed 390 chars |  | 0
52 | [12 empty] | Additional Comments:
53 | [12 empty] | Do not exceed 390 chars |  | 0
54 | [12 empty] | Additional Comments:
55 | [12 empty] | Do not exceed 390 chars |  | 0
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID: