## QC Test Summary-Lorad
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Lorad

Site Name					Report Date
Address					Survey Date	0
Medical Physicist's Name			Eugene Mah					Signature
X-Ray Unit Manufacturer			Lorad/Hologic					Model
Date of Installation					Room ID

QC Manual Version #					(use any version applicable to model; contact mfr if questions)

Accessory Equipment			Manufacturer			Model		Location		QC Manual Version #
		Review Workstation*
		Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type			Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist)					Annual Survey

Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
					PASS/FAIL
1. Mammographic Unit Assembly Evaluation
2. Collimation Assessment
3. Artifact Evaluation
4. kVp Accuracy and Reproducibility
5. Beam Quality Assessment - HVL Measurement
6. Evaluation of System Resolution
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
		Average glandular dose for average breast is ≤3 mGy (300 mrad)					mGy
9. Radiation Output Rate
10. Phantom Image Quality Evaluation
		Phantom image scores:			Fibers	0.0	Specks	0.0	Masses	0.0
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
		SNR (value)
		CNR (value)		(Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
		CNR should not vary by more than ±15% (NA for Equipment Evaluation)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
13. DICOM Printer QC (Mammography Equipment Evaluations only)
14. Detector Flat Field Calibration (Mammography Equipment Evaluations only)
15. Compression Thickness Indicator (Mammography Equipment Evaluations only)
16. Compression (Mammography Equipment Evaluations only)

*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***

## QC Test Summary-Tomo
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Lorad

Site Name		0					Report Date	0
Address		0					Survey Date	0
Medical Physicist's Name			Eugene Mah					Signature
X-Ray Unit Manufacturer			Lorad/Hologic					Model
Date of Installation					Room ID

QC Manual Version #			0					(use any version applicable to model; contact mfr if questions)

Accessory Equipment			Manufacturer			Model		Location		QC Manual Version #
		Review Workstation*	0			0				0
		Film Printer*	0			0				0
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type			Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist)					Annual Survey

Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
					PASS/FAIL
1. Mammographic Unit Assembly Evaluation
	Autodecompression can be overridden to maintain compression (and status displayed)
	Manual emergency compression release can be activated in the event of a power failure
2. Collimation Assessment
	Deviation between x-ray field and light field is less than 2% SID
	X-ray field dose not extend beyond any side of the image receptor by more than 2% SID
3. Artifact Evaluation
	Artifacts were not apparent or not significant
4. kVp Accuracy and Reproducibility
	Measured average kVp within +/-5% of indicated kV
	kVp coefficient of variation <= 0.02
5. Beam Quality Assessment - HVL Measurement
	Half value layer is within acceptable lower limit at all kVp values tested
6. Evaluation of System Resolution
	Measured performance within acceptable limits
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
	Measured performance within acceptable limits
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
		Average glandular dose for average breast is ≤3 mGy (300 mrad)
		Average glandular dose to a 4.2 cm thick breast					Conventional		mGy
					Tomosynthesis		mGy
9. Radiation Output Rate
	Radiation output rate >= 2 mGy/s					mGy/s
10. Phantom Image Quality Evaluation					Fibers		Specks		Masses
		Phantom image scores (Conventional):					0.0		0.0		0.0
		Phantom image scores (Tomosynthesis):					0.0		0.0		0.0
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
		SNR (value)	0	SNR >= 40
		CNR (value)	0	(Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
		CNR should not vary by more than ±15% (NA for Equipment Evaluation)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
	White level performance
	Black level performance (CRT displays only)
	Quality level performance (GSDF compliance)
	Uniformity performance (CRT displays only)
13. DICOM Printer QC (Mammography Equipment Evaluations only)
14. Detector Flat Field Calibration (Mammography Equipment Evaluations only)
15. Compression Thickness Indicator (Mammography Equipment Evaluations only)
16. Compression (Mammography Equipment Evaluations only)

*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***

## Tech QC Eval-Tomo
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Lorad, continued)

Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.

					FREQUENCY		PASS/FAIL
1.	DICOM Printer Quality Control					Weekly
2.	Viewboxes and Viewing Conditions					Weekly
3.	Artifact Evaluation					Weekly
4.	Signal-To-Noise and Contrast-To-Noise Measurements					Weekly
5.	Phantom Image Quality Evaluation					Weekly
6.	Detector Flat-Field Calibration					Weekly
7.	Compression Thickness Indicator					Bi-weekly
8.	Visual Checklist					Monthly
9.	Repeat/Reject Analysis					Quarterly
10.	Compression					Semi-annually
11.	Geometry Calibration (Tomosynthesis Option)					Semi-annually
12.	Diagnostic Review Workstation QC (NA if only hardcopy read)					See Lorad QC Manual

Medical Physicist's Recommendations for Quality Improvement

Important: / 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. / 2. Use any QC manual version that is applicable to the Lorad model surveyed. / 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) / 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) / 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day / than the survey date, note the date above. / 6. See the FDA-approved alternative standard for Lorad FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.

## MQSA Requirements
MEDICAL PHYSICIST'S CHECKLIST / MQSA REQUIREMENTS FOR MAMMOGRAPHY EQUIPMENT

Facility Name:	0
Unit Manufacturer:			Model:
Serial number:			Year Mfr:	#VALUE!
Medical Physicist:	Eugene Mah		Room ID:
Signature:			Survey Date:	30-Dec-1899

Feature	FDA Rule / Section	Requirement	Applies / to	Meets FDA Requirements? (if NA, please explain)
Motion of tube-image / receptor assembly	3(i)	The assembly shall be capable of being fixed in any position / where it is designed to operate. Once fixed in any such position, it shall not undergo unintended motion.	S-F & / FFDM
	3(ii)	This mechanism shall not fail in the event of power interruption.	S-F & / FFDM
Image receptor sizes	4(i)	Systems using screen-film image receptors shall provide, at a minimum, for operation with image receptors of 18 x 24 cm and 24 x 30 cm.	S-F
	4(ii)	Systems using screen-film image receptors shall be equipped / with moving grids matched to all image receptor sizes provided.	S-F
	4(iii)	Systems used for magnification procedures shall be capable of operation with the grid removed from between the source and image receptor.	S-F & / FFDM
Beam limitation and / light fields	5(i)	All systems shall have beam limiting devices that allow the / useful beam to extend to or beyond the chest wall edge of the / image receptor.	S-F & / FFDM
	5(ii)	For any mammography system with a light beam that passes through the X-ray beam-limiting device, the light shall provide an average illumination of not less than 160 lux (15 ft-candles) at 100 cm or the maximum source-image receptor distance (SID), whichever is less.	S-F & / FFDM (except Fischer)
Magnification	6(i)	Systems used to perform noninterventional problem-solving / procedures shall have radiographic magnification capability / available for use by the operator.	S-F & / FFDM
	6(ii)	Systems used for magnification procedures shall provide, at a minimum, at least one magnification value within the range of 1.4 to 2.0.	S-F & / FFDM
Focal spot selection	7(i)	When more than one focal spot is provided, the system shall indicate, prior to exposure, which focal spot is selected.	S-F & / FFDM
	7(ii)	When more than one target material is provided, the system shall indicate, prior to exposure, the preselected target material.	S-F & / FFDM
	7(iii)	When the target material and/or focal spot is selected by a system algorithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure.	S-F & / FFDM
Application of compression	8(i)(A)	Each system shall provide an initial power-driven compression activated by hands-free controls operable from both sides of the patient.	S-F & / FFDM
	8(i)(B)	Each system shall provide fine adjustment compression controls operable from both sides of the patient.	S-F & / FFDM
Compression paddle	8(ii)(A)	Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system.	S-F & / FFDM
	8(ii)(B)	The compression paddle shall be flat and parallel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when compression is applied.	S-F & / FFDM (except Fischer)
	8(ii)(C)	Equipment intended by the manufacturer’s design to not be flat and parallel to the breast support table during compression shall meet the manufacturer’s design specifications and maintenance requirements.	S-F & / FFDM
	8(ii)(D)	The chest wall edge of the compression paddle shall be straight and parallel to the edge of the image receptor.	S-F & / FFDM
	8(ii)(E)	The chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image.	S-F & / FFDM
Technique factor selection and display	9(i)	Manual selection of mAs or at least one of its component parts (mA and/or time) shall be available.	S-F & / FFDM
	9(ii)	The technique factors (kVp and either mA and seconds or mAs) to be used during an exposure shall be indicated before the exposure begins, except when AEC is used, in which case the technique factors that are set prior to the exposure shall be indicated.	S-F & / FFDM
	9(iii)	Following AEC mode use, the system shall indicate the actual kVp and mAs (or mA and time) used during the exposure.	S-F & / FFDM
Automatic exposure control	10(i)	Each screen-film system shall provide an AEC mode that is operable in all combinations of equipment configuration provided, e.g., grid, non-grid; magnification, nonmagnification; and various target-filter combinations.	S-F
	10(ii)	The positioning or selection of the detector shall permit flexibility in the placement of the detector under the target tissue. The size and the available positions of the detector shall be clearly indicated at the X-ray input surface of the breast compression paddle. The selected position of the detector shall be clearly indicated.	S-F
	10(iii)	The system shall provide means for the operator to vary the selected optical density from the normal (zero) setting.	S-F
X-ray film*	11	The facility shall use X-ray film for mammography that has been designated by the film manufacturer as appropriate for mammography.	S-F
Intensifying screens*	12	The facility shall use intensifying screens for mammography that have been designated by the screen manufacturer as appropriate for mammography and shall use film that is matched to the screen’s spectral output as specified by the manufacturer.	S-F
Film processing solutions*	13	For processing mammography films, the facility shall use chemical solutions that are capable of developing the films used by the facility in a manner equivalent to the minimum requirements specified by the film manufacturer.	S-F
Lighting*	14	The facility shall make special lights for film illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the interpreting physicians.	S-F & / FFDM (for hardcopy comparison)
Film masking devices*	15	Facilities shall ensure that film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are available to all interpreting physicians interpreting for the facility.	S-F & / FFDM (for hardcopy comparison)
* NA is acceptable for new units at existing facilities if these were previously evaluated and have not changed

## Sheet1
1					Revision 2.0-20160302					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Mammography System Compliance Inspection					Mammography System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah				kV Set	mAs Set	Add Filt	Target	Filter	kVp	ms	Dose (mGy)	Dose rate (mGy/s)	Corr kV
10				Facility:					Site Number:					Facility:				Site Number:					Facility:					24	50	0.0
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:					24	50	0.0
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:					24	50	0.2
13				Survey ID:					Number of X-ray Tubes:					Survey ID:				Number of X-ray Tubes:					Survey ID:					24	50	0.2
14					Room Number:				Accession Number:					Room Number:					24	50	0.3
15			X-Ray Generator					Site Number:					24	50	0.3
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:					24	50	0.4
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:					24	50	0.4
18				Max kVp:					Max mA:					Model:				Serial Number:					Number of X-ray Tubes:					25	50	0.0
19					Max kVp:				Max mA:					Accesssion Number:					25	50	0.0
20			X-Ray Tube 1					X-Ray Generator					25	50	0.2
21				Tube Designation/Use:					Target:					X-Ray Tube 1					Target:					Manufacturer:					25	50	0.2
22				Manufacture Date:					Tube Designation/Use:					Model:					25	50	0.3
23			Insert					Filter:					Manufacture Date:					Max kVp:					25	50	0.3
24				Manufacturer:					Insert					Filter:					Manufacture Date:					25	50	0.4
25				Model:					Manufacturer:					Serial Number:					25	50	0.4
26				Serial Number:				Focal Spot Sizes (mm)					Model:					Max mA:					26	50	0.0
27			Housing					Large:					Serial Number:				Focal Spot Sizes (mm)					X-Ray Tube 1					28	20	0.0
28				Manufacturer:					Small:					Housing					Large:					Tube Designation/Use:					28	50	0.0
29				Model:				Filtration	Inherent:					Manufacturer:				Small:					Manufacture Date:					28	50	0.0
30				Serial Number:					Added:					Model:					Manufacturer:					28	50	0.0
31					Serial Number:				Filtration					Model:					28	50	0.0
32					Inherent:					Serial Number:					28	50	0.3
33					Technique Chart					Added:					Manufacturer:					28	50	0.3
34					Is this a breast tomosynthesis unit					Model:					28	50	0.4
35		Compressed	Fatty Breast			50% Fatty - 50% Dense Breast			Dense Breast					Is this unit used for screening only?					Serial Number:					28	50	0.4
36		Breast					Large:					28	50	0.5
37		Thickness	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs					Inspection Results					Small:					28	50	0.5
38		<3 cm					Enter 1 for YES, 2 for NO, 3 for NA					Target 1:					28	100	0.0
39		3 to 5 cm					DHEC Registration sticker is present, clearly visible and legible					Target 2:					28	300	0.0
40		5 to 7 cm					DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Filter 1:					30	50	0.0
41		>7 cm					Radiation warning label posted on the generator control panel					Filter 2:					32	50	0.0
42					Operator manuals are available.					Filter 3:					32	50	0.0
43					Inspection Results					Monthly radiation monitoring reports are posted.					Inherent:					32	50	0.5
44		Rule Number					Compliance					Added					32	50	0.5
45		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Mammography Unit Assembly Evaluation					Technique Chart					32	50	0.6
46		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD				Free-standing unit is mechanically stable					Fatty Breast					32	50	0.6
47		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD				All moving parts move smoothly, without obstructions to motion					Target/Filter:					32	50	0.7
48				Operator manuals are available.					TBD				All locks and detents work properly					kVp:					32	50	0.7
49				Monthly radiation monitoring reports are posted.					TBD				Image receptor holder assembly is free from vibrations					Density:					34	50	0.0
50					Mammography Unit Assembly Evaluation					Image receptor slides smoothly into holder assembly					Target/Filter:					34	50	0.0
51				Free-standing unit is mechanically stable					TBD				Image receptor is held securely by assembly in any orientation					kVp:					34	50	0.5
52				All moving parts move smoothly, without obstructions to motion					TBD				Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm					Density:					34	50	0.5
53				All locks and detents work properly					TBD				Patient or operator is not exposed to sharp or rough edges, or other hazards					Target/Filter:					34	50	0.6
54				Image receptor holder assembly is free from vibrations					TBD				Operator technique control charts are posted					kVp:					34	50	0.6
55				Image receptor slides smoothly into holder assembly					TBD				Operator protected during exposure by adequate radiation shielding					Density:					34	50	0.7
56				Image receptor is held securely by assembly in any orientation					TBD				All indicator lights working properly					Target/Filter:					34	50	0.7
57				Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm					TBD				Auto decompression can be overridden to maintain compression (and status displayed)					kVp:					28	50	0.0
58				Patient or operator is not exposed to sharp or rough edges, or other hazards					TBD				Manual emergency compression release can be activated in the event of a power failure					Density:					28	50	0.0
59				Operator technique control charts are posted					TBD					50/50 Breast					28	50	0.4
60				Operator protected during exposure by adequate radiation shielding					TBD					MQSA Equipment Standards					Target/Filter:					28	50	0.4
61				All indicator lights working properly					TBD				Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position					kVp:					28	50	0.5
62				Autodecompression can be overridden to maintain compression					TBD				Tube-receptor assembly locking mechanism does not fail during power interruption					Density:					28	50	0.5
63				Manual emergency compression release in the event of a power failure					TBD				Image Receptor Sizes 18 x 24 & 24 x 30 are available					Target/Filter:					28	50	0.6
64					Moving grids are present for all image receptor sizes					kVp:					28	50	0.6
65					Magnification can be done without grid					Density:					30	50	0.0
66					Beam limiting device allows useful beam to extend beyond chest wall edge of receptor					Target/Filter:					30	50	0.0
67		Date:					Inspector:	Eugene Mah		1	Average illumination of the light field is at least 160 lux					kVp:					30	50	0.5
68		Room Number					Survey ID:				Magnification mode is available for non-interventional problem solving procedures					Density:					30	50	0.5
1					Medical University of South Carolina			System provides magnification factor between 1.4 – 2.0					Target/Filter:					30	50	0.6
2					Inspection Results					Mammography System Compliance Inspection			Prior indication of selected focal spot is provided					kVp:					30	50	0.6
3		Rule Number					MQSA Equipment Standards					Prior indication of selected target is provided					Density:					30	50	0.7
4		900.12(b)(3i)		Tube-receptor can be fixed into any operating position and does not move					TBD				Post exposure display of target/focal spot is provided					Dense Breast					30	50	0.7
5		900.12(b)(3ii)		Tube-receptor locking mechanism does not fail during power interruption					TBD				Initial power driven compression can be activated hands-free from both sides					Target/Filter:					32	50	0.0
6		900.12(b)(4i)		Image Receptor Sizes 18 x 24 & 24 x 30 are available					TBD				Fine adjustment of compression can be activated from both sides					kVp:					32	50	0.0
7		900.12(b)(4ii)		Moving grids are present for all image receptor sizessizes					TBD				Compression paddles match the sizes of all image receptors					Density:					32	50	0.5
8		900.12(b)(4iii)		Magnification can be done without grid					TBD				Compression paddles do not flex by more than 1 cm					Target/Filter:					32	50	0.5
9		900.12(b)(5i)		Beam limiting device allows useful beam to extend beyond chest wall edge					TBD				Compression paddles are flat and parallel to the breast support table					kVp:					32	50	0.6
10		900.12(b)(5ii)		Average illumination of the light field is at least 160 lux					YES				Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor					Density:					32	50	0.6
11		900.12(b)(6i)		Magnification mode available for non-interventional problem solving					TBD				Chest wall edge may be bent upwards for patient comfort and does not appear in the image					Target/Filter:					32	50	0.7
12		900.12(b)(6ii)		System provides magnification factor between 1.4 – 2.0					TBD				X-ray field extends beyond the chest wall edge of the receptor					kVp:					32	50	0.7
13		900.12(b)(7i)		Prior indication of selected focal spot is provided					TBD				Manual selection of mAs or mA and/or time is available					Density:					34	50	0.0
14		900.12(b)(7i)		Prior indication of selected target is provided					TBD				Prior indication of selected technique factors is provided					Target/Filter:					34	50	0.0
15		900.12(b)(7iii)		Post exposure display of target/focal spot is provided					TBD				Post exposure display of technique factors is provided for AEC exposures					kVp:					34	50	0.5
16		900.12(b)(8)(iB)		Initial power driven compression can be activated hands-free from both sides					TBD				AEC mode is operable in all combinations of equipment configuration					Density:					34	50	0.5
17				Fine adjustment of compression can be activated from both sides					TBD				AEC sensor positioning is flexible					34	50	0.6
18		900.12(b)(8)(iiA)		Compression paddles match the sizes of all image receptors					TBD				AEC sensor position and size is indicated on the compression paddle					MGD:					34	50	0.6
19		900.12(b)(8)(iiB)		Compression paddles do not flex by more than 1 cm					TBD				Selected position of the AEC sensor is clearly indicated					mGy/mAs:					34	50	0.7
20		900.12(b)(8)(iiC)		Compression paddles are flat and parallel to the breast support table					TBD				AEC density control setting is provided					mGy/s:					34	50	0.7
21		900.12(b)(8)(iiD)		Chest wall edge of the compression paddle is straight and parallel to receptor					TBD			3	X-ray film used is appropriate for mammography					Image Quality (Printer)					36	50	0.0
22		900.12(b)(8)(iiE)		Chest wall edge may be bent upwards and does not appear in the image					TBD			3	Intensifying screens used are appropriate for mammography and matched to the film used					kV:					38	50	0.0
23				X-ray field extends beyond the chest wall edge of the receptor					TBD			3	Film processing chemicals are appropriate for the film being used					mA:					28	50	0.0
24		900.12(b)(9i)		Manual selection of mAs or mA and/or time is available					TBD				Lights for film illumination i.e. hot lights are available					Bkg Density:					28	50	0.0
25		900.12(b)(9ii)		Prior indication of selected technique factors is provided					TBD				Film masking devices are available to all interpreting physicians					Inside disk:					28	50	0.4
26		900.12(b)(9iii)		Post exposure display of technique factors is provided for AEC exposures					TBD					Outside disk:					28	50	0.4
27		900.12(b)(10i)		AEC mode is operable in all combinations of equipment configuration					TBD					Mammography Phototimer Technique Chart					Difference:					28	50	0.5
28		900.12(b)(10ii)		AEC sensor positioning is flexible					TBD					Fibers:					28	50	0.5
29				AEC sensor position and size is indicated on the compression paddle					TBD			Compressed	Fatty Breast			50% Fatty - 50% Dense Breast			Dense Breast					Specks:					28	50	0.6
30				Selected position of the AEC sensor is clearly indicated					TBD			Breast					Masses:					28	50	0.6
31		900.12(b)(10iii)		AEC density control setting is provided					TBD			Thickness	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs			Image Quality (Acq Unit)					30	50	0.0
32		900.12(b)(11)		X-ray film used is appropriate for mammography					NA			<3 cm					2D kV:					30	50	0.0
33		900.12(b)(12)		Intensifying screens used are appropriate for mammography and matched to film					NA			3 to 5 cm					mA:					30	50	0.4
34		900.12(b)(13)		Film processing chemicals are appropriate for the film being used					NA			5 to 7 cm					Exp Ind:					30	50	0.4
35		900.12(b)(14)		Lights for film illumination i.e. hot lights are available					TBD			>7 cm					Fibers:					30	50	0.5
36		900.12(b)(15)		Film masking devices are available to all interpreting physicians					TBD					Specks:					30	50	0.5
37					Comments					Compressed	Fatty Breast			50% Fatty - 50% Dense Breast			Dense Breast					Masses:					30	50	0.6
38					Breast					Mag kV:					30	50	0.6
39					Thickness	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs	Target/Filter	kVp	mAs			mA:					32	50	0.0
40					<3 cm					Exp Ind:					32	50	0.0
41					3 to 5 cm					Fibers:					32	50	0.5
42					5 to 7 cm					Specks:					32	50	0.5
43					>7 cm					Masses:					32	50	0.6
44					3D kV:					32	50	0.6
45					Inspection Results					mA:					32	50	0.7
46					Compression Force Indicator					Exp Ind:					32	50	0.7
47					Indicated	Measured	After 20s	Slip					Fibers:					34	50	0.0
48					Power					Specks:					34	50	0.0
49					Manual					Masses:					34	50	0.5
50					Criteria:	Minimum compression force – 25 lbs					34	50	0.5
51					Maximum compression force – 45 lbs					SNR:					34	50	0.6
52					CNR:					34	50	0.6
53					Compression Thickness Indicator					34	50	0.7
54					BR-12 (cm)	Indicated	Difference					Comment Page 1					34	50	0.7
55					1					38	50	0.0
56					2					38	50	0.0
57					4			Mean:					38	50	0.6
58					4			StDev:					38	50	0.6
59					4					38	50	0.7
60					4				Criteria:	Difference between indicated and actual thickness is ± 5mm					38	50	0.7
61					6					38	50	0.8
62					8					38	50	0.8
63
64					Paddle flex is less than 1 cm				Compression thickness indicator:
65					Overall unit assembly
66
67		Date:					Inspector:	Eugene Mah		Artifact Evaluation
68		Room Number					Survey ID:				Target:
1					Medical University of South Carolina			Filter:
2					Inspection Results					Mammography System Compliance Inspection			Focal spot:
3		Compression Force Indicator					kV:
4				Indicated	Measured	After 20s	Slip					mAs:
5			Power					EI:
6			Manual					Acceptable:					Acceptable:
7			Criteria:	Minimum compression force – 25 lbs
8				Maximum compression force – 45 lbs
9					Comments:
10		Compression Thickness Indicator					Do not exceed 390 chars
11			BR-12 (cm)	Indicated	Difference				Paddle flex is less than 1 cm:					New:
12			1					Overall unit assembly:
13			2					Compression thickness indicator:					Light Field Brightness				Enter 1 for YES, 2 for NO, 3 for NA
14			4			Mean:					Measurements (lux)					Average Illumination
15			4			StDev:					1st quad.	2nd quad.	3rd quad.	4th quad.		(lux)
16			4			Criteria:	Difference between indicated and actual thickness is ± 5mm
17			4
18			6					Light Field/X-ray Field Alignment
19			8					SID (cm):	65
20					24x30	18x24			18x24	Small breast
21		Artifact Evaluation					24x30	Left	Center	Right	Tomo	24x30
22			Receptor:					Left:
23			Target:					Right:
24			Filter:					Anterior:
25			Focal spot:					Chest Wall:
26			kV:					Enter deviation in cm
27			mAs:
28			EI:					% Deviation
29			Acceptable:					24x30	18x24			18x24	Small breast
30			Criteria:	No visible artifacts in the images					24x30	Left	Center	Right	Tomo	24x30
31			Comments:					Left:
32					Right:
33					Anterior:
34					Chest Wall:
35		Light Field Brightness					Left-Right
36			Measurements (lux)					Average Illumination					Ant-Chest
37			1st quad.	2nd quad.	3rd quad.	4th quad.		(lux)					Criteria:	Sum of left and right edge deviation is less than 2% SID
38					Sum of anterior and chest wall edge deviation is less than 2% SID
39			Criteria:	Average light field illumination >= 160 lux.					Enter deviation in cm. Positive deviation indicates x-ray field extends beyond light field
40					X-ray Field/Image Receptor Alignment
41		Light fField/X-ray Field Alignment				X-ray Field/Image Receptor Alignment
42		SID (cm):					24x30	18x24			24x30	Small breast
43		65	Light Field/X-ray Field % Deviation					24x30	Left	Center	Right	Tomo	24x30
44			24x30	18x24			24x30	Small breast					Left:
45			24x30	Left	Center	Right	Tomo	Tomo					Right:
46		Left:					Anterior:
47		Right:					Chest Wall:
48		Anterior:					Paddle:
49		Chest Wall:					Enter deviation in cm. Positive deviation indicates x-ray field extends beyond image receptor
50		Left-Right:
51		Ant-Chest:					% Deviation
52		Acceptable:					24x30	18x24			24x30	Small breast
53			X-ray Field/Image Receptor % Deviation					24x30	Left	Center	Right	Tomo
54		Left:					Left:
55		Right:					Right:
56		Anterior:					Anterior:
57		Chest Wall:					Chest Wall:
58		Paddle:					Paddle:
59		Acceptable:					X-ray field extends beyond the chest wall edge of the receptor
60		Criteria:	Sum of left/right and chest wall/anterior deviation is less than 2% SID					Chest wall edge of the compression paddle is not visible in the image
61			X-ray field does not extend beyond image receptor by more than 2% SID
62			X-ray field extends beyond the chest wall edge of image receptor.					Criteria:	X-ray field does not extend beyond image receptor by more than 2% SID
63			X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side					X-ray field extends beyond the chest wall edge of image receptor.
64			X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side					X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side
65			Difference between chest wall edge of the receptor and compression paddle is less than 1% SID					X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side
66					Difference between chest wall edge of the receptor and compression paddle is less than 1% SID
67		Date:					Inspector:	Eugene Mah
68		Room Number					Survey ID:			System Resolution
1					Medical University of South Carolina			Mode (2D/3D):
2					Inspection Results					Mammography System Compliance Inspection			Target:
3		AEC Thickness Tracking – 2D					kV:
4		AEC Mode:			AEC Position:					mAs:
5			Thickness					Reported					lp/mm:
6			(cm)	Filter	kV	mAs	EI	MGD (mGy)	Deviation					Acceptable:
7			2					Criteria:	Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode
8			4					Limiting system resolution must be 3 lp/mm or higher in 3D mode
9			4					AEC Thickness Tracking – 2D
10			4					AEC Mode:			AEC Position:			Dose Table:	5
11			4					Thickness					Corrected	Reported
12			6					(cm)	Filter	kV	mAs	EI	EI	MGD (mGy)	Deviation
13			8					2
14			4					4
15					Average:					4	0	0
16			Criteria:	Mean pixel value of each image is within 10% of the average pixel value					4	0	0
17					4	0	0
18		AEC Thickness Tracking – 3D					6
19		AEC Mode:			AEC Position:					8
20			Thickness					Reported					Mag:	4
21			(cm)	Filter	kV	mAs	EI	MGD (mGy)	Deviation					Average:
22			2					Criteria:	Mean pixel value of each image is within 10% of the average pixel value
23			4
24			4					AEC Thickness Tracking – 3D
25			4					AEC Mode:			AEC Position:			Dose Table:	0
26			4					Thickness					Corrected	Reported
27			6					(cm)	Filter	kV	mAs	EI	EI	MGD (mGy)	Deviation
28			8					2
29					Average:					4
30			Criteria:	Mean pixel value of each image is within 10% of the average pixel value					4	0	0
31					4	0	0
32		AEC Density Control					4	0	0
33		AEC Mode:			AEC Position:			kV:					6
34					Reported					8
35			Density	mAs	EI	MGD (mGy)	Change	Acceptable					Average:
36			-3					Criteria:	Mean pixel value of each image is within 10% of the average pixel value
37			-2
38			-1
39			0					AEC Density Control
40			1					AEC Mode:	0		AEC Position:	0	Reported				Selenia		Selenia		Selenia Dimensions
41			2					Density	kV	mAs	EI	MGD (mGy)	Change			Low	High	Low	High	Low	High
42			3					-3					0.5	0.61	0.5	0.61	0.56	0.66
43			4					-2					0.63	0.77	0.63	0.77	0.66	0.78
44			Criteria:	Change in pixel value at each density step is within specified ranges					-1					0.77	0.94	0.77	0.94	0.78	0.92
45					0
46		System Resolution					1					1.04	1.27	1.04	1.27	1.06	1.24
47			Mode (2D/3D):					2					1.17	1.43	1.17	1.43	1.22	1.43
48			Target:					3					1.31	1.6	1.31	1.6	1.40	1.64
49			kV:					Acceptable:					4					1.44	1.76	1.44	1.76	1.61	1.89
50			mAs:
51			lp/mm:					Criteria:	Change in pixel value at each density step is within specified ranges
52			Criteria:	Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode
53					Mean Glandular Dose – 2D
54					Phantom:			kV:
55					Target/Filter:			mA:
56					Reported		Air KERMA
57					mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:
58					ESE (mGy):
59					DGN (mrad/R):
60					MGD (mGy):
61					Prev MGD (mGy):
62					Average:					% Change:
63					Coeff of Var:
64
65					Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:
66					Deviation from reported MGD is less than15%					mAs at 3 mGy:
67		Date:					Inspector:	Eugene Mah
68		Room Number					Survey ID:			Mean Glandular Dose – 3D
1					Medical University of South Carolina			Phantom:			kV:
2					Mammography System Compliance Inspection			Target/Filter:			mA:
3		Mean Glandular Dose – 2D					Reported		Air KERMA
4			Phantom:			kV:					mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:
5			Target/Filter:			mA:					ESE (mGy):
6					Reported		Air KERMA					DGN (mrad/R):
7				mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:					MGD (mGy):
8					ESE (mGy):					Prev MGD (mGy):
9					DGN (mrad/R):					Average:					% Change:
10					MGD (mGy):					Coeff of Var:
11					Prev MGD (mGy):
12			Average:					% Change:					Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:
13			Coeff of Var:					Deviation from reported MGD is less than15%					mAs at 3 mGy:
14			Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:					MGD for 2D + 3D acquisitions is less than 3 mGy
15				Deviation from reported MGD is less than15%					mAs at 3 mGy:					2D + 3D MGD:	#VALUE!
16
17					kVp Accuracy and Output
18		Mean Glandular Dose – 3D					Meter:				Calibration Date:
19			Phantom:			kV:					Serial #:				Calibration Due:
20			Target/Filter:			mA:
21					Reported		Air KERMA					Measured Values
22				mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
23					ESE (mGy):					24	50
24					DGN (mrad/R):					25	50
25					MGD (mGy):					26	50
26					Prev MGD (mGy):					28	50
27			Average:					% Change:					30	50
28			Coeff of Var:					32	50
29			Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:					34	50
30				Deviation from reported MGD is less than15%					mAs at 3 mGy:					Criteria:	kV accuracy is <= 5%
31
32					2D + 3D MGD:	#VALUE!					Measured Values
33					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
34		Image Quality Evaluation					34	50
35			Printer			Acq Workstation				Review Workstation					30	50
36			kV:			Normal	Mag	Tomo		Left	Right					32	50
37			mA:		kV:					Workstation:					34	50
38			Background Density:		mA:				Fibers:					36	50
39			Inside disk:		EI:				Specks:					38	50
40			Outside disk:		Fibers:				Masses:					Criteria:	kV accuracy is <= 5%
41			Difference:		Specks:					Workstation:
42			Fibers:		Masses:				Fibers:					Measured Values
43			Specks:		Criteria:				Specks:					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
44			Masses:		2D – Must be able to see 5 fibers, 4 speck groups, 4 masses				Masses:					28	50
45					3D – Must be able to see 4 fibers, 3 speck groups, 3 masses					30	50
46		Signal to Noise Ratio					32	50
47			kV:				Current	Previous	Change	Acceptable					34	50
48			mA:			SNR:				Pass					38	50
49			Target:			CNR:				NA					Criteria:	kV accuracy is <= 5%
50			Filter:
51			Criteria:	SNR must be >= 40					kV, Exposure, and Output Reproducibility
52				CNR does not change by more than 15%
53					Measured Values
54					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
55					28	50
56					28	50
57					28	50
58					28	50
59					Mean:
60					Std Dev:
61					Coeff of Var:
62					Previous:
63					Criteria:	Coefficient of variation <= 2%
64					Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
65					Minimum acceptable output is 2 mGy/s at 28 kVp for W target
66					Exposure Linearity
67		Date:					Inspector:	Eugene Mah					Measured Values
68		Room Number					Survey ID:				Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
1					Medical University of South Carolina					28	20
2					Inspection Results					Mammography System Compliance Inspection					28	50
3		Meter:					Calibration Date:					28	100
4		SN:					Calibration Due:					28	300
5		kVp Accuracy					Criteria:	Coefficient of linearity is <= 10%					Coeff of Linearity:
6		Target:				Target:				Target:
7		Filter:				Filter:				Filter:					Unit Can Maintain Output for 3 sec at 28 kVp
8		mAs:	50			mAs:	50.00			mAs:	50.00
9		Indicated	Measured			Indicated	Measured			Indicated	Measured			Half Value Layer – 2D		1	Compression paddle used?
10		kVp	kVp	Error (%)		kVp	kVp	Error (%)		kVp	kVp	Error (%)			Target/Filter	/	/	/	/	/	/	/	/
11		24				34				28					kV	24	25	28	32	28	30	32	34
12		25				30				30					Al (mm)	Air Kerma (mGy)
13		26				32				32					0.0
14		28				34				34					0.0
15		30				36				38					HVL
16		32				38					0 mm
17		34					Min	0.27	0.28	0.31	0.35	0.31	0.33	0.35	0.37
18			Acceptable:				Acceptable:				Acceptable:				Max	0.36	0.37	0.40	0.44	0.47	0.49	0.51	0.53
19
20			Criteria:	kVp accurate to within 5% of indicated
21					Criteria:	HVL is between the minimum and maximum limits
22		kVp and Output Reproducibility					Half Value Layer – 3D
23		Target:		Filter:					Target/Filter	/	/	/	/	/
24		mAs:	50					kV	28	30	32	34	38
25		Indicated	Measured					Al (mm)	Air Kerma (mGy)
26		kVp	kVp	mGy	mGy/mAs	mGy/s					0.0
27		28					0.0
28					HVL
29					0 mm
30					Min	0.31	0.33	0.35	0.37	0.41
31		Mean:					Max	0.40	0.42	0.44	0.46	0.57
32		Std Dev:
33		Coeff of Var:					Criteria:	Mo/Rh targets – HVL is between the minimum and maximum limits
34		Acceptable:					W targets – HVL is above the minimum limit
35			Criteria:	Coefficient of variation <= 2%					Image Quality Evaluation
36				Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target					Printer	Model:		SN:			QC Manual Version:
37				Minimum acceptable output is 2 mGy/s at 28 kVp for W target
38					Previous				Previous
39		Exposure Linearity					kV:	0			Fibers:
40		Target:		Filter:					mAs:	0			Specks:
41					Background Density:				Masses:
42		Set mAs	kVp	mGy	mGy/mAs	mGy/s					Inside disk:				Acceptable:
43		20					Outside disk:
44		50					Difference:
45		100
46		300					Workstation
47				Coeff of Linearity:					Workstation:				Workstation:
48			Criteria:	Coefficient of linearity is <= 10%					Left	Right			Left	Right
49					Fibers:				Fibers:
50		Half Value Layer					Specks:				Specks:
51		Target/Filter	/	/	/	/	/	/	/	/					Masses:				Masses:
52		kV	24	25	28	32	28	30	32	34
53		HVL					Acquisition Unit
54		0 mm					2D	Previous	Mag	Previous	3D	Previous
55		Min	0.27	0.28	0.31	0.35	0.31	0.33	0.35	0.37				kV:	0
56		Max	0.36	0.37	0.40	0.44	0.47	0.49	0.51	0.53				mAs:	0
57		Acceptable:					EI:	0
58					Fibers:
59		Target/Filter	/	/	/	/	/					Specks:
60		kV	28	30	32	34	38					Masses:
61		HVL
62		0 mm					Criteria:	2D – Must be able to see 5 fibers, 4 speck groups, 4 masses
63		Min	0.31	0.33	0.35	0.37	0.41					3D – Must be able to see 4 fibers, 3 speck groups, 3 masses
64		Max	0.40	0.42	0.44	0.46	0.57					Signal to Noise
65		Acceptable:					Mean	SD
66			Criteria:	Mo/Rh targets – HVL is between the minimum and maximum limits					kV:	0		In acrylic:
67		Date:					Inspector:	Eugene Mah			mAs:	0		Background:
68		Room Number					Survey ID:				Target:
				Filter:				Current	Previous	Change	Acceptable
				SNR:				Pass
				CNR:				NA
				Criteria:	SNR must be >= 40
				CNR does not change by more than 15%

				Previous Year Comments
				Comments:
				Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
				Additional Comments:
				Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
				Additional Comments:
				Do not exceed 390 chars		0			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
				Additional Comments:
				Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
				Additional Comments:
				Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
				Additional Comments:
				Do not exceed 390 chars		0
				Additional Comments:
				Do not exceed 390 chars		0
				Additional Comments:
				Do not exceed 390 chars		0
				Additional Comments:
				Do not exceed 390 chars		0

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast					Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast					Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
	kVp					kVp					kVp
HVL	23	24	25	26	27	28	29	30	31	HVL	23	24	25	26	27	28	29	30	31	HVL	23	24	25	26	27	28	29	30	31
0.23	116					0.23					0.23
0.24	121	124					0.24					0.24
0.25	126	129	131					0.25					0.25
0.26	130	133	135	138					0.26					0.26
0.27	135	138	140	142	143					0.27					0.27
0.28	140	142	144	146	147	149				0.28			149	151	154					0.28			150	155	159
0.29	144	146	148	150	151	153	154			0.29			154	156	158	159				0.29			155	160	164	168
0.3	149	151	153	155	156	157	158	159		0.3			158	160	162	162	163			0.3			160	164	168	172	176
0.31	154	156	157	159	160	161	162	163	164	0.31			163	164	166	166	167	167		0.31			165	168	172	174	180	182
0.32	158	160	162	163	164	166	167	168	168	0.32			167	169	171	171	171	172	172	0.32			169	173	177	181	184	186	188
0.33	163	165	166	168	169	170	171	173	173	0.33			171	173	175	176	176	176	176	0.33			174	178	181	185	188	190	192
0.34	168	170	171	172	173	174	175	176	177	0.34			176	178	179	179	180	180	180	0.34			179	183	186	190	193	195	196
0.35		174	175	176	177	178	179	180	181	0.35			180	181	183	183	184	185	185	0.35			184	187	190	194	197	199	201
0.36			179	181	182	183	184	185	185	0.36			185	186	187	187	188	188	189	0.36			189	192	195	198	201	204	205
0.37				185	186	187	188	189	190	0.37			189	190	191	191	192	193	193	0.37			193	196	199	202	205	207	209
0.38					190	191	192	193	194	0.38			193	194	196	196	197	197	197	0.38			198	201	204	207	209	211	213
0.39					196	197	198	198	0.39			198	199	200	200	201	201	202	0.39			203	206	208	211	214	216	217
0.4					201	202	203	0.4			202	203	204	204	205	205	206	0.4			208	211	213	216	218	220	221
0.41					206	207	0.41			206	207	208	208	209	209	210	0.41			213	215	217	220	222	224	225
0.42					211	0.42			211	211	212	212	213	213	214	0.42			218	220	222	224	226	228	229

Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast					Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
	kVp					kVp
HVL	22	23	24	25	26	27	28	29	30	31	32	33	HVL	27	28	29	30	31	32	33	34	35	36	37	38	39
0.3	152	157	163	166	170	173	175	177	179	182	184	187	0.4	222	226	229	231	234	236	239	241	244	246	248	250	252
0.325	163	169	174	177	181	183	186	188	190	192	195	197	0.425	233	236	239	242	244	246	248	251	253	256	258	260	262
0.35	175	180	185	188	191	194	196	198	200	202	205	207	0.45	244	247	249	252	254	256	258	260	263	265	267	269	271
0.375	186	191	196	199	202	205	207	209	211	213	215	218	0.475	254	257	260	262	264	266	268	270	273	275	277	279	281
0.4	198	203	207	210	213	215	217	219	221	223	226	228	0.5	265	267	270	272	274	276	278	280	282	284	286	288	290
0.425	209	214	218	221	224	226	228	230	232	234	236	238	0.525	275	278	280	282	284	286	288	290	292	294	296	298	300
0.45	221	226	230	232	235	237	238	240	242	244	246	248	0.55	286	288	290	292	294	296	298	299	301	303	305	307	309
0.475	233	237	241	243	245	247	249	251	253	254	256	258	0.575	296	298	3090	302	304	305	307	309	311	313	315	317	318
0.5	244	248	252	254	256	258	260	261	263	265	267	269	0.6	306	308	310	312	313	315	317	319	320	322	324	326	328
0.525	256	260	263	265	267	269	270	272	273	275	277	279	0.625	316	318	320	322	323	325	326	328	330	332	333	335	337
0.55	267	271	274	276	278	279	281	282	284	285	287	289	0.65	326	328	330	331	333	334	336	338	339	341	343	344	346
0.575	279	282	285	287	288	290	291	292	294	296	297	299	0.675	336	338	339	341	342	344	345	347	349	350	352	354	355
0.6	290	293	296	297	299	300	301	303	304	306	308	310	0.7	346	348	349	350	352	353	355	356	358	359	361	363	364
0.625	301	304	306	308	309	310	312	313	315	316	318	320	0.725	356	357	358	360	361	362	364	365	367	368	370	372	373
0.65	312	314	317	318	320	321	322	323	325	326	328	330	0.75	365	367	368	369	370	372	373	375	376	378	379	381	382
0.675	322	325	327	328	330	331	333	333	335	336	338	340	0.775	374	376	377	378	379	381	382	383	385	386	388	390	391
0.7	333	335	337	339	340	341	342	343	345	346	348	350	0.8	384	385	386	387	388	390	391	392	394	395	397	398	400
0.725	342	345	347	348	349	351	352	353	354	356	358	360	0.825	393	394	395	396	397	399	400	401	403	404	406	407	408
0.75	352	355	357	358	359	360	361	363	364	366	368	369	0.85	402	403	404	405	406	407	409	410	411	413	414	416	417
0.775	361	365	367	368	369	370	371	372	374	375	377	379	0.875	410	411	412	413	415	416	417	418	420	421	423	424	425
0.8	369	374	376	377	378	379	380	382	383	385	387	389	0.9	419	420	421	422	423	424	425	427	428	429	431	432	434

Exposure to Dose Conversion Factors W/Al for 4.2cm 50/50 Breast
	kVp
HVL	25	26	27	28	29	30	31	32	33	34	35	36	37	38	39	40	41	42	43	44	45	46	47	48	49
0.2	120	125	133	138	143	148	156	160	165	169	177	181	185	188	195	198	200	203	208	216	219	222	225	231	234
0.25	143	148	155	160	165	169	176	181	185	189	196	200	204	207	213	216	219	221	226	233	236	239	241	247	250
0.3	166	171	177	182	187	191	197	202	206	210	216	220	223	226	232	235	237	239	244	250	253	256	258	264	266
0.35	189	195	200	204	209	213	218	222	226	230	236	239	243	246	250	253	255	257	262	267	270	273	275	280	282
0.4	212	217	222	226	231	235	239	243	247	251	255	258	262	265	269	271	273	275	280	285	287	290	292	296	298
0.45	234	239	244	248	252	256	260	264	268	271	275	278	281	284	287	290	292	294	298	302	304	307	309	313	315
0.5	256	261	265	269	273	277	280	284	288	291	294	297	300	303	306	308	310	312	316	319	321	324	326	329	331
0.55	278	282	286	290	293	297	300	304	307	310	313	316	319	322	324	326	328	330	334	336	339	341	343	345	347
0.6	300	303	307	310	313	317	320	323	326	329	332	335	337	340	342	344	346	348	351	354	356	358	360	362	363
0.65	321	324	327	330	333	336	339	342	345	347	350	353	355	358	360	362	363	365	368	370	372	374	376	378	380
0.7	341	344	347	350	352	355	358	361	363	366	368	370	373	375	377	379	380	382	385	387	389	390	392	394	395
0.75	360	363	366	369	370	374	376	379	381	383	386	388	390	392	394	395	397	398	401	403	404	406	408	409	411
0.8	379	382	385	387	389	392	394	396	398	401	403	405	407	408	410	412	413	414	417	418	420	421	423	424	426
0.85	398	400	403	405	407	409	411	413	415	417	419	421	423	425	426	427	429	430	432	434	435	436	438	439	440
0.9	415	418	420	422	424	46	428	430	432	434	435	437	439	440	442	443	444	445	447	449	450	451	452	454	455
0.95	432	435	437	438	440	442	444	446	448	449	451	453	454	455	457	458	459	460	462	463	464	465	467	468	469
1	448	451	453	454	456	458	460	461	463	465	466	467	469	470	471	472	474	475	476	477	478	479	480	481	482

kV Accuracy/Output		Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV^2					Radcal kV correction for Hologic Selenia/Dimension
Target	kV	kV^2	mGy/mAs	mGy/s		Target	kV	kV^2	mGy/mAs	mGy/s		Target	kV	kV^2	mGy/mAs	mGy/s				Selenia		Dimension
	24	576					34	1156					28	784					Measured	W/Rh	W/Ag	W/Rh	W/Ag	W/Al
Filter	25	625				Filter	30	900				Filter	30	900
	26	676					32	1024					32	1024
	28	784					34	1156					34	1156
	30	900					36	1296					38	1444
	32	1024					38	1444					Slope:	#VALUE!	#VALUE!
	34	1156					Slope:	#VALUE!	#VALUE!				Intercept:	#VALUE!	#VALUE!
		Slope:	#VALUE!	#VALUE!				Intercept:	#VALUE!	#VALUE!
		Intercept:	#VALUE!	#VALUE!
				#VALUE!
Half Value Layer
Target	kV	HVL	Target	kV	HVL	Target	kV	HVL
	24			28			28
Filter	25		Filter	30		Filter	30
	28			32			32
	32			34			34
	34			Slope:	#VALUE!		38
	Slope:	#VALUE!		Intercept:	#VALUE!		Slope:	#VALUE!
	Intercept:	#VALUE!					Intercept:	#VALUE!
Target	Filter	C		Critiera:	Min HVL > kV/100 (no paddle)
Mo	Mo	0.12			Min HVL > kV/100 + 0.03 (with paddle)
Mo	Rh	0.19			Max HVL < kV/100 + C
Rh	Rh	0.22

DGN	#VALUE!
Mo/Mo	Mo/Rh	Rh/Rh	W/Rh	W/Ag	W/Al
#N/A	#N/A	#N/A	#N/A	#N/A	#N/A

Radcal kV Correction for Hologic Selenia
	kV		kV^0	kV^1	kV^2	kV^3	kV^4	kV^5	kV^6
W/Rh	<	27.586	-8375.0727645926	975.9254356043	-37.913729682	0.4908658347	0	0	0
	>=	27.586	-9984.6167916494	1436.5245457141	-82.5051021853	2.3655908176	-0.0338672434	0.0001936869	0
		Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
		Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
	kV		kV^0	kV^1	kV^2	kV^3	kV^4	kV^5	kV^6
W/Ag	<	30.1	-540847.695500773	100186.233642731	-7418.4790179813	274.4766092958	-5.0743695436	0.0375005748	0
	>=	30.1	-11057.7739361992	1297.2285673767	-56.9891889897	1.1115828564	-0.0081233997	0	0
		Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
		Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.

Radcal kV Correction for Hologic Dimension
	kV		kV^0	kV^1	kV^2	kV^3	kV^4
W/Rh	<	26.9	138.88667	-10.72639	0.26216	-0.00082
	>=	26.9	-5009.7751652	605.732006	-27.3018617	0.5467114	-0.0040986

W/Ag	kV		kV^0	kV^1	kV^2	kV^3	kV^4
	<	28.7	296.34185	-31.62925	1.18025	-0.01417
	>=	28.7	4.834469	0.919242

W/Al			kV^0	kV^1	kV^2	kV^3	kV^4
	<	28.7	49.31115	-2.93017	0.07379
		28.7 < kV < 30.1	-24.875	1.8032
	>=	30.1	-4.83461	1.15715

CNR Correction (XX8xxxxx series detectors)
Conventional
Thickness (cm)	0	1	2	3	4	5	6
2	1	1	1	1	1	1	1
4	1	1	1	1	1	1	1
6	1.04	1.02	1.01	1.08	1.35	1.35	1.31
8	1.23	1.2	1.18	1.1	1.82	1.85	1.82
Magnification
Thickness (cm)	0	1	2	3	4
2	1.15	1.15	1.15	1.15	1.5
4	1.15	1.15	1.15	1.15	1.5
6	1.19	1.18	1.18	1.18	1.55
8	1.28	1.24	1.22	1.29	1.67
Stereotactic biopsy			CE2D
Thickness (cm)	0		Thickness	0	1
2	0.91		2	1	0.72
4	1		4	1	0.9
6	1.32		6	1.91	1.71
8	1.88		8	1.81	2.22
Tomo 15 Proj STD				Tomo 15 Proj ENH
Thickness (cm)	0	1		Thickness	0
2	0.7	0.7		2	0.57
4	0.91	0.91		4	0.91
6	1.55	1.55		6	1.68
8	2.78	2.78		8	1.93

## HVLProcessing
24	kVp					28	kVp					28	kVp
mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!
0.0					0.0					0.0
0.2					0.4					0.4

			Data Exp	LN(EXP)					Data Exp	LN(EXP)					Data Exp	LN(EXP)

25	kVp					30	kVp					30	kVp
mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!
0.0					0.0					0.0
0.2					0.5					0.4

			Data Exp	LN(EXP)					Data Exp	LN(EXP)					Data Exp	LN(EXP)

28	kVp					32	kVp					32	kVp
mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!
0.0					0.0					0.0
0.3					0.5					0.5

			Data Exp	LN(EXP)					Data Exp	LN(EXP)					Data Exp	LN(EXP)

32	kVp					34	kVp					34	kVp
mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!		mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!
0.0					0.0					0.0
0.5					0.5					0.5

			Data Exp	LN(EXP)					Data Exp	LN(EXP)					Data Exp	LN(EXP)

34	kVp					38	kVp
mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!					mm Al.	Avg. Exp.	DATA (1/2)	Data Al	kVp		#VALUE!
0.0					0.0
0.5					0.6

			Data Exp	LN(EXP)					Data Exp	LN(EXP)

## Corrected kV
GENERAL
	Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
	Read the true corrected kV in the adjacent cell to the right.
	If that cell is light green, the corrected kV is accurate.
	If the cell is red, the entered kV is out of range and the correction should not be used.
	If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.

Tungsten-Rhodium (W / Rh)
	Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
	Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
	Radcal kV	True kV
	27.27	25.50

Tungsten-Silver (W / Ag)
	Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
	Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
	Radcal kV	True kV
	37.1	38.97

27.586	-8375.0727645926	975.9254356043	-37.913729682	0.4908658347	0	0	0
	-9984.6167916494	1436.5245457141	-82.5051021853	2.3655908176	-0.0338672434	0.0001936869	0

30.1	-540847.695500773	100186.233642731	-7418.4790179813	274.4766092958	-5.0743695436	0.0375005748	0
	-11057.7739361992	1297.2285673767	-56.9891889897	1.1115828564	-0.0081233997	0	0

## dropdowns
Pass/Fail
Pass
Fail

Pass/Fail/NA
Pass
Fail
NA

Phantom Scores
6.0
5.5
5.0
4.5
4.0
3.5
3.0
2.5
2.0
1.5
1.0
0.5
0.0