## Gen_form
1 | [10 empty] | Revision 100.7-20200521 (Has 42 data records) | [10 empty] | PRINT AREAS
2 |  |  |  |  | Medical University of South Carolina | [9 empty] | Medical University of South Carolina | [6 empty] | FIRST,EXP,BLD,TB,WB,FS_KV,REPRO,TIM,LIN,HALFVAL,PT_DEN,PT_TB,PT_WB,KAP_AK,LAST
3 |  |  |  |  | Charleston, South Carolina | [9 empty] | Charleston, South Carolina | [6 empty] | ALL: | FIRST,COMPG2,COMPG3,COMPG4,EXP,BLD,TB,WB,FS_KV,REPRO,TIM,LIN,HALFVAL,PT_DEN,PT_TB,PT_WB,KAP_AK,PT_TB_CR,PT_WB_CR,LAST,DATA1,DATA2
4 | [23 empty] | REM: | (1) Go to last year's spreadsheet (see "**Note", below) and copy from ITS "THIS YEAR" range
5 |  |  |  |  | Radiographic System Compliance Inspection | [8 empty] | Radiographic System Compliance Inspection | [10 empty] | (2) Come back to this sheet and paste (use paste special, "values") into "LAST YEAR" range
6 | [24 empty] | LAST | Cell AD6 | THIS
7 | [23 empty] | Measurement Parameter | YEAR |  | YEAR | (Locked) | RANGES: | LASTYEAR (Data, starting with cell AB8)
8 |  |  |  |  | System Information | [6 empty] | Date: |  |  |  | System Information |  |  | Inspector: | Eugene Mah |  |  |  | Print Area: |  | Change | FIRST,EXP,BLD,TB,WB,FS_KV,REPRO,TIM,LIN,HALFVAL,PT_DEN,PT_TB,PT_WB,KAP_AK,LAST | Y2 |  | THISYEAR (Data, starting with cell AD8)
9 | [11 empty] | Last Year: | [6 empty] | Last Year: |  |  |  |  | Date: |  |  |  | O8
10 |  | Location | [9 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  |  | Inspector: |  | Change | Eugene Mah | V8
11 |  |  | Facility: |  |  |  | Site: | [5 empty] | Facility: |  |  |  | Site Number: | [6 empty] | Facility: |  |  |  | P11
12 |  |  | Department: |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [6 empty] | Department: |  |  |  | P12
13 |  |  | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [6 empty] | Area/Division: |  |  |  | P13
14 |  |  | Survey ID: |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: | [6 empty] | Survey ID: |  |  |  | P14
15 | [12 empty] | Room Number: |  |  |  | Accession Number: | [6 empty] | Room Number: |  |  |  | P15
16
17 |  | X-Ray Generator | [9 empty] | X-Ray Generator | [11 empty] | X-Ray Generator
18 |  |  | Manufacturer: |  |  |  | Date of Manufacture: | [5 empty] | Manufacturer: |  |  |  | Date of Manufacture: | [6 empty] | Manufacturer: |  |  |  | P18
19 |  |  | Model: |  |  |  | Serial Number: | [5 empty] | Model: |  |  |  | Serial Number: | [6 empty] | Model: |  |  |  | P19
20 |  |  | Max. kVp: |  |  |  | Max. mA: | [5 empty] | Max. kVp: |  |  |  | Max. mA: | [6 empty] | Max. kVp: |  |  |  | P20
21 |  |  | Software Version: | [9 empty] | Software Version: | [10 empty] | Software Version:
22
23 |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes |  |  |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube 1
24 |  |  | Tube Designation/Use: |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Large: | [6 empty] | Tube Designation/Use: |  |  |  | P24
25 |  |  | Date of Manufacture: |  |  |  | Small: | [5 empty] | Date of Manufacture: |  |  |  | Small: | [6 empty] | Date of Manufacture: |  |  |  | P25
26 | [16 empty] | Micro:
27 |  |  | Insert Manufacturer: |  |  |  |  | Filtration |  |  |  |  | Insert Manufacturer: |  |  |  |  | Filtration | [5 empty] | Insert Manufacturer: |  |  |  | P27
28 |  |  | Insert Model: |  |  |  | Inherent: | [5 empty] | Insert Model: |  |  |  | Inherent: | [6 empty] | Insert Model: |  |  |  | P28
29 |  |  | Insert Serial Number: |  |  |  | Added: | [5 empty] | Insert Serial Number: |  |  |  | Added: | [6 empty] | Insert Serial Number: |  |  |  | P29
30
31 |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor |  |  |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor | [5 empty] | Housing Manufacturer: |  |  |  | P31
32 |  |  | Housing Model: |  |  |  | Film: | [5 empty] | Housing Model: |  |  |  | Film: | [6 empty] | Housing Model: |  |  |  | P32
33 |  |  | Housing Serial Number: |  |  |  | Screen: | [5 empty] | Housing Serial Number: |  |  |  | Screen: | [6 empty] | Housing Serial Number: |  |  |  | P33
34
35
36 |  | X-Ray Tube 2 | [5 empty] | Focal Spot Sizes |  |  |  | X-Ray Tube 2 | [5 empty] | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube 2
37 |  |  | Tube Designation/Use: |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Large: | [6 empty] | Tube Designation/Use: |  |  |  | P37
38 |  |  | Date of Manufacture: |  |  |  | Small: | [5 empty] | Date of Manufacture: |  |  |  | Small: | [6 empty] | Date of Manufacture: |  |  |  | P38
39 | [16 empty] | Micro:
40 |  |  | Insert Manufacturer: |  |  |  |  | Filtration |  |  |  |  | Insert Manufacturer: |  |  |  |  | Filtration | [5 empty] | Insert Manufacturer: |  |  |  | P40
41 |  |  | Insert Model: |  |  |  | Inherent: | [5 empty] | Insert Model: |  |  |  | Inherent: | [6 empty] | Insert Model: |  |  |  | P41
42 |  |  | Insert Serial Number: |  |  |  | Added: | [5 empty] | Insert Serial Number: |  |  |  | Added: | [6 empty] | Insert Serial Number: |  |  |  | P42
43
44 |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor |  |  |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor | [5 empty] | Housing Manufacturer: |  |  |  | P44
45 |  |  | Housing Model: |  |  |  | Film: | [5 empty] | Housing Model: |  |  |  | Film: | [6 empty] | Housing Model: |  |  |  | P45
46 |  |  | Housing Serial Number: |  |  |  | Screen: | [5 empty] | Housing Serial Number: |  |  |  | Screen: | [6 empty] | Housing Serial Number: |  |  |  | P46
47
48 | [23 empty] | Site Number: |  |  |  | T11
49 |  |  |  |  | Inspection Results | [10 empty] | Inspection Results | [7 empty] | Authorized Use: |  |  |  | T12
50 | [23 empty] | Date of Installation: |  |  |  | T13
51 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [12 empty] | Number of X-ray Tubes: |  |  |  | T14
52 | Rule Number |  |  |  | Labels, Notices, Postings |  |  |  | Compliance |  | 2 | 2 | Is this a new installation? |  |  |  | Text if not tube 1 is: | "(See tube 1 report)" | [5 empty] | Accession Number: |  |  |  | T15
53 | [10 empty] | 2 | 2 | Is this the second or more report on the same generator (i.e. more than one tube?)
54 | DHEC RHB 2.5.1.1 |  | DHEC registration sticker is present, clearly visible and legible. | [5 empty] | TBD | [5 empty] | Labels, Notices, Postings |  |  | 1 | 1 | Tube Number** |  |  |  | Date of Manufacture: |  |  |  | T18
55 | DHEC RHB 10.2.1 |  | DHEC form SC-RHA-20 "Notice to Employees" posted or referenced. | [5 empty] | TBD | [10 empty] | for this room or |  |  |  | Serial Number: |  |  |  | T19
56 | DHEC RHB 4.2.7 |  | Pregnancy Warning sign is posted | [5 empty] | TBD |  |  |  | DHEC registration sticker is present, clearly visible and legible. | [6 empty] | generator |  |  |  | Max. mA: |  |  |  | T20
57 | DHEC RHB 4.3.1 |  | Radiation warning label present on the generator control panel. | [5 empty] | TBD |  |  |  | DHEC form SC-RHA-20 "Notice to Employees" posted or referenced. | [5 empty] | **Note: Input only one of the
58 | DHEC RHB 4.2.6 |  | Technique chart is present and accurate. |  |  |  | DATE: |  | TBD |  |  |  | Pregnancy Warning sign is posted | [5 empty] | following numerals or names: |  |  |  |  | Focal Spot Sizes
59 | JCAHO |  | Patient exposure/dose measurements are posted. |  |  |  | DATE: |  | TBD |  |  |  | Radiation warning label present on the generator control panel. | [5 empty] | 1 | 2 |  |  |  | Large: |  |  |  | T24
60 |  |  | Tube loading charts for each focal spot are posted. | [5 empty] | TBD |  |  | DATE: |  | Technique chart present. |  |  |  |  | AP | 3 |  |  |  | Small: |  |  |  | T25
61 |  |  | Tube loading charts are correct for installed tube(s). | [5 empty] | TBD |  |  | DATE: |  | Patient exposure/dose measurements posted. |  |  |  |  | Front | Lat |  |  |  | Micro: |  |  |  | T26
62 |  |  | Operators manuals are available. | [5 empty] | TBD |  |  |  | Tube loading charts for each focal spot are posted. | [5 empty] | Frontal | Lateral |  |  |  | Filtration
63 |  |  | Monthly radiation monitoring reports are posted. | [5 empty] | TBD |  |  |  | Tube loading charts are correct for installed tube(s). | [5 empty] | Ceph | Pan |  |  |  | Inherent: |  |  |  | T28
64 | [12 empty] | Operators manuals are available. | [5 empty] | Mobile | BCM |  |  |  | Added: |  |  |  | T29
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  |  | Monthly radiation monitoring reports are posted.
66 | Room Number: | [6 empty] | Survey ID: | [15 empty] | Image Receptor
1 | [9 empty] | Medical University of South Carolina | [8 empty] | NEW VALUES |  |  |  |  | Film: |  |  |  | T32
2 |  |  |  |  | Inspection Results |  |  |  |  | Radiographic System Compliance Inspection | [13 empty] | Screen: |  |  |  | T33
3
4 | [23 empty] | Focal Spot Sizes
5 | Rule Number |  |  |  | Beam Limiting Device |  |  |  | Compliance | [14 empty] | Large: |  |  |  | T37
6 | DHEC RHB 4.7.1.1 |  | A beam limiting device is provided to restrict the size of the useful beam. | [5 empty] | TBD | [14 empty] | Small: |  |  |  | T38
7 |  |  | Fixed aperture device is marked with operating SID (i.e. Dental). | [5 empty] | NA | [14 empty] | Micro: |  |  |  | T39
8 | DHEC RHB 4.7.1.4 |  | Means are provided to indicate the projected field size at a specified SID. | [5 empty] | TBD
9 | DHEC RHB 4.3.6 |  | Positive means is provided to ensure the minimum filtration is in beam. | [5 empty] | TBD | [14 empty] | Filtration
10 | DHEC RHB 4.7.12 |  | SID indicator is present and accurate to within 2% SID. | [17 empty] | Medical University of South Carolina |  |  | Inherent: |  |  |  | T41
11 | 21CFR 1020.31(d)(1) |  | Minimum field size is <= 5cm on each side. | [5 empty] | TBD | [11 empty] | Radiographic System Compliance Inspection |  |  | Added: |  |  |  | T42
12 | 21CFR 1020.31(d)(2)(ii) |  | Average Light field illumination >= 160 lux at 100cm. | [5 empty] | TBD |  |  | NEW
13 | DHEC RHB 4.3.5 |  | The measured HVL meets DHEC minimum requirements. | [5 empty] | TBD |  | 10 |  | Enter the distance (cm) from the distal side of the patient to the image receptor for typical patient exposure conditions. | [10 empty] | Image Receptor
14 | [10 empty] | 400 | 400 | Typical Large Focus mA (For technique chart) | [10 empty] | Film: |  |  |  | T45
15 |  |  |  |  | Table Bucky | [5 empty] | 200 | 200 | Typical Small Focus mA |  |  | Exposure/Dose Measurement | [7 empty] | Screen: |  |  |  | T46
16 |  |  | Central ray is perpendicular to image receptor to within 2 degrees | [5 empty] | NA | [14 empty] | New install? |  | Change | 2 | M52
17 | DHEC RHB 4.7.1.7 |  | Center of the x-ray field to center of image receptor <= 2% SID. | [5 empty] | TBD |  | Film/Screen: | ?? | [5 empty] | Date: | [5 empty] | Second Tube? |  | Change | 2 | M53
18 | DHEC RHB 4.7.1.5 |  | The indicated field size is accurate to within 2% of the SID. | [5 empty] | NA |  |  |  |  | Patient |  |  |  |  | Entrance |  |  |  |  | Tube Designation |  | Change | 1 | U54
19 | DHEC RHB 4.7.1.3 |  | X-ray field and light field length/width are congruent within 2% SID. | [5 empty] | NA |  |  |  | SID | thickness | Field size |  |  |  | exposure | KAP
20 | DHEC RHB 4.7.13.1 |  | For PBL, x-ray field length/width and film size are equal to within 3% SID. | [5 empty] | NA |  | Procedure | Grid | (in) | (cm) | (in) | kVp | mA/mAs | time (s) | (mGy) | (Gy-cm^2) | mAs |  |  | Fixed aperture SID marked |  | Change | 3 | M469
21 |  |  | Radiographic image is free of O.D. variations due to grid cutoff. | [5 empty] | TBD |  | CXR | Yes | 72 | 23 | 14x17 | 120 |  |  |  |  | 2
22 |  |  | Radiographic image is free of grid lines. | [5 empty] | TBD |  | Abdomen AP | Yes | 40 | 23 | 14x17 | 76 |  |  |  |  | 30 |  |  | SID Indicator Accuracy |  |  |  | P482
23 | [10 empty] | Ribs | Yes | 40 | 20 | 14x17 | 66 |  |  |  |  | 12 | [6 empty] | P483
24 | [10 empty] | C-Spine AP | Yes | 40 | 13 | 10x12 | 72 |  |  |  |  | 10 | [6 empty] | P484
25 |  |  |  |  | Wall Bucky | [5 empty] | Skull | Yes | 40 | 15 | 8x10 | 72 |  |  |  |  | 10
26 |  |  | Center of the x-ray field to center of image receptor <= 0.5 inch. | [5 empty] | TBD |  | Foot | No | 40 | 8 | 8x10 | 60 |  |  |  |  | 2.5 |  |  | Min. field-Longitudinal |  |  |  | Q494
27 | DHEC RHB 4.7.1.5 |  | The indicated field size is accurate to within 2% of the SID. | [5 empty] | NA |  | Knee | No | 40 | 12 | 10x12 | 66 |  |  |  |  | 8 |  |  | Min. field-Transverse |  |  |  | R494
28 | DHEC RHB 4.7.1.3 |  | X-ray field and light field length/width are congruent within 2% SID. | [5 empty] | NA |  | Pelvis AP | Yes | 40 | 20 | 14x17 | 78 |  |  |  |  | 30
29 | DHEC RHB 4.7.13.1 |  | For PBL, x-ray field length/width and film size are equal to within 1% SID. | [5 empty] | NA |  | Shoulder | Yes | 40 | 15 | 10x12 | 74 |  |  |  |  | 14 |  |  | Table Bucky
30 |  |  | Radiographic image is free of O.D. variations due to grid cutoff. | [5 empty] | TBD |  | LS Spine AP | Yes | 40 | 23 | 10x12 | 78 |  |  |  |  | 25 |  |  | TB used? |  |  |  | M409
31 |  |  | Radiographic image is free of grid lines. | [5 empty] | TBD | [14 empty] | Beam alignment on table top? |  |  |  | M410
32 | [11 empty] | Comments:
33 | [11 empty] | Do NOT exceed 390 |  | 0 | [9 empty] | Phosphor Size
34 |  |  |  |  | Generator | [6 empty] | New | [11 empty] | Transverse |  |  |  | P414
35 | DHEC RHB 4.3.9.1 |  | Means are provided to indicate technique factors prior to exposure. | [5 empty] | TBD | [14 empty] | Longitudinal |  |  |  | Q414
36 |  |  | All technique indicators functional and are free of excessive wear. | [5 empty] | TBD | [14 empty] | Size on Film
37 |  |  | Technique overload circuits function to block excessive technique. | [5 empty] | TBD |  |  | Note: 10 | cm distance between patient and image receptor | [10 empty] | Transverse |  |  |  | P415
38 |  |  | Anode heat calculator/indicator is functioning properly. | [5 empty] | TBD | [14 empty] | Longitudinal |  |  |  | Q415
39 | DHEC RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational. | [5 empty] | TBD
40 | DHEC RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational. | [5 empty] | TBD | [14 empty] | Source to Table Top (cm) |  |  |  | M420
41 | DHEC RHB 4.7.4.2.4 |  | Audible exposure end indicator present and operational. | [5 empty] | TBD | [14 empty] | Table top to cassette (cm) |  |  |  | M421
42 |  |  | Exposure delay time is not excessive. | [5 empty] | TBD | [10 empty] | LAST
43 | DHEC RHB 4.7.7 |  | kVp accurate to within 5% of indicated. | [5 empty] | TBD |  |  |  |  | New Values | [5 empty] | YEAR |  |  |  | Table bucky cm/in |  | Change | cm | V423
44 | DHEC RHB 4.7.5 |  | kVp and exposure reproducibility coeff. of var. < 5%. | [5 empty] | NA |  |  |  |  | Patient |  |  |  | Entrance | ESE | Entrance |  |  |  | Bucky SID |  |  |  | U423
45 | DHEC RHB 4.7.7 |  | Timer accuracy within 10%. | [5 empty] | NA |  |  |  | SID | thickness | Field size |  |  | exposure | within | exposure | Percent
46 | DHEC 4.7.4.2.7 |  | Timer reproducibility coeff. of var. < 5%. | [5 empty] | NA |  | Procedure | Grid | (in) | (cm) | (in) | kVp | mAs | (mGy) | limits | (mGy) | Change |  |  | kVp |  | Change | 80 | O542
47 | DHEC RHB 4.7.8 |  | Exposure coefficient of linearity < 10%. | [5 empty] | TBD |  | CXR | Yes | 72 | 23 | 14x17 | 120 | 2 | [6 empty] | mA |  |  |  | Q542
48 | [10 empty] | Abdomen AP | Yes | 40 | 23 | 14x17 | 76 | 30 | [6 empty] | Time |  |  |  | S542
49 |  |  |  |  | Radiation Safety | [5 empty] | Ribs | Yes | 40 | 20 | 14x17 | 66 | 12 | [6 empty] | New mAs value |  | Change | 0.5 | U542
50 | DHEC RHB 4.7.4.2.1 |  | Continuous pressure on exposure switch required to make an exposure. | [5 empty] | TBD |  | C-Spine AP | Yes | 40 | 13 | 10x12 | 72 | 10
51 |  |  | Spot film exposure blocked when there is no cassette | [5 empty] | NA |  | Skull | Yes | 40 | 15 | 8x10 | 72 | 10 | [6 empty] | Indicated field size
52 |  |  | Spot film exposure blocked when serial exposures are completed | [5 empty] | NA |  | Foot | No | 40 | 8 | 8x10 | 60 | 2.5 | [6 empty] | Transverse 1 |  |  |  | M554
53 | 21CFR 1020.31(g)(4) |  | Spot film device provides acceptable centering to the image receptor | [5 empty] | NA |  | Knee | No | 40 | 12 | 10x12 | 66 | 8 | [6 empty] | Longitudinal 1 |  |  |  | N554
54 | DHEC RHB 4.7.4.2.2 |  | Exposure switch for selected tube is secured at a safe location | [5 empty] | TBD |  | Pelvis AP | Yes | 40 | 20 | 14x17 | 78 | 30 | [6 empty] | Transverse 2 |  |  |  | M555
55 | DHEC RHB 4.7.4.2.5 |  | Operator has full visibility of the exposure factors and patient. | [5 empty] | TBD |  | Shoulder | Yes | 40 | 15 | 10x12 | 74 | 14 | [6 empty] | Longitudinal 2 |  |  |  | N555
56 | DHEC RHB 4.2.14.7 |  | Lead aprons available. |  | Half: 0 |  | Full: 0 |  | TBD |  | LS Spine AP | Yes | 40 | 23 | 10x12 | 78 | 25
57 | DHEC RHB 4.2.14.7 |  | Lead gloves available. |  |  | Number (Left + Right): | 0 |  | TBD | [8 empty] | ESE Limits | Pass |  |  |  |  | PBL Present |  |  |  | M572
58 |  |  | A properly designed and installed apron rack is present. | [5 empty] | TBD |  | ESE Limits - Title B Part IV Appendix D |  |  | Film speed |  | Film speed | [7 empty] | Stationary Grid? |  |  |  | M589
59 | DHEC RHB 4.2.14.7 |  | Documentation of annual protective apparel integrity inspection is available. | [5 empty] | TBD |  |  |  | Thickness | 200/Dig | 400 | 200/Dig | 400
60 |  |  | Patient restraint devices available. | [5 empty] | TBD |  | Projection |  | (cm) | mR |  | mGy | [7 empty] | Wall Bucky
61 | DHEC RHB 4.2.12 |  | Gonadal shielding is present and functional. | [5 empty] | TBD |  | PA Chest - Grid |  | 23 | 38 | 23 | 0.33288 | 0.20148 | [6 empty] | WB used? |  |  |  | M434
62 |  |  | Radiation warning light at room entrance is functional with selected tube | [5 empty] | TBD |  | PA Chest - Non grid |  | 23 | 23 | 8 | 0.20148 | 0.07008
63 | DHEC RHB 4.2.13 |  | Entrance skin exposures are within limits | [5 empty] | YES |  | AP Abdomen |  | 23 | 735 | 450 | 6.4386 | 3.942 | [6 empty] | Phosphor Size
64 | [10 empty] | AP Lumbar spine |  | 23 | 675 | 525 | 5.913 | 4.599 | [6 empty] | Transverse |  |  |  | P438
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  | AP full spine |  | 23 | 390 | 218 | 3.4164 | 1.9097 | [6 empty] | Longitudinal |  |  |  | Q438
66 | Room Number: | [6 empty] | Survey ID: |  |  | AP C-spine |  | 13 | 203 | 142 | 1.7783 | 1.2439 | [6 empty] | Size on Film
1 | [9 empty] | Medical University of South Carolina | Lateral skull |  | 15 | 218 | 105 | 1.9097 | 0.9198 | [6 empty] | Transverse |  |  |  | P439
2 |  |  |  |  | Inspection Results |  |  |  |  | Radiographic System Compliance Inspection | AP Ret Pyelogram |  | 23 | 893 | 893 | 7.8227 | 7.8227 | [6 empty] | Longitudinal |  |  |  | Q439
3 | [10 empty] | AP T-spine |  | 23 | 612 | 612 | 5.3611 | 5.3611
4 | [10 empty] | DP Foot |  | 8 | 111 | 111 | 0.97236 | 0.97236 | [6 empty] | kVp |  | Change | 80 | O601
5 |  |  |  |  | X-ray Tube |  |  |  | Compliance |  | Cephalometric |  | 15 | 45 | 45 | 0.3942 | 0.3942 | [6 empty] | mA |  |  |  | Q601
6 |  |  | Focal spot resolution is within acceptable parameters | [5 empty] | NA | [14 empty] | Time |  |  |  | S601
7 |  |  | The x-ray output @40" within +/-25% from previous measurements. | [5 empty] | NA | [14 empty] | New mAs value |  | Change | 2 | U601
8
9 |  |  |  |  | Phototiming - General | [18 empty] | Wall bucky cm/in |  | Change | cm | V448
10 | 21CFR 1020.31(a)(3)(i) |  | Indication given on control panel when AEC is selected. | [5 empty] | NA | [14 empty] | Wall Bucky SID 1 |  |  |  | M616
11 | 21CFR 1020.31(a)(3)(iii) |  | Backup limit is <= 600 mAs or 60 kWs |  |  |  | Limit: |  | NA | [14 empty] | Wall Bucky SID 2 |  |  |  | M617
12 | 21CFR 1020.31(a)(3)(iv) |  | Visible signal indicates backup limit and subsequent exposures blocked. | [5 empty] | NA
13 |  |  | Density control produces a percentage increment between 15-50%. | [5 empty] | NA | [14 empty] | Indicated field size
14 | [23 empty] | Transverse 1 |  |  |  | N616
15 |  |  |  |  | Phototiming - Table Bucky | [18 empty] | Vertical 1 |  |  |  | O616
16 |  |  | Average Optical Density is within MUSC standards |  |  |  |  | O.D.:_____ | NA | [14 empty] | Transverse 2 |  |  |  | N617
17 |  |  | Response of detectors is the same to within MUSC standards. | [5 empty] | NA | [14 empty] | Vertical 2 |  |  |  | O617
18 |  |  | Reproducibility for each detector provides coeff. of var. <=5%. | [5 empty] | NA
19 |  |  | Performance with varying kVp and thickness is within MUSC standards. | [5 empty] | NA | [14 empty] | PBL Present |  |  |  | M635
20 | [23 empty] | Stationary Grid? |  |  |  | M653
21 |  |  |  |  | Phototiming - Wall Bucky
22 |  |  | Average Optical Density is within MUSC standards |  |  |  |  | O.D.:_____ | NA | [14 empty] | X-RAY TUBE
23 |  |  | Response of detectors is the same to within MUSC standards. | [5 empty] | NA | [13 empty] | Large Focus: | mA/mAs |  | Change | mA | P668
24 |  |  | Reproducibility for each detector provides coeff. of var. <=5%. | [5 empty] | NA | [14 empty] | kVp |  |  |  | O669
25 |  |  | Performance with varying kVp and thickness is within MUSC standards. | [5 empty] | NA | [14 empty] | mA |  |  |  | P669
26 | [23 empty] | Time |  |  |  | Q669
27 |  |  |  |  | Physical Inspection | [18 empty] | SID |  |  |  | R669
28 |  |  | System clean and free of debris. | [5 empty] | TBD | [14 empty] | OFD |  |  |  | S669
29 |  |  | All cables free from excessive wear or restraints. | [5 empty] | TBD
30 |  |  | System free of missing or damaged components. | [5 empty] | TBD | [14 empty] | Last year Lrg FS resolution |  |  |  | U669
31 |  |  | Table movements and locks acceptable | [5 empty] | TBD | [14 empty] | Previous 1 yr Lrg FS resolution |  |  |  | V669
32 |  |  | Table bucky cassette tray movements and locks acceptable | [5 empty] | TBD
33 |  |  | Wall bucky movements and locks acceptable | [5 empty] | TBD | [13 empty] | Small Focus: | mA/mAs |  | Change | mA | P673
34 |  |  | Wall bucky cassette tray movements and locks acceptable | [5 empty] | TBD | [14 empty] | kVp |  |  |  | O674
35 |  |  | Tube crane movements and locks acceptable. | [5 empty] | TBD | [6 empty] | COMMENTS | [7 empty] | mA |  |  |  | P674
36 |  |  | All other movements and locks functional. | [5 empty] | TBD |  |  | Note: Type all comments in the shaded cells. | [11 empty] | Time |  |  |  | Q674
37 | [14 empty] | Previous Year Comments: | [8 empty] | SID |  |  |  | R674
38 | Comments: | [10 empty] | Comments: | [11 empty] | OFD |  |  |  | S674
39 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
40 | [11 empty] | Additional comments: | [11 empty] | Last year Sm FS resolution |  |  | 0 | U674
41 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **A technique chart was not posted in the room. A technique chart reflecting techniques currently used with the CR system must be provided (DHEC RHB 4.2.8). | [6 empty] | Previous 1 yr Sm FS resolution |  |  |  | V674
42 | [11 empty] | Additional comments:
43 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console. | [6 empty] | Test Equipment Utilized: |  | Change | Piranha CB2-17090320 | O687
44 | [11 empty] | Additional comments: | [11 empty] | Piranha CB2-17090320
45 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Tube loading charts for each focal spot were not available at the time of inspection. These charts should be obtained and posted near the operator's console.
46 | [11 empty] | Additional comments: | [11 empty] | LFSDD |  | Change | cm | N693
47 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Operator's manuals for the imaging equipment were not available at the time of inspection. These should be obtained and made available to the operator of the equipment. | [6 empty] | SFSDD |  | Change | cm | N711
48 | [11 empty] | Additional comments:
49 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office. | [6 empty] | Timer
50 | [11 empty] | Additional comments: | [11 empty] | Generator Type |  | Change | HF | O805
51 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **The light field illumination failed to meet the minimum specification of 160 lux. The light source should be adjusted or replaced so as to meet this requirement. | [6 empty] | Timer Test Valid? |  |  |  | M806
52 | [11 empty] | Additional comments:
53 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Phototime Backup |  |  |  | O1027
54 | [11 empty] | Additional comments:
55 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Large Focus Output
56 | [23 empty] | SDD(in.) |  | Change | 39.3701 | AB188
57
58 | [23 empty] | kVp |  |  |  | P1397
59 | [27 empty] | P1398
60 | [14 empty] | DO NOT USE THIS SPACE | [12 empty] | P1399
61 | [27 empty] | P1400
62 | [27 empty] | P1401
63 | [27 empty] | P1402
64 | [27 empty] | P1403
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [18 empty] | P1404
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina |  |  | mGy/mAs |  |  |  | R1397
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [10 empty] | Radiographic System Compliance Inspection |  |  | @ 39.4 in. |  |  |  | R1398
3 | [27 empty] | R1399
4 |  |  |  |  | COMMENTS PAGE 2 | [10 empty] | COMMENTS PAGE 2 | [11 empty] | R1400
5 | [27 empty] | R1401
6 | [14 empty] | Previous Year Comments: | [12 empty] | R1402
7 | Comments: | [10 empty] | Comments: | [15 empty] | R1403
8 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1404
9 | [11 empty] | Additional comments:
10 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Small Focus Output
11 | [11 empty] | Additional comments: | [11 empty] | SDD(in.) |  | Change | 39.3701 | N1411
12 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0
13 | [11 empty] | Additional comments: | [11 empty] | kVp |  |  |  | P1414
14 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1415
15 | [11 empty] | Additional comments: | [15 empty] | P1416
16 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1417
17 | [11 empty] | Additional comments: | [15 empty] | P1418
18 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1419
19 | [11 empty] | Additional comments:
20 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | mGy/mAs |  |  |  | R1414
21 | [11 empty] | Additional comments: | [11 empty] | @ 39.4 in. |  |  |  | R1415
22 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1416
23 | [11 empty] | Additional comments: | [15 empty] | R1417
24 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1418
25 | [11 empty] | Additional comments: | [15 empty] | R1419
26 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | mGy/mAs Output
27 | [11 empty] | Additional comments: | [11 empty] | Lrg Focus Max. |  | Change | NA | Q1406
28 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Lrg Focus Median |  | Change | NA | R1406
29 | [11 empty] | Additional comments: | [11 empty] | Lrg Focus Min |  | Change | NA | S1406
30 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Sml Focus Max. |  | Change | NA | Q1421
31 | [11 empty] | Additional comments: | [11 empty] | Sml Focus Median |  | Change | NA | R1421
32 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Sml Focus Min |  | Change | NA | S1421
33 | [11 empty] | Additional comments:
34 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Exposure Chart
35 | [11 empty] | Additional comments: | [11 empty] | Patient to film distance (cm) |  | Change | 10 | M1456
36 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | LF mA |  | Change | 400 | M1457
37 | [11 empty] | Additional comments: | [11 empty] | SF mA |  | Change | 200 | M1458
38 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0
39 | [11 empty] | Additional comments: | [11 empty] | Procedure |  | Change | CXR | N1464
40 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | Abdomen AP | N1465
41 | [11 empty] | Additional comments: | [13 empty] | Change | Ribs | N1466
42 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | C-Spine AP | N1467
43 | [11 empty] | Additional comments: | [13 empty] | Change | Skull | N1468
44 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | Foot | N1469
45 | [25 empty] | Change | Knee | N1470
46 | [25 empty] | Change | Pelvis AP | N1471
47 | [25 empty] | Change | Shoulder | N1472
48 | [25 empty] | Change | LS Spine AP | N1473
49 | [14 empty] | DO NOT USE THIS SPACE
50 | [23 empty] | Grid |  | Change | Yes | O1464
51 | [25 empty] | Change | Yes | O1465
52 | [25 empty] | Change | Yes | O1466
53 | [25 empty] | Change | Yes | O1467
54 | [25 empty] | Change | Yes | O1468
55 | [25 empty] | Change | No | O1469
56 | [25 empty] | Change | No | O1470
57 | [25 empty] | Change | Yes | O1471
58 | [25 empty] | Change | Yes | O1472
59 | [25 empty] | Change | Yes | O1473
60
61 | [23 empty] | SID (in.) |  | Change | 72 | P1464
62 | [25 empty] | Change | 40 | P1465
63 | [25 empty] | Change | 40 | P1466
64 | [25 empty] | Change | 40 | P1467
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [16 empty] | Change | 40 | P1468
66 | Room Number: | [6 empty] | Survey ID: | [17 empty] | Change | 40 | P1469
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina |  |  |  |  | Change | 40 | P1470
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [10 empty] | Radiographic System Compliance Inspection |  |  |  |  | Change | 40 | P1471
3 | [25 empty] | Change | 40 | P1472
4 |  |  |  |  | COMMENTS PAGE 2 | [10 empty] | COMMENTS PAGE 2 | [9 empty] | Change | 40 | P1473
5
6 | [14 empty] | Previous Year Comments: | [8 empty] | thickness (cm) |  | Change | 23 | Q1464
7 | Comments: | [10 empty] | Comments: | [13 empty] | Change | 23 | Q1465
8 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 20 | Q1466
9 | [11 empty] | Additional comments: | [13 empty] | Change | 13 | Q1467
10 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 15 | Q1468
11 | [11 empty] | Additional comments: | [13 empty] | Change | 8 | Q1469
12 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 12 | Q1470
13 | [11 empty] | Additional comments: | [13 empty] | Change | 20 | Q1471
14 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 15 | Q1472
15 | [11 empty] | Additional comments: | [13 empty] | Change | 23 | Q1473
16 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0
17 | [11 empty] | Additional comments: | [11 empty] | Field Size (in.) |  | Change | 14x17 | R1464
18 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 14x17 | R1465
19 | [11 empty] | Additional comments: | [13 empty] | Change | 14x17 | R1466
20 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 10x12 | R1467
21 | [11 empty] | Additional comments: | [13 empty] | Change | 8x10 | R1468
22 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 8x10 | R1469
23 | [11 empty] | Additional comments: | [13 empty] | Change | 10x12 | R1470
24 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 14x17 | R1471
25 | [11 empty] | Additional comments: | [13 empty] | Change | 10x12 | R1472
26 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 10x12 | R1473
27 | [11 empty] | Additional comments:
28 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | kVp |  | Change | 120 | S1464
29 | [11 empty] | Additional comments: | [13 empty] | Change | 76 | S1465
30 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 66 | S1466
31 | [11 empty] | Additional comments: | [13 empty] | Change | 72 | S1467
32 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 72 | S1468
33 | [11 empty] | Additional comments: | [13 empty] | Change | 60 | S1469
34 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 66 | S1470
35 | [11 empty] | Additional comments: | [13 empty] | Change | 78 | S1471
36 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 74 | S1472
37 | [11 empty] | Additional comments: | [13 empty] | Change | 78 | S1473
38 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0
39 | [11 empty] | Additional comments: | [11 empty] | mAs |  | Change | 2 | W1464
40 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 30 | W1465
41 | [11 empty] | Additional comments: | [13 empty] | Change | 12 | W1466
42 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 10 | W1467
43 | [11 empty] | Additional comments: | [13 empty] | Change | 10 | W1468
44 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [9 empty] | Change | 2.5 | W1469
45 | [25 empty] | Change | 8 | W1470
46 | [25 empty] | Change | 30 | W1471
47 | [25 empty] | Change | 14 | W1472
48 | [25 empty] | Change | 25 | W1473
49 | [14 empty] | DO NOT USE THIS SPACE
50 | [23 empty] | COMMENT PAGE 2 |  |  | 0 | O205
51
52 | [26 empty] | 0 | O207
53
54 | [26 empty] | 0 | O209
55
56 | [26 empty] | 0 | O211
57
58 | [26 empty] | 0 | O213
59
60 | [26 empty] | 0 | O215
61
62 | [26 empty] | 0 | O217
63
64 | [26 empty] | 0 | O219
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID: | [18 empty] | 0 | O221
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [10 empty] | Radiographic System Compliance Inspection | [5 empty] | 0 | O223
3
4 |  |  |  |  | COMMENTS PAGE 3 | [10 empty] | COMMENTS PAGE 3 | [10 empty] | 0 | O225
5
6 | [14 empty] | Previous Year Comments: | [11 empty] | 0 | O227
7 | Comments: | [10 empty] | Comments:
8 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O229
9 | [11 empty] | Additional comments:
10 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O231
11 | [11 empty] | Additional comments:
12 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O233
13 | [11 empty] | Additional comments:
14 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O235
15 | [11 empty] | Additional comments:
16 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O237
17 | [11 empty] | Additional comments:
18 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O239
19 | [11 empty] | Additional comments:
20 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O241
21 | [11 empty] | Additional comments:
22 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | COMMENT PAGE 3 |  |  | 0 | O271
23 | [11 empty] | Additional comments:
24 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O273
25 | [11 empty] | Additional comments:
26 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O275
27 | [11 empty] | Additional comments:
28 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O277
29 | [11 empty] | Additional comments:
30 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O279
31 | [11 empty] | Additional comments:
32 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O281
33 | [11 empty] | Additional comments:
34 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O283
35 | [11 empty] | Additional comments:
36 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O285
37 | [11 empty] | Additional comments:
38 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O287
39 | [11 empty] | Additional comments:
40 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O289
41 | [11 empty] | Additional comments:
42 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O291
43 | [11 empty] | Additional comments:
44 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [10 empty] | 0 | O293
45
46 | [26 empty] | 0 | O295
47
48 | [26 empty] | 0 | O297
49 | [14 empty] | DO NOT USE THIS SPACE
50 | [26 empty] | 0 | O299
51
52 | [26 empty] | 0 | O301
53
54 | [26 empty] | 0 | O303
55
56 | [26 empty] | 0 | O305
57
58 | [26 empty] | 0 | O307
59
60 | [23 empty] | COMMENT PAGE 4 |  |  | 0 | O337
61
62 | [26 empty] | 0 | O339
63
64 | [26 empty] | 0 | O341
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID: | [18 empty] | 0 | O343
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Exposure/Dose Measurement |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | 0 | O345
3
4 | [26 empty] | 0 | O347
5 | Film/Screen: | ?? |  |  |  |  | Date: | [8 empty] | SETUP FOR PERFORMANCE TESTING OF THE CASSETTE BUCKIES
6 |  |  |  | Patient |  |  |  |  | Entrance | [17 empty] | 0 | O349
7 |  |  | SID | thickness | Field size |  |  |  | air kerma | KAP
8 | Procedure | Grid | (in) | (cm) | (in) | kVp | mA | time (s) | (mGy) | (Gy-cm^2) | [16 empty] | 0 | O351
9 | CXR | Yes | 72 | 23 | 14x17 | 120 |  |  |  |  | TABLE BUCKY
10 | Abdomen AP | Yes | 40 | 23 | 14x17 | 76 | [20 empty] | 0 | O353
11 | Ribs | Yes | 40 | 20 | 14x17 | 66
12 | C-Spine AP | Yes | 40 | 13 | 10x12 | 72 |  |  |  |  | Enter 1 for YES, 2 for NO | [15 empty] | 0 | O355
13 | Skull | Yes | 40 | 15 | 8x10 | 72 | [6 empty] | Is this tube used with a table bucky or other horizontal type film image receptor?
14 | Foot | No | 40 | 8 | 8x10 | 60 | [6 empty] | Mag. Correction - Is beam alignment test being done with the film on the table top? | [13 empty] | 0 | O357
15 | Knee | No | 40 | 12 | 10x12 | 66
16 | Pelvis AP | Yes | 40 | 20 | 14x17 | 78 | [20 empty] | 0 | O359
17 | Shoulder | Yes | 40 | 15 | 10x12 | 74 | [5 empty] | Computed Radiog. |  | Transverse | Longitud |  | Transverse | Longitud
18 | LS Spine AP | Yes | 40 | 23 | 10x12 | 78 | [6 empty] | Phosphor Size: |  |  | CR |  |  | New | [7 empty] | 0 | O361
19 | [12 empty] | Size on Film: |  |  | Avg. CF |  |  | New
20 | Comment: | [11 empty] | Correction Factor: | NA | NA | 1 | [10 empty] | 0 | O363
21
22 | [26 empty] | 0 | O365
23 | [10 empty] | Beam Alignment |  |  |  |  | NOTE: | Use consistent units (cm/in.)
24 | Note: 10 | cm distance between patient and image receptor | [10 empty] | (cm) Source to Table |  |  |  | throughout |  | Last Year: | ?? | [6 empty] | 0 | O367
25 | [12 empty] | (cm) Table top to cassette
26 | [18 empty] | Push button to select cm/in. | [7 empty] | 0 | O369
27 | [17 empty] | Bucky SID: |  | cm
28 | [17 empty] | Last Year | [8 empty] | 0 | O371
29 | [17 empty] | New
30 | [26 empty] | 0 | O373
31
32 | [23 empty] | COMMENTS
33 | [23 empty] | Beam Limiting Device |  |  |  | O475
34 | [10 empty] | WALL BUCKY | [12 empty] | SID Indicator |  |  |  | O488
35 | [23 empty] | Min. field size |  |  |  | O498
36 | [23 empty] | Light illumination |  |  |  | O509
37 | [10 empty] | Enter 1 for YES, 2 for NO | [12 empty] | Table Bucky Collimator |  |  |  | O565
38 | [12 empty] | Is this tube used with a wall bucky or other vertical type film image receptor? | [10 empty] | Table Bucky PBL |  |  |  | O582
39 | [23 empty] | Table Bucky Grid |  |  |  | O591
40 | [23 empty] | Wall Bucky Collimator |  |  |  | O628
41 | [11 empty] | Computed Radiog. |  | Transverse | Longitudinal |  | Transverse | Longitudinal | [5 empty] | Wall Bucky PBL |  |  |  | O645
42 | [12 empty] | Phosphor Size: |  |  | Avg. |  |  | New |  |  |  |  | Wall Bucky Grid |  |  |  | O656
43 | [12 empty] | Size on Film: |  |  | CF |  |  | New |  |  |  |  | Focal Spot Resolution |  |  |  | O683
44 | [12 empty] | Correction Factor: | NA | NA | 1 | [7 empty] | Large Focus kV accuracy |  |  |  | O703
45 | [23 empty] | Small Focus kV accuracy |  |  |  | O719
46 | [23 empty] | Large Focus Timer Reprod-1 |  |  |  | O751
47 | [23 empty] | Large Focus Timer Reprod-2 |  |  |  | O753
48 | [15 empty] | NOTE: | Use consistent units (cm/in.) | [6 empty] | Small Focus Timer Reprod-1 |  |  |  | P770
49 | [16 empty] | throughout |  | Last Year | ?? |  |  |  | Small Focus Timer Reprod-2 |  |  |  | P772
50 | [23 empty] | Exposure Timer |  |  |  | O824
51 | [18 empty] | Push button to select cm/in. |  |  |  |  | Large Focus Linearity |  |  |  | O877
52 | [19 empty] | cm |  |  |  | Small Focus Linearity |  |  |  | O892
53 | [23 empty] | Half Value Layer |  |  |  | O946
54 | [23 empty] | Phototiming-table bucky |  | Change | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall | O1105
55 | [23 empty] | Phototiming-wall bucky |  | Change | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall | O1171
56 | [23 empty] | Patient Exposure Data |  |  |  | P1475
57 | [23 empty] | Phototime-Table bucky (CR) |  |  |  | O1301
58 | [23 empty] | Phototime-Wall bucky (CR) |  |  |  | O1367
59
60 | [23 empty] | Tube Designation(column 1): |  | Change | 1 | U59
61 | [25 empty] | Change | AP | U60
62 | [25 empty] | Change | Front | U61
63 |  | LFS Output Slope: |  |  | SFS Output Slope: |  |  | Overall Slope: | [17 empty] | Change | Frontal | U62
64 | [25 empty] | Change | Ceph | U63
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [16 empty] | Change | Mobile | U64
66 | Room Number: | [6 empty] | Survey ID: | [19 empty] | U65
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  | Tube Designation(column 2): |  | Change | 2 | V59
3 | [25 empty] | Change | 3 | V60
4 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [14 empty] | Change | Lat | V61
5 | Rule Number |  |  |  | Beam Limiting Device |  |  |  | Compliance | [6 empty] | Beam Limiting Device | [9 empty] | Change | Lateral | V62
6 | DHEC RHB 4.7.1 |  | A beam limiting device is provided to restrict the size of the useful beam. | [5 empty] | TBD |  |  |  | A beam limiting device is provided to restrict the size of the useful beam. | [12 empty] | Change | Pan | V63
7 |  |  | Fixed aperture device is marked with operating SID (i.e. Dental). | [5 empty] | NA |  | 3 | 3 | Fixed aperture device is marked with operating SID (i.e. Dental). | [12 empty] | Change | BCM | V64
8 | DHEC RHB 4.7.1.5 |  | Means are provided to indicate the projected field size at a specified SID. | [5 empty] | TBD |  |  |  | Means are provided to indicate the projected field size at a specified SID. | [14 empty] | V65
9 | DHEC RHB 4.3.6 |  | Positive means is provided to ensure the minimum filtration is in beam. | [5 empty] | TBD |  |  |  | Positive means is provided to ensure the minimum filtration is in beam.
10 | DHEC RHB 4.7.12 |  | SID Indicator is present on the system. | [5 empty] | TBD |  |  |  | SID Indicator is present on the system. | [10 empty] | SID Indicator Accuracy: |  |  |  | S482
11 |  |  | A light field is provided to indicate location and size of radiation field | [5 empty] | TBD |  |  |  | A light field is provided to indicate location and size of radiation field | [10 empty] | Average Illumination: |  | Change | TBD | U504
12 | DHEC RHB 4.8.12 |  | Minimum SSD is >= 30 cm for portable units | [5 empty] | TBD |  |  |  | Minimum SSD is >= 30 cm for portable units | [10 empty] | Illumination units:
13 | [23 empty] | Alignment Error (Degrees): |  | Change | NA | R537
14 | Comments: | [10 empty] | Comments: | [11 empty] | TB Rad/Film Center: |  |  |  | Q547
15 | [11 empty] | Do NOT exceed 390 |  | 0
16 | [11 empty] | New | [11 empty] | WB Rad/Film Center: |  |  |  | Q608
17
18 | SID Indicator |  | SID | SID | Percent | [8 empty] | New SID | Prev SID | SID | Percent | [6 empty] | HVL Measured kVp: |  | Change | HVL @80 kVp (Max kV) | S950
19 | Accuracy |  | Indicated | Measured | Error | Acceptable: | [5 empty] | SID Indicator |  | Indicated | Indicated | Measured | Error | Acceptable: | [5 empty] | HVL Measured: |  | Change | TBD | S951
20 | [11 empty] | Accuracy
21 | [23 empty] | Table Bucky PT OD: |  | Change | NA | W1063
22 | [23 empty] | Wall Bucky PT OD: |  | Change | NA | W1129
23 | [16 empty] | Last Year:
24 | Criteria: | SID indicator accurate to within 2% | [9 empty] | Criteria: | SID indicator accurate to within 2% | [10 empty] | No. Half Aprons: |  |  | 0 | R1430
25 | [23 empty] | No. Full Aprons: |  |  | 0 | T1430
26 | Comments: | [10 empty] | Comments: | [11 empty] | No. Lead Gloves: |  |  | 0 | R1431
27 | [11 empty] | Do NOT exceed 390 |  | 0 | [9 empty] | Patient Entrance Exposure:
28 | [11 empty] | New | [11 empty] | CXR |  |  |  | V1464
29 | [23 empty] | Abdomen AP |  |  |  | V1465
30 | Minimum |  |  |  | Measurements (cm) | [6 empty] | Minimum |  |  | Measurements (cm) | [8 empty] | Ribs |  |  |  | V1466
31 | Field Size |  |  |  | Longitudinal | Transverse | [5 empty] | Field Size |  |  | Longitudinal | Transverse | [7 empty] | C-Spine AP |  |  |  | V1467
32 | [7 empty] | Acceptable: | TBD | [9 empty] | Acceptable: | TBD |  |  |  | Skull |  |  |  | V1468
33 | [13 empty] | Last Year: | [9 empty] | Foot |  |  |  | V1469
34 | Criteria: | Each dimension of the minimum field size at an SID of 100 cm shall be <= 5 cm | [9 empty] | Criteria: | Each dimension of the minimum field size at an SID of 100 cm shall be <= 5 cm | [10 empty] | Knee |  |  |  | V1470
35 | [23 empty] | Pelvis AP |  |  |  | V1471
36 | Comments: | [10 empty] | Comments: | [11 empty] | Shoulder |  |  |  | V1472
37 | [11 empty] | Do NOT exceed 390 |  | 0 | [9 empty] | LS Spine AP |  |  |  | V1473
38 | [11 empty] | New
39 | [23 empty] | COMMENT PAGE 1 |  |  | 0 | O170
40 | Light Field |  |  | Measurements (lux) at 100cm* |  |  |  | Average Illumination |  |  |  | Light Field |  |  | Measurements (lux) at 100cm* |  |  |  | Average Illumination
41 | Illumination |  | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux) |  |  |  | Illumination |  | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux) | [7 empty] | 0 | O172
42 | [7 empty] | TBD | [10 empty] | TBD
43 |  |  | *Or at maximum SID, whichever is smaller | [10 empty] | *Or at maximum SID, whichever is smaller | [12 empty] | 0 | O174
44 | Criteria: | Average light field illumination >= 160 lux. | [5 empty] | Acceptable: | TBD |  |  | Units: |  | Enter 1 for ft-cd or 2 for lux |  |  |  |  | Acceptable: | TBD
45 |  | Test Equipment Used: | Piranha CB2-17090320 |  |  | Calibration Date: |  |  | Calibration Due: | [17 empty] | 0 | O176
46 | [11 empty] | Criteria: | Average light field illumination >= 160 lux.
47 | Comments: | [16 empty] | Last Year: | [8 empty] | 0 | O178
48 | [11 empty] | Comments:
49 | [11 empty] | Do NOT exceed 390 |  | 0 | [12 empty] | 0 | O180
50 | [11 empty] | New
51 | [26 empty] | 0 | O182
52
53 | [26 empty] | 0 | O184
54
55 | [26 empty] | 0 | O186
56
57 | [23 empty] | PHOTOTIMING DENS TECH
58 | [23 empty] | SID: |  | Change | 101.6 cm | M1003
59 | [23 empty] | Detector: |  | Change | C | N1003
60 | [23 empty] | kVp: |  | Change | 80 | M1008
61 | [23 empty] | Atten. Thickness: |  | Change | 7.874 | N1008
62
63 | [23 empty] | PHOTOTIMING TB TECH
64 | [23 empty] | SID: |  | Change | 101.6 cm | Q1063
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [14 empty] | Detector: |  | Change | C | Q1064
66 | Room Number: | [6 empty] | Survey ID: | [15 empty] | kVp: |  | Change | 80 | Q1065
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina |  |  | Atten. Thickness: |  | Change | 7.874 | Q1066
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection
3 | [23 empty] | PHOTOTIMING WB TECH
4 |  |  |  |  | Table Bucky |  |  |  | CR Field Size Correction Factor: | NA |  |  |  | **TEST** |  | TABLE BUCKY | [7 empty] | SID: |  | Change | 101.6 cm | R1127
5 | [6 empty] | Distance (cm) from the distal side of the |  |  |  | X-ray Centering and Congruence | [5 empty] | IMPORTANT!! Select below if radiation field is | [6 empty] | Detector: |  | Change | C | R1128
6 | [6 empty] | patient to the image receptor for typical | [9 empty] | centered to upper or lower bead of the test object. | [6 empty] | kVp: |  | Change | 80 | R1129
9 | Beam Alignment |  |  | Measured Separation |  | Error |  |  | patient exposure conditions: | NA |  |  | Measured Separation |  |  | Error |  | Lower |  |  | Upper |  |  | Atten. Thickness: |  | Change | 7.874 | R1130
8 |  |  |  | Of Steel Balls (mm) |  | (Degrees) | [6 empty] | Of Steel Balls (mm) |  |  | (Degrees) |  |  |  |  | Lower
9 | [5 empty] | NA |  | Acceptable: | NA | [6 empty] | NA |  | Acceptable: | NA |  |  |  |  | Contact/access Info.:
10 | [14 empty] | Last Year: | [12 empty] | T1
11 | Criteria: | Central ray perpendicular to image receptor to within 2 degrees. | [9 empty] | Criteria: | Central ray perpendicular to image receptor to within 2 degrees. | [14 empty] | T2
12 | [27 empty] | T3
13 | X-ray Centering and Congruence | [9 empty] | Technique: | [6 empty] | Last Year: | [9 empty] | T4
14 | [11 empty] | kVp: | 80 | mA: |  | time: |  | mAs: | 0.5 | [8 empty] | T5
15 | kVp: | 80 | mA: |  | time: |  | mAs: | 0.5 | SID: |  |  | New: | 80 | New: |  | New: |  | New: | [9 empty] | T6
16
17 |  | Rad/Film Center | [9 empty] | Visual Check |  |  | Rad/Film Center
18 |  | Distance (cm) | [5 empty] | Rad/Film Center |  |  |  | Centering okay? |  |  | Distance (cm) |  |  | Rad/Film Center
19 |  | 0 | [5 empty] | Acceptable: | TBD |  |  | Field size indicators okay? | [5 empty] | Acceptable: | TBD
20 | [11 empty] | Edge Congruence Okay? |  |  | Last Year:
21 | Criteria: | Center of x-ray field to center of image receptor <= 0.5 inches. | [9 empty] | Criteria: | Center of x-ray field to center of image receptor <= 2% SID.
22
23 | Indicated Field Size @SID |  | Measured* Radiation |  | Measured** Light |  | Field Size Indicators |  | Light/Rad. Congruence |  | Indicated Field Size @SID |  | Measured* Radiation |  | Measured** Light |  |  | Field Size Indicators |  | Light/Rad. Congruence
24 | (cm) |  | (cm) |  | (cm) |  | Error (% SID) |  | Error (% SID) |  | (cm) |  | (cm) |  | (cm) |  |  | Error (% SID) |  | Error (% SID)
25 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | CR | Transverse | Longitudinal | Transverse | Longitudinal
26 | [6 empty] | NA | NA | NA | NA | [6 empty] | Avg. CF | NA | NA | NA | NA
27 | [6 empty] | NA | NA | NA | NA | [6 empty] | 1 | NA | NA | NA | NA
28 | *Dimension of highest optical density region | [9 empty] | *Dimension of highest optical density region
29 | **Dimension between the image of the coins marking the edge of the light field | [6 empty] | Acceptable |  |  | **Dimension between the image of the coins marking the edge of the light field | [7 empty] | Acceptable
30 |  |  |  |  | Indicated Field Size (cm) |  | Field Size Indicators |  | Light/Rad. Congruence |  | New Values |  |  |  |  | Indicated Field Size (cm) |  | Field Size Indicators |  | Light/Rad. Congruence
31 |  |  |  |  | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal |  |  |  | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal
32 | [6 empty] | NA | NA | NA | NA | [7 empty] | NA | NA | NA | NA
33 | [6 empty] | NA | NA | NA | NA | [7 empty] | NA | NA | NA | NA
34
35 | Criteria: | The indicated field length/width are accurate to within 2% of the SID. | [9 empty] | Criteria: | The indicated field length/width are accurate to within 2% of the SID.
36 |  | X-ray field and light field length/width are congruent within 2% SID. | [10 empty] | X-ray field and light field length/width are congruent within 2% SID.
37 | Comments: | [10 empty] | Comments:
38 | [11 empty] | Do NOT exceed 390 |  | 0
39 | [11 empty] | New
40
41 | Automatic (PBL) Collimation
42 | [8 empty] | Compliance |  | Automatic (PBL) Collimation
43 |  |  |  | Positive beam limitation (PBL) is present and operational. |  |  |  |  | TBD |  | Enter 1 for YES, 2 for PBL present but not functioning, 3 for No PBL or Intentionally disabled
44 | [12 empty] | Positive beam limitation (PBL) is present and operational.
45 | Cassette/Film Size (cm) |  | Measured* Radiation (cm) |  | Error (% SID) |  |  | Acceptable |  |  |  | Cassette/Film Size (cm) |  | Measured* Radiation (cm) |  | Error (% SID) |  |  | Acceptable
46 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Sum | Transverse | Longitudinal | Sum |  | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Sum | Transverse | Longitudinal | Sum
47 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
48 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
49 |  | *Dimension of highest optical density region | [9 empty] | *Dimension of highest optical density region
50 | [10 empty] | New
51 | Criteria: | For PBL, x-ray field length/width and film size are equal to within 2% SID. | [8 empty] | Values
52 |  | Sum of length/width differences is less than 3% SID | [9 empty] | Criteria: | For PBL, x-ray field length/width and film size are equal to within 2% SID.
53 | Comments: | [11 empty] | Sum of length/width differences is less than 3% SID
54 | [11 empty] | Comments:
55 | [11 empty] | Do NOT exceed 390 |  | 0
56 | [11 empty] | New
57 | Rule Number |  |  | Antiscatter Grid |  |  |  |  | Compliance
58 |  |  | Radiographic image is free of O.D. variations due to grid cutoff. | [5 empty] | TBD
59 |  |  | Grid type | [5 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA |  |  |  | Antiscatter Grid
60 |  |  | Radiographic image is free of grid lines. | [5 empty] | TBD |  |  |  | Radiographic image is free of O.D. variations due to grid cutoff.
61 | Comments: | [11 empty] | Is this a stationary grid?
62 | [12 empty] | Radiographic image is free of grid lines.
63 | [11 empty] | Comments:
64 | [11 empty] | Do NOT exceed 390 |  | 0
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | New
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection
3
4 | [13 empty] | **TEST** |  | WALL BUCKY
5 |  |  |  |  | Wall Bucky
6 | [10 empty] | Technique: | [6 empty] | Last Year
7 | [11 empty] | kVp: | 80 | mA: |  | time: |  | mAs: | 2
8 | X-ray Centering and Congruence | [10 empty] | New: | 80 | New: |  | New: |  | New:
9
10 | kVp: | 80 | mA: |  | time: |  | mAs: | 2
11 | [10 empty] | X-ray Centering and Congruence
12 |  | Rad/Film Center | [9 empty] | Visual Check |  |  | Rad/Film Center
13 |  | Distance (cm) | [5 empty] | Rad/Film Center |  |  |  | Centering okay? |  |  | Distance (cm) |  |  | Rad/Film Center
14 |  | 0 | [5 empty] | Acceptable: | TBD |  |  | Field size indicators okay? | [5 empty] | Acceptable: | TBD
15 | [11 empty] | Edge Congruence Okay? |  |  | Last Year:
16 | Criteria: | Center of x-ray field to center of image receptor <= 0.5 inches. | [9 empty] | Criteria: | Center of x-ray field to center of image receptor <= 2% SID.
17
18 |  |  | Indicated Field Size @SID |  | Measured* Radiation |  | Measured** Light
19 |  | SID | (cm) |  | (cm) |  | (cm) |  |  |  | Bucky | Indicated Field Size @SID |  | Measured* Radiation |  | Measured** Light |  | Field Size Indicators |  | Light/Rad. Congruence
20 |  | (cm) | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal |  |  | SID | (cm) |  | (cm) |  | (cm) |  | Error (% SID) |  | Error (% SID)
21 | [6 empty] | 14 | 18 |  |  | (cm) | Transverse | Vertical | Transverse | Vertical | Transverse | Vertical | Transverse | Vertical | Transverse | Vertical
22 | [15 empty] | 14 | 18 | NA | NA | NA | NA
23 |  | *Dimension of highest optical density region | [15 empty] | NA | NA | NA | NA
24 |  | **Dimension between the image of the coins marking the edge of the light field | [8 empty] | *Dimension of highest optical density region
25 | Field Size Indicators |  | Light/Rad. Congruence |  |  |  |  | Acceptable |  |  | **Dimension between the image of the coins marking the edge of the light field | [7 empty] | Acceptable
26 | Error (% SID) |  | Error (% SID) |  | Indicated Field Size (cm) |  | Field Size Indicators |  | Light/Rad. Congruence |  | New Values |  |  |  |  | Indicated Field Size (cm) |  | Field Size Indicators |  | Light/Rad. Congruence
27 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | SID | Transverse | Vertical |  |  | Transverse | Vertical | Transverse | Vertical | Transverse | Vertical
28 | NA | NA | NA | NA |  |  | NA | NA | NA | NA | [7 empty] | NA | NA | NA | NA
29 | NA | NA | NA | NA |  |  | NA | NA | NA | NA | [7 empty] | NA | NA | NA | NA
30 | Criteria: | The indicated field length/width are accurate to within 2% of the SID.
31 |  | X-ray field and light field length/width are congruent within 2% SID. | [9 empty] | Criteria: | The indicated field length/width are accurate to within 2% of the SID.
32 | [12 empty] | X-ray field and light field length/width are congruent within 2% SID.
33 | Comments:
34 | [11 empty] | Comments:
35 | [11 empty] | Do NOT exceed 390 |  | 0
36 | [11 empty] | New
37 | Automatic (PBL) Collimation
38 | [8 empty] | Compliance
39 |  |  |  | Positive beam limitation (PBL) is present and operational. |  |  |  |  | TBD |  | Automatic (PBL) Collimation
40 | [10 empty] | Enter 1 for YES, 2 for PBL present but not functioning, 3 for No PBL or Intentionally disabled | [14 empty] | 3 | Enter 1 - Manual, 2 - Radcal, 3 - Barracuda |  |  |  | Manual |  |  |  | Radcal System 2000 |  |  | Barracuda MPD
41 | SID | Cassette/Film Size (cm) |  | Measured* Radiation (cm) |  | Error (% SID) |  |  | Acceptable |  |  |  | Positive beam limitation (PBL) is present and operational. | [19 empty] | Exp |  |  | Exp
42 | (cm) | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Sum | Transverse | Longitudinal | [15 empty] | kVp Set | mA Set | Time | mAs Set | Add Filt | kVp Eff | kVp Max | Time | Exposure (mGy) | Exposure (mGy) | Time | kVp Max | kVp | ms | Dose (mGy) | Dose rate (mGy/s) | Total filtration (mm Al) | HVL (mm Al)
43 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | SID | Cassette/Film Size (cm) |  | Measured* Radiation (cm) |  | Error (% SID) |  |  | Acceptable | [6 empty] | 60 | 400 | 0.05 | 0 | 0
44 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | (cm) | Transverse | Vertical | Transverse | Vertical | Transverse | Vertical | Sum | Transverse | Vertical | Sum |  |  |  |  | 60 | 400 | 0.05 | 0 | 0
45 |  | *Dimension of highest optical density region | [8 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  |  | 80 | 400 | 0.05 | 0 | 0
46 | [10 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  |  | 80 | 400 | 0.05 | 0 | 0
47 | Criteria: | For PBL, x-ray field length/width and film size are equal to within 2% SID. | [9 empty] | *Dimension of highest optical density region | [13 empty] | 80 | 400 | 0.05 | 0 | 0
48 |  | Sum of length/width differences is less than 3% SID | [13 empty] | New | [9 empty] | 80 | 400 | 0.05 | 0 | 0
49 | [15 empty] | Values | [9 empty] | 80 | 400 | 0.05 | 0 | 0
50 | Comments: | [10 empty] | Criteria: | For PBL, x-ray field length/width and film size are equal to within 2% SID. | [12 empty] | 80 | 400 | 0.05 | 0 | 3
51 | [12 empty] | Sum of length/width differences is less than 3% SID | [12 empty] | 80 | 400 | 0.05 | 0 | 3
52 | [11 empty] | Comments: | [13 empty] | 80 | 400 | 0.05 | 0 | 3.5
53 | [11 empty] | Do NOT exceed 390 |  | 0 | [11 empty] | 80 | 400 | 0.05 | 0 | 3.5
54 | Rule Number |  |  | Antiscatter Grid |  |  |  |  | Compliance |  |  | New | [13 empty] | 80 | 400 | 0.05 | 0 | 2.5
55 |  |  | Radiographic image is free of O.D. variations due to grid cutoff. | [5 empty] | TBD | [16 empty] | 80 | 400 | 0.05 | 0 | 2.5
56 |  |  | Grid type | [5 empty] | TBD | [6 empty] | Antiscatter Grid | [9 empty] | 80 | 500 | 0.05 | 0 | 0
57 |  |  | Radiographic image is free of grid lines. | [5 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA | [14 empty] | 80 | 50 | 0.05 | 0 | 0
58 | [12 empty] | Radiographic image is free of O.D. variations due to grid cutoff. | [12 empty] | 80 | 800 | 0.05 | 0 | 0
59 | Comments: | [11 empty] | Is this a stationary grid? | [12 empty] | 80 | 250 | 0.05 | 0 | 0
60 | [12 empty] | Radiographic image is free of grid lines. | [12 empty] | 100 | 400 | 0.05 | 0 | 0
61 | [25 empty] | 100 | 400 | 0.05 | 0 | 0
62 | [11 empty] | Comments: | [13 empty] | 120 | 400 | 0.05 | 0 | 0
63 | [11 empty] | Do NOT exceed 390 |  | 0 | [8 empty] | Huttner Grating |  |  | 120 | 400 | 0.05 | 0 | 0
64 | [11 empty] | New | [10 empty] | 1.1 | 0.5 |  | 140 | 400 | 0.05 | 0 | 0
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [13 empty] | 1.2 | 0.56 |  | 50 | 100 | 0.1 | 0 | 0
66 | Room Number: | [6 empty] | Survey ID: | [14 empty] | 1.3 | 0.63 |  | 70 | 100 | 0.1 | 0 | 0
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina |  | 2.1 | 0.71 |  | 90 | 100 | 0.1 | 0 | 0
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  |  |  | Measurement Data | [5 empty] | Radiographic System Compliance Inspection |  | 2.2 | 0.8 |  | 90 | 100 | 0.1 | 0 | 0
3 | [22 empty] | 2.3 | 0.9 |  | 90 | 100 | 0.1 | 0 | 0
4 |  |  |  |  | X-RAY TUBE | [17 empty] | 3.1 | 1 |  | 90 | 100 | 0.1 | 0 | 0
5 | [14 empty] | X-ray Tube |  |  |  |  | 1 Year | 2 Year |  | 3.2 | 1.12 |  | 90 | 100 | 0.1 | 0 | 0
6 | [10 empty] | Focal Spot Resolution | [7 empty] | Measured | Previous | Previous |  | 3.3 | 1.25 |  | 80 | 100 | 0.1 | 0 | 0
7 | Focal Spot Resolution | [9 empty] | Large: |  | mm | mA |  |  |  |  | Resolution | Resolution | Resolution |  | 4.1 | 1.4 |  | 80 | 110 | 0.1 | 0 | 0
8 | [8 empty] | 1 Year | 2 Year |  |  | kVp | mA | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  | 4.2 | 1.6 |  | 80 | 250 | 0.1 | 0 | 0
9 | Large: |  | mm |  |  |  |  | Measured | Previous | Previous | [12 empty] | 4.3 | 1.8 |  | 80 | 50 | 0.1 | 0 | 0
10 | [7 empty] | Resolution | Resolution | Resolution |  | New Values: | [10 empty] | 5.1 | 2 |  | 80 | 160 | 0.1 | 0 | 0
11 |  | kVp | mA | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) | [8 empty] | Measured | Previous | Previous |  | 5.2 | 2.24 |  | 110 | 100 | 0.1 | 0 | 0
12 | [10 empty] | Small: |  | mm | mA |  |  |  |  | Resolution | Resolution | Resolution |  | 5.3 | 2.5 |  | 130 | 100 | 0.1 | 0 | 0
13 | [12 empty] | kVp | mA | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  | 6.1 | 2.8 |  | 80 | 200 | 0.01 | 0 | 0
14 | Small: |  | mm |  |  |  |  | Measured | Previous | Previous | [12 empty] | 6.2 | 3.15 |  | 80 | 200 | 0.02 | 0 | 0
15 | [7 empty] | Resolution | Resolution | Resolution |  | New Values: | [10 empty] | 6.3 | 3.55 |  | 80 | 200 | 0.04 | 0 | 0
16 |  | kVp | mA | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  |  | * SID = Source to Image (film) Distance. | [9 empty] | 7.1 | 4 |  | 80 | 200 | 0.1 | 0 | 0
17 | [12 empty] | ** OFD = Object (resolution pattern) to Film Distance. | [9 empty] | 7.2 | 4.5 |  | 80 | 200 | 0.25 | 0 | 0
18 |  | * SID = Source to Image (film) Distance. | [8 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [11 empty] | 7.3 | 5 |  | 80 | 200 | 0.4 | 0 | 0
19 |  | ** OFD = Object (resolution pattern) to Film Distance. | [5 empty] | Acceptable: | NA |  | 3 | Focal spot resolution is within acceptable parameters
20 | [12 empty] | Criteria: | No significant change in resolution since previous measurement.
21 |  | Criteria: | No significant change in resolution since previous measurement.
22 | [31 empty] | Medical University of South Carolina
23 | Comments: | [10 empty] | Comments: | [19 empty] | Radiographic System Compliance Inspection
24 | [11 empty] | Do NOT exceed 390 |  | 0 | [12 empty] | Measurement Data
25 | [11 empty] | New | [9 empty] | Note: See column BD for a description of how the measurement data is utilized.
26 | [15 empty] | Generator Performance | [33 empty] | Exp |  | mAs | Measurement |  | =HVL data
27 | [11 empty] | Test Equipment Used: | Piranha CB2-17090320 |  |  | Calibration Date: | [6 empty] | Mode | kVp Set | mA Set | Time | mAs Set | Filter Set | Add Filt | Material | Distance | Units | Exp. CF. | Sensitivity | Chamber Type | Delay | Wave | Focal Spot | Technique Comments | Linearity | Accuracy | Beam Quality | Parameters | Reproduci... | Survey | Wave File | kVp Avg | kVp Eff | kVp Max | Time (s) | Exposure (mGy) |  | Utilization |  | =Small focus data
28 | [11 empty] | New: | Piranha CB2-17090320 |  |  | Calibration Due: | [6 empty] | Radio | 60 | 400 | 0.05 | 0 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV) | Accu(kV)=Accuracy of kVp
29 |  |  |  |  | GENERATOR PERFORMANCE | [5 empty] | Large Focus kV Accuracy and Output | [11 empty] | Radio | 60 | 400 | 0.05 | 0 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | Accu (kV)
30 | [22 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 1 | 0 | 0 | 1 | [8 empty] | Repr, Accu (kV), Lin
31 | [14 empty] | Note: See focal spot data above for mA/mAs selector button | [7 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
32 |  | Test Equipment Used: | Piranha CB2-17090320 |  |  | Calibration Date: |  |  | Calibration Due: |  | Test | SDD* | Focal | Filter Pack/ |  |  | Set |  | Measured values |  |  |  | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
33 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 1 | [8 empty] | Repr, HVL
34 | kVp Accuracy | [10 empty] | 100 | Large | 50-85/Low | 60 | 400 | 0.05 | [5 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 1 | [8 empty] | Repr, HVL
35 | [11 empty] | 100 | Large | 70-120/Low | 80 | 400 | 0.05 | [5 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL
36 | [11 empty] | 100 | Large | 70-120/Low | 100 | 400 | 0.05 | [5 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | Accu(T)=Accuracy of timer
37 |  | Large Focal Spot |  |  |  |  | Small Focal Spot |  |  |  |  | 100 | Large | 100-155/Low | 120 | 400 | 0.05 | [5 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | Repr=Reproducibility (kVp, Expos, timer)
38 | [11 empty] | 100 | Large | 100-155/Low | 140 | 400 | 0.05 | [5 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | HVL=Half Value Layer
39 | SDD: | 100 | cm |  |  | SDD: | 100 | cm | [14 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL | Lin=Linearity
40 | mA: | 400 |  |  |  | mA: | 100 | [15 empty] | Radio | 80 | 400 | 0.05 | 0 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 0 | 1 | 0 | 0 | [8 empty] | HVL
41 | mAs: |  |  |  |  | mAs: | [16 empty] | Radio | 80 | 500 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
42 | time (s): | 0.05 |  |  |  | time (s): | 0.1 |  |  |  | *SDD=Source to detector distance | [11 empty] | Radio | 80 | 50 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
43 | *SDD=Source to detector distance |  |  |  |  | *SDD=Source to detector distance | [5 empty] | Comments: | [10 empty] | Radio | 80 | 800 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
44 | [11 empty] | Do NOT exceed 190 |  | 0 |  |  |  |  | Acceptable: | TBD |  |  | Radio | 80 | 250 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
45 | Indicated | Measured |  | kV (max) |  | Indicated | Measured |  | kV (max) |  |  | New | [10 empty] | Radio | 100 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
46 | kVp | kV (max) | kV (avg) | % error |  | kVp | kV (max) | kV (avg) | % error |  | Small Focus kV Accuracy and Output | [11 empty] | Radio | 100 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 1 | 0 | 0 | [8 empty] | Accu (kV)
47 | 60 |  |  |  |  | 50 | [16 empty] | Radio | 120 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
48 | 80 |  |  |  |  | 70 | [16 empty] | Radio | 120 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 1 | 0 | 0 | [8 empty] | Accu (kV)
49 | 100 |  |  |  |  | 90 | [8 empty] | Note: See focal spot data above for mA/mAs selector button | [7 empty] | Radio | 140 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
50 | 120 |  |  |  |  | 110 |  |  |  |  | Test | SDD* | Focal | Filter Pack/ |  |  | Set |  | Measured values |  |  |  | Radio | 50 | 100 | 0.1 | 0 | 35-60 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
51 | 140 |  |  |  |  | 130 |  |  |  |  | Setup | cm | Spot | Sensitivity | kV | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  | Radio | 70 | 100 | 0.1 | 0 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
52 | [11 empty] | 100 | Small | 35-60/Low | 50 | 100 | 0.1 | [5 empty] | Radio | 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 1 | [8 empty] | Repr, Accu (kV)
53 | [11 empty] | 100 | Small | 50-85/Low | 70 | 100 | 0.1 | [5 empty] | Radio | 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
54 | [11 empty] | 100 | Small | 70-120/Low | 90 | 100 | 0.1 | [5 empty] | Radio | 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
55 | [11 empty] | 100 | Small | 70-120/Low | 110 | 100 | 0.1 | [5 empty] | Radio | 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
56 |  |  | Acceptable: | TBD |  |  |  | Acceptable: | TBD |  |  | 100 | Small | 100-155/Low | 130 | 100 | 0.1 | [5 empty] | Radio | 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 0 | 0 | 0 | 1 | [8 empty] | Repr
57 | [22 empty] | Radio | 80 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
58 |  | Criteria: | kVp accurate to within 5% of indicated. | [7 empty] | *SDD=Source to detector distance | [11 empty] | Radio | 80 | 110 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
59 | [11 empty] | Comments: | [10 empty] | Radio | 80 | 250 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
60 | [11 empty] | New |  | 0 | [8 empty] | Radio | 80 | 50 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
61 | Comments: | [12 empty] | (Combined comments) |  |  |  |  | Acceptable: | TBD |  |  | Radio | 80 | 160 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 1 | 0 | 0 | 0 | 0 | [8 empty] | Lin
62 | [11 empty] | Do NOT exceed 190 |  | 0 | [8 empty] | Radio | 110 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
63 | [11 empty] | Criteria: | kVp accurate to within 5% of indicated. | [9 empty] | Radio | 130 | 100 | 0.1 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (kV)
64 | [22 empty] | Radio | 80 | 200 | 0.01 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [13 empty] | Radio | 80 | 200 | 0.02 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
66 | Room Number: | [6 empty] | Survey ID: | [14 empty] | Radio | 80 | 200 | 0.04 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina |  | Radio | 80 | 200 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [10 empty] | Radiographic System Compliance Inspection |  | Radio | 80 | 200 | 0.25 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
3 | [15 empty] | Measurement Data | [6 empty] | Radio | 80 | 200 | 0.4 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large |  | 0 | 1 | 0 | 0 | 0 | [8 empty] | Accu (T)
4
5 |  |  |  |  | Generator Performance
6 | [15 empty] | Generator Performance
7 |  | Test Equipment Used: | Piranha CB2-17090320
8 | [11 empty] | Test Equipment Used: | Piranha CB2-17090320
9 | kVp and Exposure Reproducibility
10 | [10 empty] | Large Focus kV, Exposure, Timer Reproducibility
11 |  |  | Large Focal Spot |  |  |  | Small Focal Spot
12 |  |  | SDD*: | 100 | cm |  | SDD*: | 100 | cm |  | Test | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
13 |  |  | mA: | 400 |  |  | mA: | 100 |  |  | Setup | cm | Spot | Sensitivity | kV | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
14 |  |  | mAs: |  |  |  | mAs: |  |  |  |  | 100 | Large | 70-120 | 80 | 400 | 0.05
15 |  |  | time (s): | 0.05 |  |  | time (s): | 0.1 |  |  |  | 100 | Large | 70-120 | 80 | 400 | 0.05
16 | [11 empty] | 100 | Large | 70-120 | 80 | 400 | 0.05
17 |  |  | Indicated | Measured |  |  | Indicated | Measured |  |  |  | 100 | Large | 70-120 | 80 | 400 | 0.05
18 |  |  | kVp | kVp | mGy |  | kVp | kVp | mGy |  |  | 100 | Large | 70-120 | 80 | 400 | 0.05
19 |  |  | 80 |  |  |  | 90
20
21
22
23
24 | [10 empty] | Note: Type comments in the below shaded cell.
25 | [11 empty] | Do NOT exceed 190
26 | [11 empty] | New |  | 0
27 | [11 empty] | Do NOT exceed 190 | [6 empty] | kVp | mGy | Time
28 | [11 empty] | New |  | 0 |  |  |  | Acceptable: | NA | NA | NA
29 | [10 empty] | Small Focus kV, Exposure, Timer Reproducibility
30 |  |  | mean: |  |  |  | mean:
31 |  |  | std dev: |  |  |  | std dev: |  |  |  | Test | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
32 |  |  | Coeff. of var. | TBD | TBD |  | Coeff. of var. | TBD | TBD |  | Setup | cm | Spot | Sensitivity | kV | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
33 | [11 empty] | 100 | Small | 70-120/Low | 90 | 100 | 0.1
34 |  |  | Acceptable: | NA | NA |  | Acceptable: | NA | NA |  |  | 100 | Small | 70-120/Low | 90 | 100 | 0.1
35 | [11 empty] | 100 | Small | 70-120/Low | 90 | 100 | 0.1
36 |  | *SDD = Source to Detector Distance | [9 empty] | 100 | Small | 70-120/Low | 90 | 100 | 0.1
37 | [11 empty] | 100 | Small | 70-120/Low | 90 | 100 | 0.1
38 |  |  | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%.
39
40
41
42
43 | Comments:
44 | [11 empty] | Do NOT exceed 190
45 | [11 empty] | New |  | 0
46 | [11 empty] | Do NOT exceed 190
47 | [11 empty] | New |  | 0 | [5 empty] | kVp | mGy
48 | [18 empty] | Acceptable: | NA | NA
49 | [11 empty] | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%.
50
51
52
53 | [11 empty] | Combined Comments:
54 | [11 empty] | Large Focus:
55
56 | [11 empty] | Small Focus:
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [10 empty] | Radiographic System Compliance Inspection
3
4 | [14 empty] | Measurement Data
5 |  | Test Equipment Used: | Piranha CB2-17090320
6
7 | [11 empty] | Test Equipment Used: | Piranha CB2-17090320
8 | [5 empty] | Timer performance (Large Focal Spot)
9
10 | [10 empty] | Timer performance (Large Focal Spot)
11 |  |  | Generator Type: | HF | (High Frequency)
12 | [11 empty] | NOTE: Enter "Dent" for Dental, "1*" for 1-phase, "3" for 3-phase, "HF" for High Frequency, "CP" for Constant Potential
13 |  |  |  |  | Accuracy |  |  |  | Reproducibility |  |  | Generator Type: | HF | HF | (High Frequency)
14 | [12 empty] | Timer Test is Valid for this generator
15 |  | SDD: | 100 | cm |  |  |  | SDD: | 100 | cm | Test | SDD | Focal | Filter Pack/ |  | Set Values* |  |  | Measured Values*
16 |  | kVp: | 80 |  |  |  |  | kVp: | 80 |  | Setup | cm | Spot | Sensitivity | kV | mA/mAs | seconds | kV (max) | kV (avg) | seconds | mGy
17 | [7 empty] | mA: | 400 |  |  | 100 | Large | 70-120/Low | 80 | 200 | 0.01
18 | [7 empty] | mAs: |  |  |  | 100 | Large | 70-120/Low | 80 | 200 | 0.02
19 | [11 empty] | 100 | Large | 70-120/Low | 80 | 200 | 0.04
20 | Measured Values |  |  | Time (sec) |  | Error |  |  | Time (sec) |  |  | 100 | Large | 70-120/Low | 80 | 200 | 0.1
21 | kV (max) | kV (avg) | mA/mAs | Indicated | Measured | (Percent) | Acceptable |  | Indicated | Measured |  | 100 | Large | 70-120/Low | 80 | 200 | 0.25
22 |  |  | 200 | 0.01 |  |  | TBD |  | 0.05 |  |  | 100 | Large | 70-120/Low | 80 | 200 | 0.4
23 |  |  | 200 | 0.02 |  |  | TBD
24 |  |  | 200 | 0.04 |  |  | TBD
25 |  |  | 200 | 0.1 |  |  | TBD
26 |  |  | 200 | 0.25 |  |  | TBD
27 |  |  | 200 | 0.4 |  |  | TBD |  |  |  |  | *Record time in seconds for 3-phase, HF, and CP
28 | [11 empty] | For 1-phase, record time in pulses up to 0.5 sec. (60 pulses), and seconds thereafter.
29
30 | [11 empty] | Comments:
31 | [10 empty] | Note: Type comments in the above shaded cell.
32 | [11 empty] | Additional comments:
33 | [8 empty] | mean: |  |  | New |  | 0
34 | [8 empty] | std dev: | [6 empty] | Time (sec) |  | Error | 3-Phase | 1-Phase | Dent
35 | [8 empty] | Coeff. of var.: | TBD | [5 empty] | Indicated | Measured | (Percent) | Acceptable | Acceptable | Acceptable
36 | [15 empty] | 0.01
37 | [5 empty] | Acceptable: | NA |  | Acceptable: | NA | [5 empty] | 0.02
38 | Comments: | [14 empty] | 0.04
39 | [15 empty] | 0.1
40 | [15 empty] | 0.25
41 | [15 empty] | 0.4
42 | [13 empty] | Combined Comments
43
44
45 |  | Timer Specifications
46 |  | Generator |  |  |  | Generator | [5 empty] | Timer Specifications
47 |  | Type | time (ms) | error |  | Type | mAs | error |  |  |  | Generator |  |  |  | Generator
48 |  | 1f | < =4 pulses | 0 | pulses | Cine | < 2 | 15% |  |  |  | Type | time (ms) | error |  | Type | mAs | error
49 |  |  | 5-24 pulses | 1 | pulses |  | 2 - 5 | 10% |  |  |  | 1f | < =4 pulses | 0 | pulses | Cine | < 2 | 15%
50 |  |  | 25-60 pulses | 2 | pulses |  | > 5 | 5% |  |  |  |  | 5-24 pulses | 1 | pulse |  | 2 - 5 | 10%
51 |  |  | >60 pulses | 3 | pulses | mAs timer | < 2 | 20% |  |  |  |  | 25-60 pulses | 2 | pulses |  | > 5 | 5%
52 |  | 3f | < 2 | 25% |  |  | 2 - 5 | 15% |  |  |  |  | >60 pulses | 3 | pulses | mAs timer | < 2 | 20%
53 |  |  | 2 - 10 | 20% |  |  | 5 - 100 | 10% |  |  |  | 3f | < 2 | 25% |  |  | 2 - 5 | 15%
54 |  |  | 10 - 50 | 15% |  |  | > 100 | 5% |  |  |  |  | 2 - 10 | 20% |  |  | 5 - 100 | 10%
55 |  |  | 50 - 100 | 10% | [8 empty] | 10 - 50 | 15% |  |  | > 100 | 5%
56 |  |  | > 100 | 5% | [8 empty] | 50 - 100 | 10%
57 |  | Dental | All | +2 | pulses | [7 empty] | > 100 | 5%
58 |  |  | All | +15% | [7 empty] | Dental | All | +2 | pulses
59 | [12 empty] | All | +15%
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection
3
4
5 |  | Test Equipment Used: | Piranha CB2-17090320 | [8 empty] | Test Equipment Used: | Piranha CB2-17090320
6
7 | [10 empty] | Large Focus Linearity
8 | Large Focus Linearity |  |  | kVp: | 80 | kV filter: | 70-120/Low
9 |  |  |  | SDD: | 100 | cm |  |  | Large Focus
10 | Test |  | Set |  | Measured values |  |  | mGy/mAs | Coeff. of |  | mA/mAs | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
11 | Setup | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | @ SDD | Linearity | Acceptable | Setup | cm | Spot | Sensitivity | kV | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
12 | Adjacent* | 400 | 0.05 | [6 empty] | Adjacent | Adjacent* | 100 | Large | 70-120/Low | 80 | 400 | 0.05
13 | Adjacent* | 500 | 0.05 | [6 empty] | TBD | Adjacent* | 100 | Large | 70-120/Low | 80 | 500 | 0.05
14 | Non-Adj. | 50 | 0.05 | [7 empty] | Non-Adj. | 100 | Large | 70-120/High | 80 | 50 | 0.05
15 | Non-Adj. | 800 | 0.05 | [6 empty] | Non-Adj. | Non-Adj. | 100 | Large | 70-120/Low | 80 | 800 | 0.05
16 | Non-Adj. | 250 | 0.05 | [6 empty] | TBD | Non-Adj. | 100 | Large | 70-120/Low | 80 | 250 | 0.05
17 | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X. | [9 empty] | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
18 | [19 empty] | Acceptable
19 | Comments: | [10 empty] | Comments: | [7 empty] | Adjacent
20 | [11 empty] | Do NOT exceed 390 |  | 0 | [5 empty] | TBD
21 | [11 empty] | New | [7 empty] | Non-Adj.
22 | [19 empty] | TBD
23 | Small Focus Linearity |  |  | kVp: | 80 | kV filter: | 70-120/Low |  |  |  | Small Focus Linearity
24 |  |  |  | SDD: | 100 | cm |  |  | Small Focus
25 | Test |  | Set |  | Measured values |  |  | mGy/mAs | Coeff. of |  | mA/mAs | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
26 | Setup | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | @ SDD | Linearity | Acceptable | Setup | cm | Spot | Sensitivity | kV | mA | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
27 | Adjacent* | 100 | 0.1 | [6 empty] | Adjacent | Adjacent* | 100 | Small | 70-120/Low | 80 | 100 | 0.1
28 | Adjacent* | 110 | 0.1 | [6 empty] | TBD | Adjacent* | 100 | Small | 70-120/Low | 80 | 110 | 0.1
29 | Non-Adj. | 250 | 0.1 | [7 empty] | Non-Adj. | 100 | Small | 70-120/High | 80 | 250 | 0.1
30 | Non-Adj. | 50 | 0.1 | [6 empty] | Non-Adj. | Non-Adj. | 100 | Small | 70-120/High | 80 | 50 | 0.1
31 | Non-Adj. | 160 | 0.1 | [6 empty] | TBD | Non-Adj. | 100 | Small | 70-120/Low | 80 | 160 | 0.1
32 | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X. | [9 empty] | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
33 | [19 empty] | Acceptable
34 | Comments: | [10 empty] | Comments: | [7 empty] | Adjacent
35 | [11 empty] | Do NOT exceed 390 |  | 0 | [5 empty] | TBD
36 | [11 empty] | New | [7 empty] | Non-Adj.
37 | [19 empty] | TBD
38 | Combined Focal Spot Linearity
39 |  |  |  | Combined
40 |  | mGy/mAs |  | Coeff. of
41 |  | @ SDD |  | Linearity |  | Acceptable | [5 empty] | Coefficient of Linearity | [7 empty] | Combined Focal Spot Linearity
42 |  | Max.: | 0 | Non-Adj. |  | Non-Adj. | [6 empty] | Adjacent: |  | (mGy/mAs)adj1 - (mGy/mAs)adj2
43 |  | Min.: | 0 | NA |  | NA | [8 empty] | (mGy/mAs)adj1 + (mGy/mAs)adj2 |  |  |  |  | Acceptable
44 | [19 empty] | Non-Adj.
45 |  | Coefficient of Linearity | [10 empty] | Non-Adjacent: |  | (mGy/mAs)max - (mGy/mAs)min |  |  |  |  | NA
46 |  |  | Adjacent: |  | (mGy/mAs)adj1 - (mGy/mAs)adj2 | [9 empty] | (mGy/mAs)max + (mGy/mAs)min
47 |  |  |  |  | (mGy/mAs)adj1 + (mGy/mAs)adj2
48 | [11 empty] | Criteria: | Coefficient of linearity <= 10% (Both adjacent and non-adjacent).
49 |  |  | Non-Adjacent: |  | (mGy/mAs)max - (mGy/mAs)min
50 |  |  |  |  | (mGy/mAs)max + (mGy/mAs)min
51
52 |  | Criteria: | Coefficient of linearity <= 10% (Both adjacent and non-adjacent).
53 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA
54 | Rule Number |  |  |  | Generator Functions |  |  |  | Compliance | [5 empty] | Generator Functions
55 | DHEC RHB 4.3.9.1 |  | Means are provided to indicate technique factors prior to exposure. | [5 empty] | TBD |  |  |  | Means provided to indicate technique factors prior to exposure.
56 |  |  | All technique indicators functional and are free of excessive wear. | [5 empty] | TBD |  |  |  | All technique indicators functional and are free of excessive wear.
57 |  |  | Technique overload circuits function to block excessive technique. | [5 empty] | TBD |  |  |  | Technique overload circuits function to block excessive technique
58 |  |  | Anode heat calculator/indicator is functioning properly. | [5 empty] | TBD |  |  |  | Anode heat calculator/indicator is functioning properly
59 | DHEC RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational. | [5 empty] | TBD |  |  |  | Indicator of the individual x-ray tube selected present and operational
60 | DHEC RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational. | [5 empty] | TBD |  |  |  | Visible "beam on" indicator present and operational
61 | DHEC RHB 4.7.4.2.4 |  | Audible exposure end indicator present and operational. | [5 empty] | TBD |  |  |  | Audible exposure end indicator present and operational
62 |  |  | Exposure delay time is not excessive. | [5 empty] | TBD |  |  |  | Exposure delay time is not excessive.
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection | [5 empty] | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection
3 |  |  |  |  | HALF VALUE LAYER MEASUREMENT | [5 empty] | 2 | Change to DHEC Standards?
4 | [10 empty] | (Type 1 to base acceptance on DHEC criteria)
5 | [17 empty] | Measured HVL based on:
6 | [10 empty] | Half Value Layer Measurement |  |  |  | Push button to select if HVL is |  |  |  | Max kV
7 | [14 empty] | based on Max kV or Avg kV
8 |  |  | Large Focal Spot | [8 empty] | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
9 |  | SDD*: | 100 | cm |  |  | mA: | 400 |  |  | mm Al | cm | Spot | Sensitivity | kV | mA/mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
10 |  | Set kVp: | 80 |  |  |  | mAs: |  |  |  | 0 | 100 | Large | 70-120/Low | 80 | 400 | 0.05
11 |  |  |  |  | Set |  | time (s): | 0.05 |  |  | 0 | 100 | Large | 70-120/Low | 80 | 400 | 0.05
12 |  | Al (mm) | mGy | mGy | kVp | [5 empty] | 3 | 100 | Large | 70-120/Low | 80 | 400 | 0.05
13 |  | 0 |  |  | 80 |  | kVp (Max kV) | 80 |  |  | 3 | 100 | Large | 70-120/Low | 80 | 400 | 0.05
14 |  | 3 |  |  |  |  | Meas. HVL: | TBD
15 | [6 empty] | Meets min HVL: | TBD
16
17 | [6 empty] | Criteria: MUSC STANDARD
18 | [6 empty] | HVL | MIN. HVL | MAX. HVL
19 |  | * SDD = Source to Detector Distance. |  |  |  | kVp | (mm Al) | (mm Al) | (mm Al)
20 | [5 empty] | 60 | 2.1855 | 1.9 | 2.4
21 | [5 empty] | 70 | 2.6109 | 2.3 | 2.9
22 |  |  | **ACCEPTABLE RESULTS (mm Al) |  |  | 80 | 3.0364 | 2.6 | 3.4 |  |  | Comments:
23 |  |  |  | Min. acceptable HVL at 80 kvp: | 2.6 | 90 | 3.4618 | 3 | 3.8 |  |  | Do NOT exceed 390 |  | 0
24 |  |  |  | Desired HVL at meas. 80 kVp: | 3 | 100 | 3.8873 | 3.4 | 4.3 |  |  | New
25 |  |  |  | Max. acceptable HVL at 80 kvp: | 3.4 | 110 | 4.3127 | 3.8 | 4.8 |  | Data Processing: |  |  |  | Max kV | [8 empty] | RADIOGRAPHIC UNITS (MUSC CRITERIA) |  |  |  |  | RAD / DENTAL UNITS (DHEC CRITERIA)
26 | [5 empty] | 120 | 4.7382 | 4.2 | 5.2 |  | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp |  | HVL @80 kVp (Max kV) |  | Criteria: MUSC STANDARD |  |  |  |  | Criteria: |  |  |  |  | Criteria:
27 | [5 empty] | 130 | 5.1636 | 4.6 | 5.7 |  | 0 |  |  |  | 80 |  | TBD |  | HVL | MIN. HVL | MAX. HVL |  |  | HVL | MIN. HVL | MAX. HVL |  |  | HVL | MIN. HVL | MAX. HVL
28 | Comments: | [9 empty] | 3 | [6 empty] | kVp | (mm Al) | (mm Al) | (mm Al) |  | kVp | (mm Al) | (mm Al) | (mm Al) |  | kVp | (mm Al) | (mm Al) | (mm Al)
29 | [15 empty] | Acceptable (mm Al) |  | 60 | 2.1855 | 1.9 | 2.4 |  | 60 | 2.1855 | 1.9 | 2.4 |  | 50 | 1.5 | 1.5 | 1.7
30 | [13 empty] | Data Exp | LN(EXP) | LOW: | 2.6 | 70 | 2.6109 | 2.3 | 2.9 |  | 70 | 2.6109 | 2.3 | 2.9 |  | 60 | 1.5 | 1.5 | 1.7
31 | [15 empty] | DESIRED: | 3 | 80 | 3.0364 | 2.6 | 3.4 |  | 80 | 3.0364 | 2.6 | 3.4 |  | 70 | 1.5 | 1.5 | 1.7
32 | [15 empty] | HIGH: | 3.4 | 90 | 3.4618 | 3 | 3.8 |  | 90 | 3.4618 | 3 | 3.8 |  | 80 | 2.3 | 2.3 | 2.7
33 | [10 empty] | TBD | TBD | [5 empty] | 100 | 3.8873 | 3.4 | 4.3 |  | 100 | 3.8873 | 3.4 | 4.3 |  | 90 | 2.5 | 2.5 | 3
34 | [15 empty] | Acceptable: | TBD | 110 | 4.3127 | 3.8 | 4.8 |  | 110 | 4.3127 | 3.8 | 4.8 |  | 100 | 2.7 | 2.7 | 3.2
35 | [15 empty] | Meets minimum HVL criteria: | TBD | 120 | 4.7382 | 4.2 | 5.2 |  | 120 | 4.7382 | 4.2 | 5.2 |  | 110 | 3 | 3 | 3.5
36 | [17 empty] | 130 | 5.1636 | 4.6 | 5.7 |  | 130 | 5.1636 | 4.6 | 5.7 |  | 120 | 3.2 | 3.2 | 3.75
37
38 | [10 empty] | 0 kV |  |  |  |  | Last Year:
39 | [10 empty] | mm Al. | Avg. Exp. | [9 empty] | kV | HVL
40 | [10 empty] | #NUM! | #NUM! | [9 empty] | 0 | #DIV/0!
41 | [10 empty] | #NUM! | #NUM! | [9 empty] | 80 | TBD
42 | [10 empty] | #NUM! | #NUM! | [9 empty] | 0 | #DIV/0!
43 | [10 empty] | #N/A | #N/A | [9 empty] | 0 | #DIV/0!
44 | [21 empty] | 0 | 0
45 | [10 empty] | 0 kV | [10 empty] | 0 | 0
46 | [10 empty] | mm Al. | Avg. Exp. | [9 empty] | #DIV/0! | #DIV/0!
47 | [10 empty] | #NUM! | #NUM! | [9 empty] | 0 | 0
48 | [10 empty] | #NUM! | #NUM! | [9 empty] | #DIV/0! | TBD
49 | [10 empty] | #NUM! | #NUM! | [9 empty] | #DIV/0! | #DIV/0!
50 | [10 empty] | #N/A | #N/A | [9 empty] | #DIV/0! | #DIV/0!
51 | [21 empty] | 0 | 0
52 | [10 empty] | 80 kV | [10 empty] | #DIV/0! | #DIV/0!
53 | [10 empty] | mm Al. | Avg. Exp. | [9 empty] | 0 | 0
54 | [10 empty] | TBD | TBD | [9 empty] | 0 | 0
55 | [10 empty] | TBD | TBD
56 | [10 empty] | TBD | TBD
57 | [10 empty] | TBD | TBD
58 | [10 empty] | TBD | TBD
59 | [10 empty] | TBD | TBD
60
61 | [10 empty] | 0 kV
62 | [10 empty] | mm Al. | Avg. Exp.
63 | [10 empty] | #NUM! | #NUM!
64 | [10 empty] | #NUM! | #NUM!
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  | #NUM! | #NUM!
66 | Room Number: | [6 empty] | Survey ID: |  |  | #N/A | #N/A
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  |  |  | Measurement Data | [5 empty] | Radiographic System Compliance Inspection
3
4
5 |  |  |  |  | PHOTOTIMING | [10 empty] | PHOTOTIMING
6
7
8
9 |  |  |  |  | DENSITY CONTROL FUNCTION | [5 empty] | Density Control Function |  |  |  | Meter
10 | [10 empty] | New Values |  |  |  | mAs
11 | [10 empty] | 101.6 |  |  |  |  | Unit has phototimer density control
12 | SID | Detector |  | Density | Meter | % change | mAs | % change |  |  | SID | Detec. |  | Density | Meter | % change | mAs | % change | Either/Or | Failure
13 | 101.6 cm | C |  | 4 | [6 empty] | 101.6 cm | C |  | 4
14 |  |  |  | 3 | [9 empty] | 3
15 | kVp | Thickness |  | 2 | [6 empty] | New Values |  |  | 2
16 | 80 | 7.874 | in. | 1 | [6 empty] | 80 | 7.874 |  | 1
17 |  |  |  | 0 | [6 empty] | kVp | Thickness |  | 0
18 |  |  |  | -1 | [6 empty] | 80 | 7.874 | in. | -1
19 |  |  |  | -2 | [9 empty] | -2
20 |  |  |  | -3 | [9 empty] | -3
21 |  |  |  | -4 | [9 empty] | -4
22 | [8 empty] | Acceptable: | TBD | [9 empty] | Acceptable
23 | [12 empty] | Criteria: |  | Minimum Step: | 15% |  | Maximum Step: | 50% | Meter: | TBD
24 |  |  | Criteria: |  | Minimum Step: | 15% |  | Maximum Step: | 50% | [10 empty] | mAs: | TBD
25 | [19 empty] | Either/Or: | TBD
26
27
28
29
30
31
32 |  |  |  |  | PHOTOTIMING CONTROL | [10 empty] | PHOTOTIMING CONTROL
33
34
35 | Rule Number | [7 empty] | Compliance |  | Enter 1 for YES, 2 for NO, 3 for NA
36 | 21CFR 1020.31(a)(3)(i) |  | Indication given on control panel when AEC is selected. | [5 empty] | TBD |  |  |  | Indication given on control panel when AEC is selected.
37 | 21CFR 1020.31(a)(3)(iii) |  | Backup limit is <= 600 mAs or 60 kWs |  |  |  | Limit: |  | TBD |  |  | Limit: |  | Backup limit is <= 600 mAs or 60 kWs |  |  |  |  | (Last Year)
38 | 21CFR 1020.31(a)(3)(iv) |  | Visible signal indicates backup limit and subsequent exposures blocked. | [5 empty] | TBD |  |  |  | Visible signal indicates backup limit and subsequent exposures blocked.
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  |  |  | Measurement Data | [5 empty] | Radiographic System Compliance Inspection
3
4 |  |  |  |  | PHOTOTIMING TABLE BUCKY | [10 empty] | PHOTOTIMING TABLE BUCKY
5 | Typical Optical Density | [9 empty] | Typical Optical Density
6 | [14 empty] | Last year | New |  |  |  | Last Year:
7 | SID: 101.6 cm | Detector: C |  | Atten. Thick: 7.87 |  | kVp: 80 | O.D.:_____ | Acceptable: | NA |  |  |  |  | SID: | 101.6 cm |  |  | mAs: |  | Meter | NA
8 | [5 empty] | mAs:_______ | [7 empty] | Detector: | C |  |  |  |  | Acceptable: | NA
9 |  | Criteria: | 1.9<=lgM<=2.1 | [10 empty] | kVp: | 80 |  |  | Criteria: | 1.9<=lgM<=2.1
10 | [13 empty] | Atten. Thick: | 7.874
11
12 | Detector Balance and Reproducibility |  |  | kVp: | 80 |  | Atten. thickness: | 7.87 in. |  |  | Detector Balance and Reproducibility |  |  |  | kVp: | 80 |  | Atten. thickness: | 7.874 | in.
13 |  |  |  | SID: | 101.6 cm |  | Film/Screen: | [7 empty] | SID: | 101.6 cm |  | Film/Screen:
14 | Chamber: | LEFT |  | CENTER |  | RIGHT | [5 empty] | Chamber: | LEFT |  | CENTER |  | RIGHT
15 |  | Meter | mAs | Meter | mAs | Meter | mAs | [5 empty] | Meter | mAs | Meter | mAs | Meter | mAs
16
17
18
19
20
21
22 | mean: | [7 empty] | Detector Balance |  |  | mean: | [7 empty] | Detector Balance
23 | stdev: | [7 empty] | Max. var: |  |  | stdev: | [7 empty] | Max. var:
24 | Coeff. of Var.: | [7 empty] | TBD |  |  | Coeff. of Var.: | [7 empty] | TBD
25 | Acceptable: | TBD | TBD | TBD | TBD | TBD | TBD | Acceptable: | TBD |  |  | Acceptable: | TBD | TBD | TBD | TBD | TBD | TBD | Acceptable: | TBD
26 | [8 empty] | NOTE: Density step 1 produced a TBD % change | [10 empty] | NOTE: Density step 1 produced a TBD % change
27 |  |  |  | Criteria: Detector Balance - | Film: Maximum variation of 0.4 O.D. | [9 empty] | Criteria: Detector Balance - | Film: Maximum variation of 0.4 O.D.
28 |  |  |  |  | Exposure & AEC Meter: Max. Variation 35% | [10 empty] | Exposure & AEC Meter: Max. Variation 35%
29 |  |  |  | Reproducibility - | Coefficient of variation <= 5% for each detector | [9 empty] | Reproducibility - | Coefficient of variation <= 5% for each detector
30
31 | kVp and Thickness Compensation
32 |  | SID: | 101.6 cm |  | Detector: | C |  | Backup mAs: |  |  |  | kVp and Thickness Compensation
33 | Attenuator | kVp1: | 60 | kVp2: | 80 | kVp3: | 100 | kVp4: | 120 |  |  |  | SID: | 101.6 cm |  | Detector: | C |  | Backup mAs:
34 | thickness (in) | Meter | mAs | Meter | mAs | Meter | mAs | Meter | mAs |  |  | Attenuator | kVp1: | 60 | kVp2: | 80 | kVp3: | 100 | kVp4: | 120
35 | 10 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | thick. (cm) | Meter | mAs | Meter | mAs | Meter | mAs | Meter | mAs
36 | 15 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | 10
37 | 20 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | 15
38 | 25 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | 20
39 | [11 empty] | 25
40 | Vary Atten: | @60 kVp: |  | @80 kVp: |  | @100 kVp: |  | @120 kVp:
41 | Acceptable: | NA |  | NA |  | NA |  | NA |  |  | TABLE | Vary Atten: | @60 kVp: |  | @80 kVp: |  | @100 kVp: |  | @120 kVp:
42 | Vary kVp: | @ 10 cm.: |  | @ 15 cm.: |  | @ 20 cm.: |  | @ 25 cm.: |  |  |  | Acceptable: | NA |  | NA |  | NA |  | NA |  |  | Max | 0
43 | Acceptable: | NA |  | NA |  | NA |  | NA |  |  |  | Vary kVp: | @ 10 cm.: |  | @ 15 cm.: |  | @ 20 cm.: |  | @ 25 cm.: |  |  | Min | 0
44 |  |  |  |  | Overall Variation: |  | Acceptable: | NA |  |  |  | Acceptable: | NA |  | NA |  | NA |  | NA
45 |  | Criteria: | Film-Max difference 0.3 O.D. for individual kV or thickness; 0.4 O.D. overall | [12 empty] | Overall Variation: |  | Acceptable: | NA
46 |  |  | Exposure & AEC Meter- | [16 empty] | NOTE: Density step 1 produced a TBD % change
47 | Comments: | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall | [10 empty] | Criteria: | Film-Max difference 0.3 O.D. for individual kV or thickness; 0.4 O.D. overall
48 | [13 empty] | Exposure & AEC Meter- | Max. Variation 30% for individual kV or thickness; 35% overall
49 | [11 empty] | Comments: | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall
50 | [11 empty] | Do NOT exceed 390 |  | 86
51 | [11 empty] | New | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall
52
53 | [11 empty] | Maximum Change:
54 | [11 empty] | 60 | 0 | 80 | 0 | 100 | 0 | 120 | 0
55 | [11 empty] | 10 | 0 | 15 | 0 | 20 | 0 | 25 | 0
56 | [13 empty] | Overall: | 0
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  |  |  | Measurement Data | [5 empty] | Radiographic System Compliance Inspection
3 | [14 empty] | PHOTOTIMING WALL BUCKY
4 |  |  |  |  | PHOTOTIMING WALL BUCKY | [10 empty] | Last year | New
5 | Typical Optical Density | [9 empty] | Typical Optical Density |  |  |  | SID: | 101.6 cm
6 | Type: | [9 empty] | Enter 1 for YES, 2 for NO, 3 for NA |  |  |  | Detector: | C |  |  |  | Last Year:
7 | SID: 101.6 cm | Detector: C |  | Atten. Thick: 7.87 |  | kVp: 80 | O.D.:_____ | Acceptable: | NA |  |  | Gen. Radiog. Imaging |  |  | kVp: | 80 |  | mAs: |  | Meter | NA
8 | [5 empty] | mAs:_______ | [5 empty] | Chest Imaging |  |  | Atten. Thick: | 7.874 |  | Criteria |  | Acceptable: | NA
9 |  |  | Criteria: Radiographic: | 1.9<=lgM<=2.1 | [13 empty] | Radiographic: | 1.9<=lgM<=2.1
10 |  |  | Chest: | 1.9<=lgM<=2.1 | [13 empty] | Chest: | 1.9<=lgM<=2.1
11
12 | Detector Balance and Reproducibility |  |  | kVp: | 80 |  | Atten. thickness: | 7.87 in. |  |  | Detector Balance and Reproducibility |  |  |  | kVp: | 80 |  | Atten. thickness: | 7.874 | in.
13 |  |  |  | SID: | 101.6 cm |  | Film/Screen: | [7 empty] | SID: | 101.6 cm |  | Film/Screen:
14 | Chamber: | LEFT |  | CENTER |  | RIGHT | [5 empty] | Chamber: | LEFT |  | CENTER |  | RIGHT
15 |  | Meter | mAs | Meter | mAs | Meter | mAs | [5 empty] | Meter | mAs | Meter | mAs | Meter | mAs
16
17
18
19
20
21
22 | mean: | [7 empty] | Detector Balance |  |  | mean: | [7 empty] | Detector Balance
23 | stdev: | [7 empty] | Max. var: |  |  | stdev: | [7 empty] | Max. var:
24 | Coeff. of Var.: | [7 empty] | TBD |  |  | Coeff. of Var.: | [7 empty] | TBD
25 | Acceptable: | TBD | TBD | TBD | TBD | TBD | TBD | Acceptable: | TBD |  |  | Acceptable: | TBD | TBD | TBD | TBD | TBD | TBD | Acceptable: | TBD
26 | [8 empty] | NOTE: Density step 1 produced a TBD % change | [10 empty] | NOTE: Density step 1 produced a TBD % change
27 |  |  |  | Criteria: Detector Balance - | Film: Maximum variation of 0.4 O.D. | [9 empty] | Criteria: Detector Balance - | Film: Maximum variation of 0.4 O.D.
28 |  |  |  |  | Exposure & AEC Meter: Max. Variation 35% | [10 empty] | Exposure & AEC Meter: Max. Variation 35%
29 |  |  |  | Reproducibility - | Coefficient of variation <= 5% for each detector | [9 empty] | Reproducibility - | Coefficient of variation <= 5% for each detector
30
31 | kVp and Thickness Compensation
32 |  | SID: | cm |  | Detector: | C |  | Backup mAs: |  |  |  | kVp and Thickness Compensation
33 | Attenuator | kVp1: | 60 | kVp2: | 80 | kVp3: | 100 | kVp4: | 120 |  |  |  | SID: | cm |  | Detector: | C |  | Backup mAs:
34 | thickness (in) | Meter | mAs | Meter | mAs | Meter | mAs | Meter | mAs |  |  | Attenuator | kVp1: | 60 | kVp2: | 80 | kVp3: | 100 | kVp4: | 120
35 | 10 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | thick. (cm) | Meter | mAs | Meter | mAs | Meter | mAs | Meter | mAs
36 | 15 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | 10
37 | 20 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | 15
38 | 25 | #N/A |  | #N/A |  | #N/A |  | #N/A |  |  |  | 20
39 | [11 empty] | 25
40 | Vary Atten: | @60 kVp: |  | @80 kVp: |  | @100 kVp: |  | @120 kVp:
41 | Acceptable: | NA |  | NA |  | NA |  | NA |  |  | WALL | Vary Atten: | @60 kVp: |  | @80 kVp: |  | @100 kVp: |  | @120 kVp:
42 | Vary kVp: | @ 10 cm.: |  | @ 15 cm.: |  | @ 20 cm.: |  | @ 25 cm.: |  |  |  | Acceptable: | NA |  | NA |  | NA |  | NA
43 | Acceptable: | NA |  | NA |  | NA |  | NA |  |  |  | Vary kVp: | @ 10 cm.: |  | @ 15 cm.: |  | @ 20 cm.: |  | @ 25 cm.:
44 |  |  |  |  | Overall Variation: |  | Acceptable: | NA |  |  |  | Acceptable: | NA |  | NA |  | NA |  | NA
45 |  | Criteria: | Film-Max difference 0.3 O.D. for individual kV or thickness; 0.4 O.D. overall | [12 empty] | Overall Variation: |  | Acceptable: | NA
46 |  |  | Exposure & AEC Meter- | [15 empty] | NOTE: Density step 1 produced a TBD % change
47 | Comments: | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall | [10 empty] | Criteria: | Film-Max difference 0.3 O.D. for individual kV or thickness; 0.4 O.D. overall
48 | [13 empty] | Exposure & AEC Meter- | Max. Variation 30% for individual kV or thickness; 35% overall
49 | [11 empty] | Comments: | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall
50 | [11 empty] | Do NOT exceed 390 |  | 86
51 | [11 empty] | New | Criteria: lgM - Max difference 0.2 lgM for individual kV or thickness; 0.2 lgM overall
52
53 | [11 empty] | Maximum Change:
54 | [11 empty] | 60 | 0 | 80 | 0 | 100 | 0 | 120 | 0
55 | [11 empty] | 10 | 0 | 15 | 0 | 20 | 0 | 25 | 0
56 | [13 empty] | Overall: | 0
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  |  |  | Measurement Data | [5 empty] | Radiographic System Compliance Inspection
3
4
5 |  |  |  |  | Kerma Area Product/Air Kerma Accuracy | [10 empty] | Kerma Area Product/Air Kerma Accuracy
6
7
8 | [10 empty] | 100 | Source-PDC distance (cm) |  |  |  | Source-Isocenter
9 |  | Test equipment: | PDC |  | Calibration date: |  |  | Calibration due: |  |  |  | Source-Ref Pt distance (cm) |  |  |  | Table-Isocenter
10 | [11 empty] | AK Ref Point correction factor |  |  |  | Ref Pt-Isocenter
11 |  |  |  | Ind KAP µGy·m² | Meas KAP µGy·m² | KAP Deviation | Ind AK @ Ref Pt | Meas AK @ Ref Pt | AK Deviation
12 |  | kV | mA | [9 empty] | Indicated KAP |  |  | Indicated AK
13 | [12 empty] | Start | End |  | Start AK | End AK
14 | [10 empty] | kV | mA | µGy·m² |  | KAP | mGy |  | AK
15
16
17
18
19
20
21 |  |  |  |  | Acceptable: | NA |  | Acceptable: | NA
22 |  | Criteria: | Measured KAP, AK is within 30% of the displayed value
23
24 | [10 empty] | Radcal PDC |  |  | Test equipment used: | PDC |  | Calibration Date: |  |  | Calibration Due:
25 | [10 empty] | kV | mA | mGy | mGy/min | µGy·m² | µGy·m²/min
26 | [17 empty] | Note: | Set Radcal PDC to High Rate mode
27
28 | [18 empty] | KAP Conversion Table
29 | [19 empty] | µGy·m²
30 | [18 empty] | mGy·cm² | 0.1
31 | [18 empty] | cGy·cm² | 1
32 | [18 empty] | dGy·cm² | 10
33 | [18 empty] | Gy·cm² | 100
34 | [18 empty] | µGy·m² | 1
35 | [10 empty] | Ind KAP µGy·m² | Meas KAP µGy·m² | KAP Deviation | Ind AK @ Ref Pt | Meas AK @ Ref Pt | AK Deviation
36
37
38
39
40
41
42
43
44
45 | [11 empty] | Acceptable: | NA |  | Acceptable: | NA
46 | [11 empty] | Criteria: | Measured KAP, AK is within 30% of the displayed value
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiographic System Compliance Inspection |  |  |  |  | Measurement Data | [5 empty] | Radiographic System Compliance Inspection
3 | Output - Large Focal Spot |  |  | Current | Previous
4 |  |  |  | mGy/mAs | mGy/mAs
5 |  |  | Current | @ Current | @ Current | [5 empty] | Output - Large Focal Spot | [5 empty] | Calculated |  | (Using |  |  | Using | [9 empty] | Using
6 |  |  | Meas. | kVp | kVp | Percent | [9 empty] | Current | Previous |  | Previous | (Using | (Using | Current | [8 empty] | (Using | Current
7 |  |  | kVp | @40" | @40" | Change | Acceptable |  |  |  | Previous* Measurements |  |  | Current Meas. |  | mGy/mAs | mGy/mAs |  | mGy/mAs | Recorded | Current | mGy/mAs | [8 empty] | Current | mGy/mAs
8 | [10 empty] | SDD (in.): |  |  | SDD (cm): | 100 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current | [8 empty] | Measured | @ Current
9 | [11 empty] | mGy/mAs | mGy/mAs |  | mGy/mAs | kVp | meas. kVp |  | 40" SID) | kVp) | kVp) | kVp | [8 empty] | kVp) | kVp
10 | [10 empty] | kVp | @SDD | @40" | kV (max) | @SDD | @40" | @40" |  | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40" | [8 empty] | LN(kVp) | @40"
11
12
13
14
15 | [7 empty] | Acceptable: | NA
16 | Output - Small Focal Spot |  |  | Current | Previous
17 |  |  |  | mGy/mAs | mGy/mAs
18 |  |  | Current | @ Current | @ Current
19 |  |  | Meas. | kVp | kVp | Percent |  |  |  |  | *From previous compliance report |  |  |  | Max | Median | Min.
20 |  |  | kVp | @40" | @40" | Change | Acceptable | [7 empty] | NA | NA | NA |  | Acceptable: | NA
21 | [13 empty] | Last Year:
22 | [10 empty] | Output - Small Focal Spot | [5 empty] | Calculated |  | (Using |  |  | Using
23 | [15 empty] | Current | Previous |  | Previous | (Using | (Using | Current
24 | [10 empty] | Previous* Measurements |  |  | Current Meas. |  | mGy/mAs | mGy/mAs |  | mGy/mAs | Recorded | Current | mGy/mAs
25 | [10 empty] | SDD (in.): |  |  | SDD (cm): | 100 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current
26 | [7 empty] | Acceptable: | NA |  |  | mGy/mAs | mGy/mAs |  | mGy/mAs | kVp | meas. kVp |  | 40" SID) | kVp) | kVp) | kVp
27 | [10 empty] | kVp | @SDD | @40" | kV (max) | @SDD | @40" | @40" |  | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40"
28 |  | Criteria: | The x-ray output @40" within +/-25% from previous measurements.
29 | Comments:
30
31
32
33 |  |  | Lrg. Focus |  |  |  |  | Sm. Focus
34 |  |  | Current |  |  |  |  | Current |  |  | *From previous compliance report |  |  |  | Max | Median | Min.
35 |  |  | mGy/mAs |  |  |  |  | mGy/mAs | [6 empty] | NA | NA | NA |  | Acceptable: | NA
36 |  | kVp | @40" |  |  |  | kVp | @40" | [5 empty] | Last Year:
37 |  | 55 |  |  |  |  | 45
38 |  | 60 |  |  |  |  | 55 |  |  |  | Enter 1 for YES, 2 for NO, 3 for NA
39 |  | 70 |  |  |  |  | 65 | [7 empty] | Radiation Safety
40 |  | 80 |  |  |  |  | 75 | [5 empty] | Continuous pressure on exposure switch required to make an exposure with selected tube
41 |  | 90 |  |  |  |  | 85 |  |  |  | 3 |  | Spot film exposure blocked when there is no cassette
42 |  | 100 |  |  |  |  | 95 |  |  |  | 3 |  | Spot film exposure blocked when serial exposures are completed
43 |  | 110 |  |  |  |  | 105 |  |  |  | 3 |  | Spot film device provides acceptable centering to the image receptor
44 |  | 120 |  |  |  |  | 115 | [5 empty] | Exposure switch for selected tube is secured at a safe location | [6 empty] | NEW
45 |  | 130 |  |  |  |  | 125 | [5 empty] | Operator has full visibility of the exposure factors and patient | [7 empty] | Half
46 |  | 140 |  |  |  |  | 135 | [5 empty] | Lead apron available |  | Number: | 0 | Half | 0 | Full |  | Full
47 | [12 empty] | Lead gloves available. |  | Number: | 0 | (Left + Right) |  |  |  | (Left + Right)
48 | [12 empty] | A properly designed and installed apron rack is present.
49 | [12 empty] | Documentation of annual protective apparel integrity inspection is available.
50 | [12 empty] | Patient restraint devices available.
51 | [12 empty] | Gonadal shielding is present and functional.
52 | [12 empty] | Radiation warning light at room entrance is functional with selected tube
53
54 | [14 empty] | Physical Inspection
55 | [12 empty] | System clean and free of debris.
56 | [12 empty] | All cables free from excessive wear or restraints.
57 | [12 empty] | System free of missing or damaged components.
58 | [12 empty] | Table movements and locks acceptable
59 | [12 empty] | Table bucky cassette tray movements and locks acceptable
60 | [12 empty] | Wall bucky movements and locks acceptable
61 | [12 empty] | Wall bucky cassette tray movements and locks acceptable
62 | [12 empty] | Tube crane movements and locks acceptable.
63 | [12 empty] | All other movements and locks functional.
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 | [9 empty] | Radiographic System Compliance Inspection
3
4 |  |  |  |  | Measurement Data
5
6
7 | [7 empty] | Exp
8 | kVp Set | mA Set | Time | Focal Spot | kVp Avg | kVp Eff | kVp Max | Time (s) | Exposure (mGy) | Wave File
9 | 60 | 400 | 0.05 | Large
10 | 60 | 400 | 0.05 | Large
11 | 80 | 400 | 0.05 | Large
12 | 80 | 400 | 0.05 | Large
13 | 80 | 400 | 0.05 | Large
14 | 80 | 400 | 0.05 | Large
15 | 80 | 400 | 0.05 | Large
16 | 80 | 400 | 0.05 | Large
17 | 80 | 400 | 0.05 | Large
18 | 80 | 400 | 0.05 | Large
19 | 80 | 400 | 0.05 | Large
20 | 80 | 400 | 0.05 | Large
21 | 80 | 400 | 0.05 | Large
22 | 80 | 500 | 0.05 | Large
23 | 80 | 50 | 0.05 | Large
24 | 80 | 800 | 0.05 | Large
25 | 80 | 250 | 0.05 | Large
26 | 100 | 400 | 0.05 | Large
27 | 100 | 400 | 0.05 | Large
28 | 120 | 400 | 0.05 | Large
29 | 120 | 400 | 0.05 | Large
30 | 140 | 400 | 0.05 | Large
31 | 50 | 100 | 0.1 | Small
32 | 70 | 100 | 0.1 | Small
33 | 90 | 100 | 0.1 | Small
34 | 90 | 100 | 0.1 | Small
35 | 90 | 100 | 0.1 | Small
36 | 90 | 100 | 0.1 | Small
37 | 90 | 100 | 0.1 | Small
38 | 80 | 100 | 0.1 | Small
39 | 80 | 110 | 0.1 | Small
40 | 80 | 250 | 0.1 | Small
41 | 80 | 50 | 0.1 | Small
42 | 80 | 160 | 0.1 | Small
43 | 110 | 100 | 0.1 | Small
44 | 130 | 100 | 0.1 | Small
45 | 80 | 200 | 0.01 | Large
46 | 80 | 200 | 0.02 | Large
47 | 80 | 200 | 0.04 | Large
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 | [9 empty] | Radiographic System Compliance Inspection
3
4 |  |  |  |  | Measurement Data
5
6
7
8 | kVp Set | mA Set | Time | Focal Spot | Filter Set | Add Filt | Material | Distance | Units | Exp. CF. |  | METHOD TO SET THE LINE LENGTH FOR THE COMMENTS: (1) Set the maximum for line 1. (2) Vary the long word factor until the three lines are fairly evenly distributed.
9 | 60 | 400 | 0.05 | Large | 50-85 | 0 | none | 100 | cm | 1 | [5 empty] | Column |  | Note: Maximum should be 407 for the parameters specified.
10 | 60 | 400 | 0.05 | Large | 50-85 | 0 | none | 100 | cm | 1 | [5 empty] | AB | Rows |  | I wanted every body to know that now is the time for all good women to come to the aid of their country, and if you don't belive that, then know that the sly brown fox jumped quickly over the lazy dog. I wanted every body to know that now is the time for all good men to come to the aid of their country, and if you don't belive that, the know that the sly brown fox jumped quickly over the lazy dog. The end.
11 | 80 | 400 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1 |  |  |  |  | Long word factor( variable): | 6 |  |  | Number of characters: | 409
12 | 80 | 400 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1 |  |  |  |  | Max line1 (variable): | 137 | 1505 | 135 | I wanted every body to know that now is the time for all good women to come to the aid of their country, and if you don't belive that,
13 | 80 | 400 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1 |  |  |  |  | Length for two lines (lower)-Calc: | 268 | 1506 | 133 | then know that the sly brown fox jumped quickly over the lazy dog. I wanted every body to know that now is the time for all good men
14 | 80 | 400 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1 |  |  |  |  | 1st line split for two lines only-Calc: | 134 | 1507 | 141 | to come to the aid of their country, and if you don't belive that, the know that the sly brown fox jumped quickly over the lazy dog. The end.
15 | 80 | 400 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1 |  |  |  |  | 1st line split (max)-Calc: | 135.5 | 1508 | 409
16 | 80 | 400 | 0.05 | Large | 70-120 | 3 | Aluminum | 100 | cm | 1 |  |  |  |  | Length for two lines (max)-Calc: | 272.5 | 1509
17 | 80 | 400 | 0.05 | Large | 70-120 | 3 | Aluminum | 100 | cm | 1 |  |  |  |  | Count for lines 1 & 2 when 3rd present-Calc: | 266.5 | 1510
18 | 80 | 400 | 0.05 | Large | 70-120 | 3.5 | Aluminum | 100 | cm | 1
19 | 80 | 400 | 0.05 | Large | 70-120 | 3.5 | Aluminum | 100 | cm | 1
20 | 80 | 400 | 0.05 | Large | 70-120 | 2.5 | Aluminum | 100 | cm | 1
21 | 80 | 400 | 0.05 | Large | 70-120 | 2.5 | Aluminum | 100 | cm | 1
22 | 80 | 500 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1
23 | 80 | 50 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1
24 | 80 | 800 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1
25 | 80 | 250 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1
26 | 100 | 400 | 0.05 | Large | 70-120 | 0 | none | 100 | cm | 1
27 | 100 | 400 | 0.05 | Large | 100-155 | 0 | none | 100 | cm | 1
28 | 120 | 400 | 0.05 | Large | 100-155 | 0 | none | 100 | cm | 1
29 | 120 | 400 | 0.05 | Large | 100-155 | 0 | none | 100 | cm | 1
30 | 140 | 400 | 0.05 | Large | 100-155 | 0 | none | 100 | cm | 1
31 | 50 | 100 | 0.1 | Small | 35-60 | 0 | none | 100 | cm | 1
32 | 70 | 100 | 0.1 | Small | 50-85 | 0 | none | 100 | cm | 1
33 | 90 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
34 | 90 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
35 | 90 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
36 | 90 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
37 | 90 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
38 | 80 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
39 | 80 | 110 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
40 | 80 | 250 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
41 | 80 | 50 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
42 | 80 | 160 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
43 | 110 | 100 | 0.1 | Small | 70-120 | 0 | none | 100 | cm | 1
44 | 130 | 100 | 0.1 | Small | 100-155 | 0 | none | 100 | cm | 1
45 | 80 | 200 | 0.01 | Large | 70-120 | 0 | none | 100 | cm | 1
46 | 80 | 200 | 0.02 | Large | 70-120 | 0 | none | 100 | cm | 1
47 | 80 | 200 | 0.04 | Large | 70-120 | 0 | none | 100 | cm | 1
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 | [9 empty] | Radiographic System Compliance Inspection
3
4 |  |  |  |  | Measurement Data
5
6
7 | [5 empty] | Chamber |  |  | Technique
8 | kVp Set | mA Set | Time | Focal Spot | Sensitivity | Type | Delay | Wave | Comments | Survey
9 | 60 | 400 | 0.05 | Large | Low | Internal | 0
10 | 60 | 400 | 0.05 | Large | Low | Internal | 0
11 | 80 | 400 | 0.05 | Large | Low | Internal | 0
12 | 80 | 400 | 0.05 | Large | Low | Internal | 0
13 | 80 | 400 | 0.05 | Large | Low | Internal | 0 | X
14 | 80 | 400 | 0.05 | Large | Low | Internal | 0
15 | 80 | 400 | 0.05 | Large | Low | Internal | 0
16 | 80 | 400 | 0.05 | Large | Low | Internal | 0
17 | 80 | 400 | 0.05 | Large | Low | Internal | 0
18 | 80 | 400 | 0.05 | Large | Low | Internal | 0
19 | 80 | 400 | 0.05 | Large | Low | Internal | 0
20 | 80 | 400 | 0.05 | Large | Low | Internal | 0
21 | 80 | 400 | 0.05 | Large | Low | Internal | 0
22 | 80 | 500 | 0.05 | Large | Low | Internal | 0
23 | 80 | 50 | 0.05 | Large | High | Internal | 0
24 | 80 | 800 | 0.05 | Large | Low | Internal | 0
25 | 80 | 250 | 0.05 | Large | Low | Internal | 0
26 | 100 | 400 | 0.05 | Large | Low | Internal | 0
27 | 100 | 400 | 0.05 | Large | Low | Internal | 0
28 | 120 | 400 | 0.05 | Large | Low | Internal | 0
29 | 120 | 400 | 0.05 | Large | Low | Internal | 0
30 | 140 | 400 | 0.05 | Large | Low | Internal | 0
31 | 50 | 100 | 0.1 | Small | Low | Internal | 0
32 | 70 | 100 | 0.1 | Small | Low | Internal | 0
33 | 90 | 100 | 0.1 | Small | Low | Internal | 0
34 | 90 | 100 | 0.1 | Small | Low | Internal | 0
35 | 90 | 100 | 0.1 | Small | Low | Internal | 0 | X
36 | 90 | 100 | 0.1 | Small | Low | Internal | 0
37 | 90 | 100 | 0.1 | Small | Low | Internal | 0
38 | 80 | 100 | 0.1 | Small | Low | Internal | 0
39 | 80 | 110 | 0.1 | Small | Low | Internal | 0
40 | 80 | 250 | 0.1 | Small | High | Internal | 0
41 | 80 | 50 | 0.1 | Small | High | Internal | 0
42 | 80 | 160 | 0.1 | Small | Low | Internal | 0
43 | 110 | 100 | 0.1 | Small | Low | Internal | 0
44 | 130 | 100 | 0.1 | Small | Low | Internal | 0
45 | 80 | 200 | 0.01 | Large | Low | Internal | 0 | X
46 | 80 | 200 | 0.02 | Large | Low | Internal | 0 | X
47 | 80 | 200 | 0.04 | Large | Low | Internal | 0 | X
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 | [9 empty] | Radiographic System Compliance Inspection
3
4 |  |  |  |  | Measurement Data
5
6
7
8 | kVp Set | mA Set | Time | Focal Spot | Linearity | Accuracy | Beam Quality | Parameters | Reproduci... | Wave File
9 | 60 | 400 | 0.05 | Large |  | X
10 | 60 | 400 | 0.05 | Large |  |  | X
11 | 80 | 400 | 0.05 | Large | X | X |  |  | X
12 | 80 | 400 | 0.05 | Large |  |  |  |  | X
13 | 80 | 400 | 0.05 | Large |  |  |  |  | X
14 | 80 | 400 | 0.05 | Large |  |  | X |  | X
15 | 80 | 400 | 0.05 | Large |  |  | X |  | X
16 | 80 | 400 | 0.05 | Large |  |  | X
17 | 80 | 400 | 0.05 | Large |  |  | X
18 | 80 | 400 | 0.05 | Large |  |  | X
19 | 80 | 400 | 0.05 | Large |  |  | X
20 | 80 | 400 | 0.05 | Large |  |  | X
21 | 80 | 400 | 0.05 | Large |  |  | X
22 | 80 | 500 | 0.05 | Large | X
23 | 80 | 50 | 0.05 | Large | X
24 | 80 | 800 | 0.05 | Large | X
25 | 80 | 250 | 0.05 | Large | X
26 | 100 | 400 | 0.05 | Large |  | X
27 | 100 | 400 | 0.05 | Large |  | X | X
28 | 120 | 400 | 0.05 | Large |  | X
29 | 120 | 400 | 0.05 | Large |  | X | X
30 | 140 | 400 | 0.05 | Large |  | X
31 | 50 | 100 | 0.1 | Small |  | X
32 | 70 | 100 | 0.1 | Small |  | X
33 | 90 | 100 | 0.1 | Small |  | X |  |  | X
34 | 90 | 100 | 0.1 | Small |  |  |  |  | X
35 | 90 | 100 | 0.1 | Small |  |  |  |  | X
36 | 90 | 100 | 0.1 | Small |  |  |  |  | X
37 | 90 | 100 | 0.1 | Small |  |  |  |  | X
38 | 80 | 100 | 0.1 | Small | X
39 | 80 | 110 | 0.1 | Small | X
40 | 80 | 250 | 0.1 | Small | X
41 | 80 | 50 | 0.1 | Small | X
42 | 80 | 160 | 0.1 | Small | X
43 | 110 | 100 | 0.1 | Small |  | X
44 | 130 | 100 | 0.1 | Small |  | X
45 | 80 | 200 | 0.01 | Large |  | X
46 | 80 | 200 | 0.02 | Large |  | X
47 | 80 | 200 | 0.04 | Large |  | X
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
Note on H941: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on U950: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on AA950: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on AF950: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Correct Time
Unprotect
Protect

## DataPage
Sheet | RAD
SurveyID
SID
Illumination | TBD
Beam Align | NA
Rad/film center table | 0
Rad/film center wall | 0
LFS | 0
SFS | 0
Table SID
Wall SID
Collimation table
PBL table | NA | NA
| NA | NA
Collimation wall |  |  |  |  | 14 | 18
PBL wall | NA | NA
| NA | NA
LFS Output
SFS Output
Gen Test | 60 | 400 | 0.05 | 0 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 60 | 400 | 0.05 | 0 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 3 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 3.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 400 | 0.05 | 0 | 70-120 | 2.5 | Aluminum | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
| 80 | 500 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 50 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 800 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 250 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 400 | 0.05 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| 120 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 120 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| 140 | 400 | 0.05 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 50 | 100 | 0.1 | 0 | 35-60 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 70 | 100 | 0.1 | 0 | 50-85 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 90 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
| 80 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 110 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 250 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 50 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | High | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 160 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 110 | 100 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 130 | 100 | 0.1 | 0 | 100-155 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Small | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.01 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.02 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.04 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 1 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.1 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.25 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 200 | 0.4 | 0 | 70-120 | 0 | none | 100 | cm | 1 | Low | Internal | 0 | 0 | Large | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
HVL | 0 | 0
| #DIV/0! | #DIV/0!
| 0 | 0
| #DIV/0! | TBD
| #DIV/0! | #DIV/0!
| #DIV/0! | #DIV/0!
| 0 | 0
| #DIV/0! | #DIV/0!
| 0 | 0
| 0 | 0