## Sheet1
1					Revision 1.0-20140806					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					CT System Compliance Inspection					CT System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1
8					System Information					Previous Date:					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Accession Number:					Survey ID:				Accession Number:					Survey ID:
14					Room Number:					Room Number:
15			CT Scanner					Site Number:
16				Manufacturer:					Manufacture Date:					CT Scanner					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18					Model:				Serial Number:					Accesssion Number:
19			X-Ray Tube					Focal Spot Sizes (mm)
20				Tube Designation/Use:					Large:					X-Ray Tube					Focal Spot Sizes (mm)					X-Ray Generator
21				Manufacture Date:					Small:					Tube Designation/Use:				Large:					Manufacturer:
22			Insert					Micro:					Manufacture Date:				Small:					Model:
23				Manufacturer:					Insert					Micro:					Max kVp:
24				Model:					Filtration					Manufacturer:					Manufacture Date:
25				Serial Number:					Inherent:					Model:				Filtration					Serial Number:
26			Housing					Added:					Serial Number:				Inherent:					Max mA:
27				Manufacturer:					Housing					Added:					X-Ray Tube 1
28				Model:					Manufacturer:					Tube Designation/Use:
29				Serial Number:					Model:					Manufacture Date:
30					Serial Number:					Manufacturer:
31					Model:
32		Rule Number			Labels, Notices, Postings					Compliance					Serial Number:
33		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Inspection Results					Manufacturer:
34		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Model:
35				Pregnancy Warning sign is posted					TBD					Serial Number:
36		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Labels, Notices, Postings					Large:
37				Operator manuals are available.					TBD				DHEC Registration sticker is present, clearly visible and legible					Small:
38				Monthly radiation monitoring reports are posted.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Micro:
39					Pregnancy Warning sign is posted					Inherent:
40					Radiation Safety					Radiation warning label posted on the generator control panel					Added:
41				Radiation warning light at room entrance is functional with selected tube					TBD				Operator manuals are available.					Measurement Protocols
42				Lead aprons available.					TBD				Monthly radiation monitoring reports are posted.					Adult Abdomen
43				A properly designed and installed apron rack is present.					TBD					kVp
44				Documentation of annual protective apparel integrity inspection is available.					TBD					Radiation Safety					mA
45				Patient restraint devices available.					TBD				Radiation warning light at room entrance is functional with selected tube					mAs
46					Lead aprons available.					Eff mAs
47				Reconstructed images are free of artifacts					TBD				A properly designed and installed apron rack is present.					Time/rotation (s)
48				Acquisition station display performance is adequate					TBD				Documentation of annual protective apparel integrity inspection is available.					Scan FOV (cm)
49				Daily QC performed by technologist					TBD				Patient restraint devices available.					Display FOV (cm)
50				Interlock systems function to prevent scanning while door is open					TBD					Reconstruction Algorithm
51		DHEC RHB 4.11.1.1		Tomographic plane is indicated by a light source					TBD				Reconstructed images are free of artifacts					Axial/Helical
52		DHEC RHB 4.11.1.3.1		Visual indication of x-ray production is present and functional					TBD				Acquisition station display performance is adequate					Z-Axis collimation
53		DHEC RHB 4.11.1.6.1		Positioning with the alignment light is accurate to within 2mm					TBD				Daily QC performed by technologist					# of channels
54		DHEC RHB 4.11.1.6.3		Table travel is accurate to within 1mm					TBD				Interlock systems function to prevent scanning while door is open					Table Increment/speed
55		DHEC RHB 4.11.2.5.1		Technologist has a clear view of the patient during the scan					TBD				Tomographic plane is indicated by a light source					Pitch
56				Patient communication system is operational					TBD				Visual indication of x-ray production is present and functional					Recon scan width (mm)
57					Positioning with the alignment light is accurate to within 2mm					Recon scan interval (mm)
58					Radiation Dose Measurements					Table travel is accurate to within 1mm					Dose reduction techniques
59				Routine head protocol CTDI is below ACR reference level of 75 mGy					NA				Technologist has a clear view of the patient during the scan					Indiated CTDIvol (mGy)
60				Routine body protocol CTDI is below ACR reference level of 25 mGy					NA				Patient communication system is operational					Adult Chest
61				Pediatric body protocol CTDI is below ACR reference level of 15 mGy					NA					kVp
62				Pediatric head protocol CTDI is below ACR reference level of 35 mGy					NA					Radiation Dose Measurements					mA
63					3	Routine head protocol CTDI is below ACR reference level of 75 mGy					mAs
64					3	Routine body protocol CTDI is below ACR reference level of 25 mGy					Eff mAs
65		Date:					Inspector:	Eugene Mah		3	Pediatric body protocol CTDI is below ACR reference level of 15 mGy					Time/rotation (s)
66		Room Number					Survey ID:			3	Pediatric head protocol CTDI is below ACR reference level of 35 mGy					Scan FOV (cm)
1					Medical University of South Carolina					Display FOV (cm)
2					Inspection Results					CT System Compliance Inspection					Reconstruction Algorithm
3					Measurement Protocols					Measurement Protocols					Axial/Helical
4					Adult			Pediatric					Adult			Pediatric					Z-Axis collimation
5					Abdomen	Chest	Head	Abdomen	Head					Abdomen	Chest	Head	Abdomen	Head					# of channels
6				kVp					kVp					Table Increment/speed
7				mA					mA					Pitch
8				mAs					mAs					Recon scan width (mm)
9				Eff mAs					Eff mAs					Recon scan interval (mm)
10				Time/rotation (s)					Time/rotation (s)					Dose reduction techniques
11				Scan FOV (cm)					Scan FOV (cm)					Indiated CTDIvol (mGy)
12				Display FOV (cm)					Display FOV (cm)					Adult Head
13				Reconstruction Algorithm					Reconstruction Algorithm					kVp
14				Axial/Helical					Axial/Helical					mA
15				Z-Axis collimation					Z-Axis collimation					mAs
16				# of channels					# of channels					Eff mAs
17				Table Increment/speed					Table Increment/speed					Time/rotation (s)
18				Pitch					Pitch					Scan FOV (cm)
19				Recon scan width (mm)					Recon scan width (mm)					Display FOV (cm)
20				Recon scan interval (mm)					Recon scan interval (mm)					Reconstruction Algorithm
21				Dose reduction techniques					Dose reduction techniques					Axial/Helical
22				Indiated CTDIvol (mGy)					Indiated CTDIvol (mGy)					Z-Axis collimation
23					Adult			Pediatric					# of channels
24		CTDI Measurements					Abdomen	Chest	Head	Abdomen	Head					Table Increment/speed
25					Adult			Pediatric					kVp					Pitch
26					Abdomen	Chest	Head	Abdomen	Head	Criteria:					mA					Recon scan width (mm)
27				Indiated CTDIvol (mGy)					CTDIvol is less than ACR Pass/Fail dose					mAs					Recon scan interval (mm)
28			CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					Reference	Pass/Fail			Eff mAs					Dose reduction techniques
29					(mGy)	(mGy)			Time/rotation (s)					Indiated CTDIvol (mGy)
30					Adult Body:	25	30			Scan FOV (cm)					Ped Abdomen
31				Average CTDIvol:					Adult Head:	75	80			Display FOV (cm)					kVp
32				Phantom size indicated (cm)					Pediatric Body:	15	20			Reconstruction Algorithm					mA
33				Phantom size used (cm)					Pediatric Head:	35	40			Axial/Helical					mAs
34				Acceptable:					Z-Axis collimation					Eff mAs
35					# of channels					Time/rotation (s)
36		Beam Profile					Table Increment/speed					Scan FOV (cm)
37					Pitch					Display FOV (cm)
38					Recon scan width (mm)					Reconstruction Algorithm
39				Channels (N)					Recon scan interval (mm)					Axial/Helical
40				Detector width (mm)					Dose reduction techniques					Z-Axis collimation
41				Table velocity (mm/s)					Indiated CTDIvol (mGy)					# of channels
42				Beam width (mm)					Table Increment/speed
43				Measured beam width (mm)					Pitch
44					Inspection Results					Recon scan width (mm)
45					CTDI Measurements					Recon scan interval (mm)
46					Adult			Pediatric			Criteria:				Dose reduction techniques
47				Acceptable:					Abdomen	Chest	Head	Abdomen	Head		CTDIvol is less than ACR Pass/Fail dose				Indiated CTDIvol (mGy)
48					Criteria:	Beam profile is within 3mm or 30% of specified value					Indicated CTDIvol (mGy)					Reference	Pass/Fail		Ped Head
49					CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					(mGy)	(mGy)		kVp
50					Adult Body:	25	30		mA
51					Adult Head:	75	80		mAs
52					Average CTDIvol:					Pediatric Body:	15	20		Eff mAs
53					Phantom size indicated (cm)					Pediatric Head:	35	40		Time/rotation (s)
54					Phantom size used (cm)					Scan FOV (cm)
55					Previous year:					Display FOV (cm)
56					Acceptable:					Reconstruction Algorithm
57					Axial/Helical
58					Beam Profile					Z-Axis collimation
59					Channels (N)					# of channels
60					Detector width (mm)					Criteria:	Beam profile is within 3mm or 30% of specified value					Table Increment/speed
61					Table velocity (mm/s)					Pitch
62					Beam width (mm)					Recon scan width (mm)
63					Measured beam width (mm)					Recon scan interval (mm)
64					Dose reduction techniques
65		Date:					Inspector:	Eugene Mah					Indiated CTDIvol (mGy)
66		Room Number					Survey ID:				Average (mm):					CTDI Measurements
1					Medical University of South Carolina			Acceptable:					Adult Abdomen
2					Inspection Results					CT System Compliance Inspection					Adult Chest
3		Alignment Accuracy					Adult Head
4			Landmark position:		mm		Slice thickness:		mm					Inspection Results					Ped Abdomen
5					Alignment Accuracy					Ped Head
6					Foot					Head					Landmark position:		mm	Slice thickness		mm
7				Location of alignment slice:		mm			Location of alignment slice:		mm
8				All markers visible					All markers visible					Foot				Head					CT Beam Profile – Siemens
9				Central wire visible					Table movement accuracy					Location of alignment slice:		mm		Location of alignment slice:		mm					NxT	Definition Flash			NxT	Definition AS			NxT	Sensation 64		NxT	Sensation 16		NxT	Emotion 16
10					All markers visible				All markers visible					Z-sharp	Nominal (mm)	Tolerance)		Z-sharp	Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)		Nominal (mm)	Tolerance)
11		CT Number Accuracy					Central wire visible				Table movement accuracy					2x1	Off	3.5	1.0	2x1	Off	4.7	1.0	30x0.6	22.5	1.5	16x1.5	30.5	1.0	4x0.6	4.9	1.0
12			kVp	80	100	110	120	130	140	150					Enter 1 for YES, 2 for NO				Enter 1 for YES, 2 for NO					1x5	Off	6.0	1.0	1x5	Off	6.6	1.0	24x1.2	33.5	1.5	16x0.75	18.5	1.0	12x0.6	12.6	1.0
13			mAs					CT Number Accuracy					10x0.6	On	10.5	1.5	10x0.6	On	11.2	1.0	12x1.2	19.5	1.5	12x1.5	24.5	1.0	16x0.6	14.5	1.0
14			Polyethylene					kVp	80	100	110	120	130	140	150		Acceptable HU			8x0.6	On	11.5	1.5	8x0.6	Off	8.5	1.0	1x10	11.5	1.5	12x0.75	15.5	1.0	2x5	12.6	1.0
15			Water					mAs					Low	High		1x10	Off	11.1	1.5	1x10	Off	11.6	1.0	1x5	6.5	1.5	2x5	12.5	1.0	12x1.2	18.5	1.0
16			Water Std Dev					Polyethylene					Polyethylene	-107	-87		20x0.6	On	18.8	1.5	8x0.6	On	11.5	1.0	2x1	4.0	1.0	2x1	3.5	1.0	2x8	18.5	1.0
17			Acrylic					Water					Water	-7	7		12x1.2	Off	19.4	1.5	6x1.2	Off	11.8	1.0	6x0.6	7.0	1.0	2x0.6	3.0	1.0	16x1.2	24.2	1.0
18			Bone					Water Std Dev					Acrylic	110	130		32x0.6	Off	24.4	2.0	12x1.2	Off	19.8	1.0
19			Air					Acrylic					Bone	850	970		32x0.6	On	27.4	2.0	60x0.6	Off	41.0	1.5
20					Acceptable:					Bone					Air	-1005	-970		32x1.2	Off	44.0	2.0	64x0.6	On	45.5	1.5
21					Critiera:	Water HU is 0 +/- 5 at all kVp					Air					64x0.6	On	45.5	2.0	32x1.2	Off	43.5	1.5
22					CT number for inserts using the adult abdomen protocol is within the acceptable range					Acceptable:		Enter 1 for YES, 2 for NO					30x0.6	Off	23.5	1.0
23					Critiera:	Water HU is 0 +/- 5 at all kVp					32x0.6	On	28.0	1.0
24		CT Slice Thickness Accuracy					CT number for inserts using the adult abdomen protocol is within the acceptable range					20x0.6	Off	18.0	1.0
25				Slice thickness (mm):	1	3	5	7	10					16x1.2	Off	25.0	1.0
26				Top (mm):					CT Slice Thickness Accuracy					20x0.6	On	20.0	1.0
27				Bottom (mm):					Slice thickness (mm):
28				Average (mm):					Top (mm):
29				Water HU:					Bottom (mm):
30				Water HU Std Dev:					Average (mm):
31				Acceptable:					Water HU:
32					Criteria:	Slice thickness is within 1.5mm of the set thickness					Water HU Std Dev:
33					CT number for water is 0 +/- 5					Acceptable:
34					Criteria:	Slice thickness is within 1.5mm of the set thickness
35		Low Contrast Performance					CT number for water is 0 +/- 5
36				Adult		Pediatric
37				Abdomen	Head	Abdomen	Head					Low Contrast Performance
38			Inside					6mm rods visible:					Adult		Pediatric
39			Adjacent					Acceptable:					Abdomen	Head	Abdomen	Head			Enter 1 for YES, 2 for NO, 3 for Not Applicable
40			Adjacent StDev					Inside					6mm rods visible:		Acceptable:
41			Contrast					Adjacent					Criteria:	All four 6mm rods are visible
42			CNR					Adjacent StDev					CNR > 1.0 for adult abdomen and head protocols
43					Criteria:	All four 6mm rods are visible					Contrast					CNR > 0.4 for pediatric body protocol
44					CNR > 1.0 for adult abdomen and head protocols					CNR					CNR > 0.7 for pediatric head protocol
45					CNR > 0.4 for pediatric body protocol
46					CNR > 0.7 for pediatric head protocol					Uniformity
47					Mean	St Dev				Acceptable
48		Uniformity					Center:				Center-Top:
49				Mean	St Dev				Acceptable					Top:				Center-Bottom:
50			Center:				Center-Top:					Bottom:				Center-Right:
51			Top:				Center-Bottom:					Right:				Center-Left:
52			Bottom:				Center-Right:					Left:				Criteria:	CT number of all ROIs is within +/- 5 HU of the central region
53			Right:				Center-Left:
54			Left:					High Contrast Resolution
55					Criteria:	CT number of all ROIs is within +/- 5 HU of the central region					lp/cm	Acceptable
56					Abdomen				Critiera:	5 lp/cm for abdomen
57		High Contrast Resolution					Hi-res Chest					6 lp/cm for high resolution
58				lp/cm	Acceptable
59			Abdomen				Critiera:	5 lp/cm for abdomen					Acquisition Console Monitor
60			Hi-res Chest					6 lp/cm for high resolution
61					SMPTE 5%/95% visible
62		Acquisition Console Monitor					Display free of artifacts
63			SMPTE 5%/95% visible			Display free of artifacts			Display uniformity					Display uniformity
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID: