## Summary
Medical Physicist CT Survey Report | [5 empty] | Pass
Fail
This report summarizes the results of tests performed in accordance with the American College of
Radiology CT QC Manual. | [5 empty] | Pass
Fail
Facility Name |  |  | Unit ID |  |  | NA
Address 1 |  |  | Manufacturer
Address 2 |  |  | Model
City, State, ZIP |  |  | Serial Number
Date of Manufacture
CTAP # (if applicable)
|  |  | Survey Date | 0
Medical Physicist | Eugene Mah |  | Report Date
Signature
Medical Physicist Tests |  | Pass/Fail |  | Technologist QC Evaluation | Pass/Fail/NA
| Review of CT Protocols |  |  | Water CT Number and SD (Daily)
| Scout Prescription Accuracy |  |  | Artifact Evaluation (Daily)
| Alignment Light Accuracy |  |  | Wet Laser QC (Weekly)
| Table Travel Accuracy |  |  | Visual Checklist (Monthly)
| Radiation Beam Width |  |  | Dry Laser QC (Monthly)
| Low-Contrast Performance |  |  | Acquisition Display QC (Monthly)
Spatial Resolution
CT Number Accuracy
Artifact Evaluation
Dosimetry
CT Number Uniformity
Acquisition Display Calibration
Comments

## Sheet1
1 | [11 empty] | Revision 1.3-202010601 | [10 empty] | Print Area |  |  |  | Measurements
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | First,CTDI,ACRPhantom,Comments
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina | [9 empty] | Set text | Light (cd/m²)
4 | [22 empty] | All: | First,CTDI,ACRPhantom,Comments |  |  | L_amb
5 | [5 empty] | CT System Compliance Inspection | [10 empty] | CT System Compliance Inspection | [9 empty] | LN01
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year | LN01
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  | Change | First,CTDI,ACRPhantom,Comments | LN18
8 | [5 empty] | System Information | [5 empty] | Previous Date: | [10 empty] | Date |  |  |  | LN18
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah | UNL10
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility: |  |  |  | UNL10
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department: |  |  |  | UNL10
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division: |  |  |  | UNL10
13 |  |  | Survey ID: |  |  |  |  | Accession Number: | [5 empty] | Survey ID: |  |  |  | Accession Number: |  |  |  |  | Survey ID: |  |  |  | UNL10
14 | [13 empty] | Room Number: | [8 empty] | Room Number: |  |  |  | UNL10
15 |  | CT Scanner | [20 empty] | Site Number: |  |  |  | UNL10
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | CT Scanner | [9 empty] | Authorized Use: |  |  |  | UNL10
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation: |  |  |  | UNL10
18 | [7 empty] | SoftwareVersion: | [5 empty] | Model: |  |  |  | Serial Number: |  |  |  |  | Accesssion Number: |  |  |  | UNL10
19 |  | X-Ray Tube | [15 empty] | Software Version: | [8 empty] | UNL10
20 |  |  | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube | [9 empty] | X-Ray Generator |  |  |  | UNL10
21 |  |  | Manufacture Date: |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Manufacturer: |  |  |  | UNL10
22 |  | Insert | [5 empty] | Small: | [5 empty] | Manufacture Date: |  |  |  | Large: |  |  |  |  | Model: |  |  |  | UNL10
23 |  |  | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Manufacture Date: |  |  |  | UNL10
24 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Micro: |  |  |  |  | Serial Number: |  |  |  | UNL10
25 |  |  | Serial Number: |  |  |  |  | Filtration | [5 empty] | Model: | [8 empty] | Software Version: |  |  |  | UNL10
26 |  | Housing | [5 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | Filtration | [8 empty] | UNL10
27 |  |  | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1 |  |  |  | UNL80
28 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Added: |  |  |  |  | Tube Designation/Use: |  |  |  | UNL80
29 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | Manufacture Date: |  |  |  | UNL80
30 | [13 empty] | Serial Number: | [8 empty] | Manufacturer: |  |  |  | UNL80
31 | [22 empty] | Model: |  |  |  | UNL80
32 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance | [12 empty] | Serial Number: |  |  |  | UNL80
33 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Inspection Results | [5 empty] | Manufacturer: |  |  |  | UNL80
34 | DHEC RHB 11.2.3 |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Model: |  |  |  | UNL80
35 |  |  | Pregnancy Warning sign is posted | [6 empty] | TBD | [12 empty] | Serial Number: |  |  |  | UNL80
36 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD | [6 empty] | Labels, Notices, Postings | [5 empty] | Large: |  |  |  | UNL80
37 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | DHEC Registration sticker is present, clearly visible and legible | [9 empty] | Small: |  |  |  | UNL80
38 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [9 empty] | Micro: |  |  |  | UNL80
39 | [12 empty] | Pregnancy Warning sign is posted | [9 empty] | Inherent: |  |  |  | UNL80
40 |  |  |  | Radiation Safety | [8 empty] | Radiation warning label posted on the generator control panel | [9 empty] | Added: |  |  |  | UNL80
41 |  |  | Radiation warning light at room entrance is functional with selected tube | [6 empty] | TBD |  |  | Operator manuals are available. | [9 empty] | Measurement Protocols |  |  |  | UNL80
42 | DHEC RHB 4.2.12.7 |  | Lead aprons available. | [6 empty] | TBD |  |  | Monthly radiation monitoring reports are posted. | [9 empty] | Adult Abdomen |  |  |  | UNL80
43 |  |  | A properly designed and installed apron rack is present. | [6 empty] | TBD | [6 empty] | Radiation Safety | [5 empty] | Protocol name |  |  |  | UNL80
44 | DHEC RHB 4.2.8 |  | Documentation of annual protective apparel integrity inspection is available. | [6 empty] | TBD |  |  | Radiation warning light at room entrance is functional with selected tube | [9 empty] | kVp |  |  |  | UNL80
45 |  |  | Patient restraint devices available. | [6 empty] | TBD |  |  | Lead aprons available. | [9 empty] | mA |  |  |  | UNL80
46 |  |  | Unit installed as shown on shielding plan | [6 empty] | TBD |  |  | A properly designed and installed apron rack is present. | [9 empty] | mAs
47 |  |  | Reconstructed images are free of artifacts | [6 empty] | TBD |  |  | Documentation of annual protective apparel integrity inspection is available. | [9 empty] | Eff mAs
48 |  |  | Acquisition station display performance is adequate | [6 empty] | TBD |  |  | Patient restraint devices available. | [9 empty] | Time/rotation (s)
49 |  |  | Daily QC performed by technologist | [6 empty] | TBD |  |  | Unit installed as shown on shielding plan | [9 empty] | Scan FOV (cm)
50 | [22 empty] | Display FOV (cm)
51 | DHEC RHB 4.11.1.1 |  | Tomographic plane is indicated by a light source | [6 empty] | TBD |  |  | Reconstructed images are free of artifacts | [9 empty] | Reconstruction Algorithm
52 | DHEC RHB 4.11.1.3.1 |  | Visual indication of x-ray production is present and functional | [6 empty] | TBD |  |  | Acquisition station display performance is adequate | [9 empty] | Axial/Helical
53 | DHEC RHB 4.11.1.5.1 |  | Positioning with the alignment light is accurate to within 5 mm | [6 empty] | TBD |  |  | Daily QC performed by technologist | [9 empty] | Z-Axis collimation
54 | DHEC RHB 4.11.1.5.3 |  | Table travel is accurate to within 1mm | [6 empty] | TBD | [12 empty] | # of channels
55 | DHEC RHB 4.11.2.4.1 |  | Technologist has a clear view of the patient during the scan | [6 empty] | TBD |  |  | Tomographic plane is indicated by a light source | [9 empty] | Table Increment/speed
56 | DHEC RHB 4.11.2.2 |  | Patient communication system is operational | [6 empty] | TBD |  |  | Visual indication of x-ray production is present and functional | [9 empty] | Pitch
57 | DHEC RHB 4.11.2.1 |  | Exposure switch is secured at a safe location | [6 empty] | TBD |  |  | Positioning with the alignment light is accurate to within 5 mm | [9 empty] | Recon scan width (mm)
58 | [12 empty] | Table travel is accurate to within 1mm | [9 empty] | Recon scan interval (mm)
59 |  |  |  | Radiation Dose Measurements | [8 empty] | Technologist has a clear view of the patient during the scan | [9 empty] | Dose reduction techniques
60 |  |  | Routine head protocol CTDI is below ACR Pass/Fail level of 80 mGy | [6 empty] | NA |  |  | Patient communication system is operational | [9 empty] | Indiated CTDIvol (mGy)
61 |  |  | Routine body protocol CTDI is below ACR Pass/Fail level of 30 mGy | [6 empty] | NA |  |  | Exposure switch is secured at a safe location | [9 empty] | Adult Chest
62 |  |  | Pediatric body protocol CTDI is below ACR Pass/fail level of 20 mGy | [6 empty] | NA | [6 empty] | Radiation Dose Measurements | [5 empty] | Protocol name
63 |  |  | Pediatric head protocol CTDI is below ACR Pass/Fail level of 40 mGy | [6 empty] | NA |  | 3 | Routine head protocol CTDI is below ACR Pass/Fail level of 80 mGy | [9 empty] | kVp
64 | [11 empty] | 3 | Routine body protocol CTDI is below ACR Pass/Fail level of 30 mGy | [9 empty] | mA
65 | Date: | [8 empty] | Inspector: | Eugene Mah | 3 | Pediatric body protocol CTDI is below ACR Pass/fail level of 20 mGy | [9 empty] | mAs
66 | Room Number | [8 empty] | Survey ID: |  | 3 | Pediatric head protocol CTDI is below ACR Pass/Fail level of 40 mGy | [9 empty] | Eff mAs
1 | [10 empty] | Medical University of South Carolina | [11 empty] | Time/rotation (s)
2 | [5 empty] | Inspection Results |  |  |  |  | CT System Compliance Inspection | [11 empty] | Scan FOV (cm)
3 | [5 empty] | Measurement Protocols | [10 empty] | Measurement Protocols | [5 empty] | Display FOV (cm)
4 |  |  |  | Adult |  |  | Pediatric | [6 empty] | Adult |  |  | Pediatric | [5 empty] | Reconstruction Algorithm
5 |  |  |  | Abdomen | Chest | Head | Abdomen | Head | [5 empty] | Abdomen | Chest | Head | Abdomen | Head |  |  |  |  | Axial/Helical
6 |  |  | kVp | [9 empty] | Protocol name | [9 empty] | Z-Axis collimation
7 |  |  | mA | [9 empty] | kVp | [9 empty] | # of channels
8 |  |  | mAs | [9 empty] | mA | [9 empty] | Table Increment/speed
9 |  |  | Eff mAs | [9 empty] | mAs | [9 empty] | Pitch
10 |  |  | Time/rotation (s) | [9 empty] | Eff mAs | [9 empty] | Recon scan width (mm)
11 |  |  | Scan FOV (cm) | [9 empty] | Time/rotation (s) | [9 empty] | Recon scan interval (mm)
12 |  |  | Display FOV (cm) | [9 empty] | Scan FOV (cm) | [9 empty] | Dose reduction techniques
13 |  |  | Reconstruction Algorithm | [9 empty] | Display FOV (cm) | [9 empty] | Indiated CTDIvol (mGy)
14 |  |  | Axial/Helical | [9 empty] | Reconstruction Algorithm | [9 empty] | Adult Head
15 |  |  | Z-Axis collimation | [9 empty] | Axial/Helical | [9 empty] | Protocol name
16 |  |  | # of channels | [9 empty] | Z-Axis collimation | [9 empty] | kVp
17 |  |  | Table Increment/speed | [9 empty] | # of channels | [9 empty] | mA
18 |  |  | Pitch | [9 empty] | Table Increment/speed | [9 empty] | mAs
19 |  |  | Recon scan width (mm) | [9 empty] | Pitch | [9 empty] | Eff mAs
20 |  |  | Recon scan interval (mm) | [9 empty] | Recon scan width (mm) | [9 empty] | Time/rotation (s)
21 |  |  | Dose reduction techniques | [9 empty] | Recon scan interval (mm) | [9 empty] | Scan FOV (cm)
22 |  |  | Indiated CTDIvol (mGy) | [9 empty] | Dose reduction techniques | [9 empty] | Display FOV (cm)
23 | [12 empty] | Indiated CTDIvol (mGy) | [9 empty] | Reconstruction Algorithm
24 | CTDI Measurements | [12 empty] | Adult |  |  | Pediatric | [5 empty] | Axial/Helical
25 | Meter: | Piranha |  |  | Calibration Date: | [8 empty] | Abdomen | Chest | Head | Abdomen | Head |  |  |  |  | Z-Axis collimation
26 | Serial #: |  |  |  | Calibration Due: | [7 empty] | Protocol name | [9 empty] | # of channels
27 | [12 empty] | kVp | [9 empty] | Table Increment/speed
28 |  |  |  | Adult |  |  | Pediatric | [5 empty] | mA | [9 empty] | Pitch
29 |  |  |  | Abdomen | Chest | Head | Abdomen | Head | Criteria: |  |  |  | mAs | [9 empty] | Recon scan width (mm)
30 |  |  | Indicated CTDIvol (mGy): | [5 empty] | CTDIvol is less than ACR Pass/Fail dose |  |  |  | Eff mAs | [9 empty] | Recon scan interval (mm)
31 |  | CTDIvol Measured (mGy) | [7 empty] | Reference | Pass/Fail |  | Time/rotation (s) | [9 empty] | Dose reduction techniques
32 | [9 empty] | (mGy) | (mGy) |  | Scan FOV (cm) | [9 empty] | Indiated CTDIvol (mGy)
33 | [8 empty] | Adult Abdomen | 25 | 30 |  | Display FOV (cm) | [9 empty] | Ped Abdomen
34 |  |  | Average CTDIvol: | [5 empty] | Adult Head: | 75 | 80 |  | Reconstruction Algorithm | [9 empty] | Protocol name
35 |  |  | Phantom size indicated (cm): | [5 empty] | Ped Abd (16cm) | 15 | 20 |  | Axial/Helical | [9 empty] | kVp
36 |  |  | Phantom size used (cm): | [5 empty] | Ped Abd (32cm) | 7.5 | 10 |  | Z-Axis collimation | [9 empty] | mA
37 |  |  | Variation from previous year: | [5 empty] | Pediatric Head: | 35 | 40 |  | # of channels | [9 empty] | mAs
38 |  |  | Variation from indicated CTDI: | [9 empty] | Table Increment/speed | [9 empty] | Eff mAs
39 |  |  | Acceptable: | [9 empty] | Pitch | [9 empty] | Time/rotation (s)
40 |  |  | CTDIvol/mAs: | [9 empty] | Recon scan width (mm) | [9 empty] | Scan FOV (cm)
41 |  |  | Criteria: | CTDI is within 5% of previous year. | [8 empty] | Recon scan interval (mm) | [9 empty] | Display FOV (cm)
42 |  |  |  | Measured CTDI is within 20% of indicated CTDI | [8 empty] | Dose reduction techniques | [9 empty] | Reconstruction Algorithm
43 | [12 empty] | Indiated CTDIvol (mGy) | [9 empty] | Axial/Helical
44 | Beam Profile | [21 empty] | Z-Axis collimation
45 |  | Channels (N) | [20 empty] | # of channels
46 |  | Detector width T (mm) | [20 empty] | Table Increment/speed
47 |  | NxT Beam width (mm) | [20 empty] | Pitch
48 |  | Nominal beam width (mm) | [14 empty] | Inspection Results | [5 empty] | Recon scan width (mm)
49 |  | Average beam width (mm) | [9 empty] | CTDI Measurements | [10 empty] | Recon scan interval (mm)
50 | [11 empty] | Meter: | Piranha |  | Calibration Date: | [7 empty] | Dose reduction techniques
51 |  | Acceptable: | [9 empty] | Serial #: |  |  | Calibration Due: | [7 empty] | Indiated CTDIvol (mGy)
52 |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value | [18 empty] | Ped Head
53 | [13 empty] | Adult |  |  | Pediatric |  |  | Criteria: |  |  | Protocol name
54 | [13 empty] | Abdomen | Chest | Head | Abdomen | Head |  | CTDIvol is less than ACR Pass/Fail dose |  |  | kVp
55 | Artifact Evaluation | [11 empty] | Indicated CTDIvol (mGy) | [7 empty] | Reference | Pass/Fail | mA
56 |  | Water phantom is free from rings, streaks, lines, cupping artifacts | [9 empty] | CTDIvol Measured (mGy) | [8 empty] | (mGy) | (mGy) | mAs
57 | [19 empty] | Adult Abdomen | 25 | 30 | Eff mAs
58 | Acquisition Console Monitor | [18 empty] | Adult Head: | 75 | 80 | Time/rotation (s)
59 |  | SMPTE 5%/95% visible |  | #DIV/0! | Lmin (cd/m^2) |  |  | Critiera: | Lmin <= 1.2 cd/m^2, Lmax >= 90 cd/m^2 |  |  |  | Average CTDIvol: | [6 empty] | Ped Abd (16cm) | 15 | 20 | Scan FOV (cm)
60 |  | Display free of artifacts |  | #DIV/0! | Lmax (cd/m^2) |  |  |  | Luminance deviation < 30% |  |  |  | Phantom size indicated (cm) | [6 empty] | Ped Abd (32cm) | 7.5 | 10 | Display FOV (cm)
61 |  | Display uniformity |  | #DIV/0! | % Non-uniformity |  |  |  | Luminance deviation from median < 30% |  |  |  | Phantom size used (cm) | [6 empty] | Pediatric Head: | 35 | 40 | Reconstruction Algorithm
62 |  |  |  | #DIV/0! | Max luminance deviation from median | [7 empty] | Previous year: | [9 empty] | Axial/Helical
63 | [12 empty] | Variation from prev year: | [9 empty] | Z-Axis collimation
64 | [12 empty] | Variation from indicated: | [9 empty] | # of channels
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Acceptable: | [9 empty] | Table Increment/speed
66 | Room Number | [8 empty] | Survey ID: |  |  | CTDIvol/100 mAs: | [9 empty] | Pitch
1 | [10 empty] | Medical University of South Carolina |  | Prev CTDIvol/100 mAs: | [9 empty] | Recon scan width (mm)
2 | [5 empty] | Inspection Results |  |  |  |  | CT System Compliance Inspection |  | Criteria: | CTDI is within 5% of previous year. | [8 empty] | Recon scan interval (mm)
3 | Alignment Accuracy | [12 empty] | Measured CTDI is within 20% of indicated CTDI | [8 empty] | Dose reduction techniques
4 |  | Landmark position: | 0 | mm |  | Slice thickness: | 1 | mm |  |  |  | Beam Profile | [10 empty] | Indiated CTDIvol (mGy)
5 | [12 empty] | Channels (N) | [9 empty] | CTDI Measurements
6 |  |  |  | Foot | Head |  |  |  | Scout prescription accuracy: |  |  |  | Detector width T (mm) |  |  |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value |  |  |  | Adult Abdomen
7 |  |  | Location of alignment slice: |  |  | mm |  |  | Central wire visible: |  |  |  | NxT Beam width (mm) | [9 empty] | Adult Chest
8 |  |  | All markers visible | [5 empty] | Table movement accuracy: |  |  |  | Nominal beam width (mm) | [9 empty] | Adult Head
9 | [8 empty] | Alignment light accuracy: |  |  |  | Measured beam width (mm) | [9 empty] | Ped Abdomen
10 | [22 empty] | Ped Head
11 | CT Number Accuracy | [21 empty] | CTDIvol/100 mAs
12 |  | kVp | 80 | 100 | 110 | 120 | 130 | 140 | 150 |  |  |  | Average (mm): | [9 empty] | Adult Abdomen
13 |  | mAs | [10 empty] | Acceptable: | [9 empty] | Adult Chest
14 |  | Polyethylene | [20 empty] | Adult Head
15 |  | Water | [9 empty] | Artifact Evaluation | [10 empty] | Ped Abdomen
16 |  | Water Std Dev | [10 empty] | Water phantom is free from rings, streaks, lines, cupping artifacts | [9 empty] | Ped Head
17 |  | Acrylic
18 |  | Bone | [20 empty] | Comment Page 1
19 |  | Air | [14 empty] | Inspection Results
20 |  |  |  |  | Acceptable: | [6 empty] | Alignment Accuracy
21 |  |  | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp | [8 empty] | Landmark position: | 0 | mm | Slice thickness | 1 | mm
22 |  |  |  | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
23 | [13 empty] | Foot | Head |  |  |  | Enter 1 for YES, 2 for NO
24 | CT Slice Thickness Accuracy | [11 empty] | Location of alignment slice: |  |  | mm |  | Scout prescription accuracy:
25 |  |  | Slice thickness (mm): | [9 empty] | All markers visible: |  |  |  |  | Central wire visible:
26 |  |  | Top (mm): | [14 empty] | Table movement accuracy:
27 |  |  | Bottom (mm): | [14 empty] | Alignment light accuracy:
28 |  |  | Average (mm): | [8 empty] | CT Number Accuracy
29 |  |  | Water HU: | [8 empty] | kVp | 80 | 100 | 110 | 120 | 130 | 140 | 150 |  | Acceptable HU
30 |  |  | Water HU Std Dev: | [8 empty] | mAs | [8 empty] | Low | High
31 |  |  | Acceptable: | [8 empty] | Polyethylene | [7 empty] | Polyethylene | -107 | -84
32 |  |  |  | Criteria: | Slice thickness is within 1.5mm of the set thickness | [6 empty] | Water | [7 empty] | Water | -7 | 7
33 |  |  |  |  | CT number for water is 0 +/- 5 | [6 empty] | Water Std Dev | [7 empty] | Acrylic | 110 | 135
34 | [11 empty] | Acrylic | [7 empty] | Bone | 850 | 970
35 | Low Contrast Performance | [10 empty] | Bone | [7 empty] | Air | -1005 | -970
36 |  |  | Adult |  | Pediatric | [6 empty] | Air
37 |  |  | Abdomen | Head | Abdomen | Head | [8 empty] | Acceptable: |  | Enter 1 for YES, 2 for NO
38 |  | Inside | [5 empty] | 6mm rods visible: |  |  |  |  | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
39 |  | Adjacent | [5 empty] | Acceptable: | [5 empty] | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
40 |  | Adjacent StDev
41 |  | Contrast | [9 empty] | CT Slice Thickness Accuracy
42 |  | CNR | [10 empty] | Slice thickness (mm):
43 |  |  |  | Criteria: | All four 6mm rods are visible | [7 empty] | Top (mm):
44 |  |  |  |  | CNR > 1.0 for adult abdomen and head protocols | [7 empty] | Bottom (mm):
45 |  |  |  |  | CNR > 0.4 for pediatric abdomen protocol | [7 empty] | Average (mm):
46 |  |  |  |  | CNR > 1.0 for pediatric head protocol | [7 empty] | Water HU:
47 | Uniformity | [11 empty] | Water HU Std Dev:
48 |  | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable | [6 empty] | Acceptable:
49 | Center: | [12 empty] | Criteria: | Slice thickness is within 1.5mm of the set thickness
50 | Top: | [13 empty] | CT number for water is 0 +/- 5
51 | Bottom:
52 | Right: | [10 empty] | Low Contrast Performance
53 | Left: | [11 empty] | Adult |  | Pediatric
54 |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region | [9 empty] | Abdomen | Head | Abdomen | Head |  |  | Enter 1 for YES, 2 for NO, 3 for Not Applicable
55 |  | Uniformity image is free from artifacts | [9 empty] | Inside | [5 empty] | 6mm rods visible: |  | Acceptable:
56 | [11 empty] | Adjacent | [5 empty] | Criteria: | All four 6mm rods are visible
57 | High Contrast Resolution | [10 empty] | Adj StDev | [6 empty] | CNR > 1.0 for adult abdomen and head protocols
58 |  |  | lp/cm | Acceptable | [7 empty] | Contrast | [6 empty] | CNR > 0.4 for pediatric abdomen protocol
59 |  | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen |  |  |  |  | CNR | [6 empty] | CNR > 1.0 for pediatric head protocol
60 |  | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution
61 | [11 empty] | Uniformity
62 | [12 empty] | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable
63 | [11 empty] | Center:
64 | [11 empty] | Top:
65 | Date: | [8 empty] | Inspector: | Eugene Mah | Bottom:
66 | Room Number | [8 empty] | Survey ID: |  | Right:
1 | [10 empty] | Medical University of South Carolina | Left:
2 | [5 empty] | Comments Page 1 |  |  |  |  | CT System Compliance Inspection |  | Uniformity image is free from artifacts |  |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region
3 | [11 empty] | High Contrast Resolution
4 | [12 empty] | lp/cm | Acceptable
5 | [11 empty] | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen
6 | [11 empty] | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution
7
8 | [11 empty] | Acquisition Console Monitor |  |  |  |  | Luminance | Dev from Median
9 | [15 empty] | Display uniformity | (cd/m^2)
10 | [12 empty] | SMPTE 5%/95% visible |  |  | Region 1: | 0 | #DIV/0!
11 | [12 empty] | Display free of artifacts |  |  | Region 2: | 0 | #DIV/0!
12 | [12 empty] | Display uniformity |  |  | Region 3: | 0 | #DIV/0!
13 | [11 empty] | #DIV/0! | Lmin (cd/m^2) |  |  | Region 4: | 0 | #DIV/0!
14 | [11 empty] | #DIV/0! | Lmax (cd/m^2) |  |  | Region 5: | 0 | #DIV/0!
15 | [11 empty] | Critiera: | Lmin <= 1.2 cd/m^2, Lmax >= 90 cd/m^2 |  |  | Region 6: | 0 | #DIV/0!
16 | [12 empty] | Luminance deviation is < 30% |  |  | Region 7: | 0 | #DIV/0!
17 | [12 empty] | Luminance deviation from median is < 30% |  |  | Region 8: | 0 | #DIV/0!
18 | [15 empty] | Region 9: | 0 | #DIV/0!
19 | [15 empty] | Average: | 0 | #DIV/0! | Max luminance deviation from median
20 | [15 empty] | SD: | 0
21 | [15 empty] | % Non-uniformity: | #DIV/0!
22
23
24
25 | [16 empty] | Comments Page 1
26 | [15 empty] | Previous Year Comments
27 | [12 empty] | Comments:
28 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
29 | [12 empty] | Additional Comments:
30 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
31 | [12 empty] | Additional Comments:
32 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
33 | [12 empty] | Additional Comments:
34 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
35 | [12 empty] | Additional Comments:
36 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
37 | [12 empty] | Additional Comments:
38 | [12 empty] | Do not exceed 390 chars |  | 0
39 | [12 empty] | Additional Comments:
40 | [12 empty] | Do not exceed 390 chars |  | 0
41 | [12 empty] | Additional Comments:
42 | [12 empty] | Do not exceed 390 chars |  | 0
43 | [12 empty] | Additional Comments:
44 | [12 empty] | Do not exceed 390 chars |  | 0
45 | [12 empty] | Additional Comments:
46 | [12 empty] | Do not exceed 390 chars |  | 0
47 | [12 empty] | Additional Comments:
48 | [12 empty] | Do not exceed 390 chars |  | 0
49 | [12 empty] | Additional Comments:
50 | [12 empty] | Do not exceed 390 chars |  | 0
51 | [12 empty] | Additional Comments:
52 | [12 empty] | Do not exceed 390 chars |  | 0
53 | [12 empty] | Additional Comments:
54 | [12 empty] | Do not exceed 390 chars |  | 0
55 | [12 empty] | Additional Comments:
56 | [12 empty] | Do not exceed 390 chars |  | 0
57 | [12 empty] | Additional Comments:
58 | [12 empty] | Do not exceed 390 chars |  | 0
59 | [12 empty] | Additional Comments:
60 | [12 empty] | Do not exceed 390 chars |  | 0
61 | [12 empty] | Additional Comments:
62 | [12 empty] | Do not exceed 390 chars |  | 0
63 | [12 empty] | Additional Comments:
64 | [12 empty] | Do not exceed 390 chars |  | 0
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:

## Sheet2
Siemens CT Beam Profiles http://128.23.56.214/MPwiki/index.php?title=Siemens_CT_Beam_Profiles
CT Beam Profile – Siemens
NxT | Definition Flash
| Z-sharp | Nominal (mm) | Tolerance)
2x1 | Off | 3.5 | 1
1x5 | Off | 6 | 1
10x0.6 | On | 10.5 | 1.5
8x0.6 | On | 11.5 | 1.5
1x10 | Off | 11.1 | 1.5
20x0.6 | On | 18.8 | 1.5
12x1.2 | Off | 19.4 | 1.5
32x0.6 | Off | 24.4 | 2
32x0.6 | On | 27.4 | 2
32x1.2 | Off | 44 | 2
64x0.6 | On | 45.5 | 2
NxT | Definition AS
| Z-sharp | Nominal (mm) | Tolerance)
2x1 | Off | 4.7 | 1
1x5 | Off | 6.6 | 1
10x0.6 | On | 11.2 | 1
8x0.6 | Off | 8.5 | 1
1x10 | Off | 11.6 | 1
8x0.6 | On | 11.5 | 1
6x1.2 | Off | 11.8 | 1
12x1.2 | Off | 19.8 | 1
60x0.6 | Off | 41 | 1.5
64x0.6 | On | 45.5 | 1.5
32x1.2 | Off | 43.5 | 1.5
30x0.6 | Off | 23.5 | 1
32x0.6 | On | 28 | 1
20x0.6 | Off | 18 | 1
16x1.2 | Off | 25 | 1
20x0.6 | On | 20 | 1
NxT | Sensation 64
| Nominal (mm) | Tolerance)
30x0.6 | 22.5 | 1.5
24x1.2 | 33.5 | 1.5
12x1.2 | 19.5 | 1.5
1x10 | 11.5 | 1.5
1x5 | 6.5 | 1.5
2x1 | 4 | 1
6x0.6 | 7 | 1
NxT | Sensation 16
| Nominal (mm) | Tolerance)
16x1.5 | 30.5 | 1
16x0.75 | 18.5 | 1
12x1.5 | 24.5 | 1
12x0.75 | 15.5 | 1
2x5 | 12.5 | 1
2x1 | 3.5 | 1
2x0.6 | 3 | 1
NxT | Emotion 16
| Nominal (mm) | Tolerance)
4x0.6 | 4.9 | 1
12x0.6 | 12.6 | 1
16x0.6 | 14.5 | 1
2x5 | 12.6 | 1
12x1.2 | 18.5 | 1
2x8 | 18.5 | 1
16x1.2 | 24.2 | 1
NxT | Somatom Force
| Nominal (mm) | Tolerance)
2x1 | 3 | 1
6x0.5 | 7 | 1
1x5 | 7.5 | 1
1x10 | 12 | 1.5
12x1.2 | 24 | 2
32x0.6 (UHR) | 22.5 | 2
32x0.6 (sUHR) | 24.5 | 2
32x0.6 (StdHR) | 28.5 | 2
64x0.6 | 47.5 | 2
96x0.6 | 66 | 2
48x1.2 | 67 | 2
NxT | Somatom Confidence
| Nominal (mm) | Tolerance)
20x0.6 | 18.8 | 1.5
2x1 | 3.5 | 1
12x1.2 | 19.8 | 1.5
16x1.2 | 25 | 2