## Start here
Philips MicroDose L50
Quality Control Procedures and Installation Test Record
Doc ID: | 1022707 | Date: | 2012-11-06
Revision: | A
Purpose of this document:
This Excel record template is to be used together with the Quality Control Procedures - Philips MicroDose Mammography SI, Model L50, Software 9.0 (This document is based on version DOC-LANN-8V7GB6-3.0).
Instructions
Go trough this sheet by sheet, which you select at the tabs bellow.
In most test cases, instruction are taken from the QCP manual, which must be available.
Beige cells | should be filled out by the tester.
Blue cells | are automatically calculated.
Red cells | use output data from Xi dose meter (this is optional).
Yellow tabs | use Xi data that is automatically written to the protocol.
Red tabs | are auxiliary tabs that are helpful but not necessary to use.
Blue tabs | contain data used for calculations. Do not edit!
Note: The system shall be calibrated according to routines in the Quality Control Procedures document (QCP).
Information about the customer/site and the system.
| Test date: |  | Software versions:
| Test responsible: |  | Bundle:
| Customer/site: |  | Service Tool
| Installation date: |  | AWS
| Serial number: |  | IPK
| Dose setting: |  | CCS
Used test equipment
Equipment type |  | Mark/model |  | ID of used equipment
Dose meter
Dose detector
CDMAM phantom
Line pair phantom
DQC phantom
Force measurement tool
Outcome of test
State the overall outcome of the test here (PASS/FAIL):
Comments

## Spatial Resolution
Spatial resolution (5.2)
Material:
CIRS 016A phantom
Instructions:
Follow instructions in the QCP manual.
Test data:
| Exposure parameters: |  | kV
mAs
Test result:
|  | Low collimator |  | High collimator
Visible line pair density
| In scan direction: |  | lp/mm |  | lp/mm
| Perpendicular to scan direction: |  | lp/mm |  | lp/mm
Expected result:
In both directions for both of the collimators, the visible line density should be at least 6 lp/mm.
Outcome of control:
---
Comments, additional information:

## Contrast-to-noise ratio
Contrast-to-noise ratio (5.3)
CNR Correction Factors
Material: | [7 empty] | Correction factor if kV used
Daily quality phantom (Part No. 1010909). | [6 empty] | kV | 26 | 29 | 32 | 35 | 38
[7 empty] | 26 | 1 | 0.94 | 0.89 | 0.84 | 0.8
Instructions: | [6 empty] | 29 | 1.07 | 1 | 0.94 | 0.9 | 0.85
Follow instructions in the QCP manual. | [6 empty] | 32 | 1.13 | 1.06 | 1 | 0.95 | 0.9
[7 empty] | 35 | 1.19 | 1.12 | 1.05 | 1 | 0.95
Test data: | [6 empty] | 38 | 1.26 | 1.18 | 1.11 | 1.06 | 1
Dose setting:
|  | Used Exposure parameters: |  |  | kV
mAs
mGy (AGD)
|  | Conversion factor: | ---
|  |  |  | Image 1 | Image 2 | Image 3
|  | Mean (m): | light circle
| Using ROI 1A | Standard deviation (s):
CNR:
|  | SNR: |  | #DIV/0! | #DIV/0! | #DIV/0!
|  | Contrast: |  | #DIV/0! | #DIV/0! | #DIV/0!
|  | Mean (m): | dark circle
| Using ROI 1B | Standard deviation (s):
CNR:
|  | SNR: |  | #DIV/0! | #DIV/0! | #DIV/0!
|  | Contrast: |  | #DIV/0! | #DIV/0! | #DIV/0!
Sum of absolute CNR:
|  | Converted sum of CNR: |  | --- | --- | ---
Test result:
|  | Measured mean CNR: |  | ---
|  | Deviation from ref. CNR: |  | ---
Expected result:
The measured CNR must be within ±10% from reference CNR
Reference CNR:
Outcome of Control:
---
Comments, additional information:
| kV | Ref CNR
26
29
32
35
38

## Visual Check
Visual Check (5.4)
Material:
None
Instructions:
Follow instructions in the QCP manual.
Test result:
(Type x in appropriate cells or NA if not applicable)
| C-arm |  | PASS | FAIL
(1) Face Shield
(2) All compression paddles
(3, 4) All five keys on all four keypads
(5) Working light
(6) Angulation Indicator on display
(7) Motion of C-arm
(8) Laterality buttons (Sin/L and Dx/R)
(7) C-arm
(10) Collimator movement
(10) Collimators (low and high)
Stand
Stand
(11) Both Emergency stop buttons
(12) Reset button
(6) Compressed breast thickness indicator on the display
(6) Compression force Indicator on the display
(13) Both sets of foot pedals
AW table
(14) Emergency stop button
(15) Exposure buttons on AW-table working
(16) Exposure foot-pedal working
Expected result:
There must be no fail.
Outcome of control:
Write Pass/Fail here:
Comments, additional information

## Thickness indication
Thickness indication (6.1)
Material:
• Compressible material to simulate a breast, such as Styrofoam. (Thickness must be varied during the test.) • Vernier calliper
Instructions:
Follow instructions in the QCP manual.
Test data:
| Thickness interval | Force interval | Comp.-force (daN) on stand | Indicated thickness on display (mm) | Measured thickness (mm) | Deviation (mm) | Absolute deviation (mm)
|  | 5–8 daN |  |  |  | --- | ---
| 20–30mm | 11-14 daN |  |  |  | --- | ---
|  | 17-20 daN |  |  |  | --- | ---
|  | 5–8 daN |  |  |  | --- | ---
| 40–60mm | 11-14 daN |  |  |  | --- | ---
|  | 17-20 daN |  |  |  | --- | ---
|  | 5–8 daN |  |  |  | --- | ---
| 70–90mm | 11-14 daN |  |  |  | --- | ---
|  | 17-20 daN |  |  |  | --- | ---
Test result:
|  |  |  | Average deviation (mm): | #DIV/0!
|  |  |  | Average abs. deviation (mm): | #DIV/0!
|  |  |  | Maximal abs. deviation (mm): | ---
Expected result:
The deviation between indicated and measured thickness shall always be less than or equal to 3 mm.
Outcome of control:
---
Comments, additional information

## Compression force
Compression force (6.2)
Material:
Analog scale or compression force measurment device
Instructions:
Follow instructions in the QCP manual.
Test data:
Part A
| Comp.-force (daN) on stand | Comp. -force measured | Absolute deviation (N)
---
---
---
---
---
---
Maintained Compression
| Compression force at start |  |  |  | daN
| Compression force after 1 min |  |  |  | daN
Part B
| Maximal compression force |  |  |  | daN
| Compression force at disengage |  |  |  | daN
Test result:
|  |  | Average force deviation (N): | ---
|  |  | Maximal force deviation (N): | ---
|  |  | Compression force change after 1 min (N): | ---
|  |  | Maximal compression force (N): | ---
Expected result:
The deviation between indicated and measured force shall always be less than 20 N.
During a minute, the compression force may not fluctuate more than 5 N
It shall not be possible to use a force larger than 200 N
Outcome of control:
---
Comments, additional information

## X-ray tube output
X-ray tube output (6.3)
Note: If Unfors Xi dose meter is used for this test, data can be automatically written to this protocol (Xi Data sheet). In that case Air kerma reproducibility, and Half value layer are automatically filled out as well.
The use of Xi is optional.
Material:
Dose meter calibrated for the correct beam quality.
Instructions:
General test instructions
- Use the AWS with manual exposure parameters for all measurements.
- Compression paddle must be present for all measurements
- Compression paddle shall always be positioned at least 9 cm above the patient support.
- X-ray tube must be warm before measurements (If not warm make 3 exposures, using 35kV max mAs).
NOTE: When making many exposures in a row; wait 1 min between successive scans.
Instructions for Unfors Xi
| 1 | Press on
| 2 | Verify values:
|  | Trig delay | 0 ms
|  | kVp delay | 1000 ms
|  | calc. delay | 0.5 s
| 3 | Press select when "Mammography" is displayed
| 4 | Press select when "W/Al" is displayed
| 5 | Start XiView
| 6 | Select this excel Workbook from the dropdownlist (A) in Figure 3 and then select the Xi Data Worksheet from droplist (B), see Figure 3 and Figure 4.
| 7 | Verify that the choosen column and row is B and 2 respectively when you start, see (C) in Figure 3.
| 8 | Press start
Test procedure and test data:
| Subtest ID | Action by tester |  |  |  |  | Measured and calculated data
Change to low collimator
| 1 | Open the AEC status window in the ST Note the reported ESAK (mGy/mAs) in the protocol.
ESAK [mGy/mAs]
26 kV
29 kV
32 kV
35 kV
38 kV
| 2 | Position the Xi detector 6 cm from the chest wall and laterally centered using the positioning jigg. (The height of the detector above the object table when using the jigg is 14 mm)
Height (mm):
[7 empty] | Height correction factor: |  |  |  | ---
| 3 | Use 26 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 4 | Use 29 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 5 | Use 35 kV and measure the air kerma threetimes with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 6 | Use 38 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 7 | Use 32 kV and measure the air kerma 6 times using the lowest available mAs. Then measure the air kerma once using an intermediate mAs-value. Then measure the air kerma 6 times with the highest available mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
Change to high collimator
| 8 | Open the AEC status window in the ST Note the reported ESAK (mGy/mAs) in the protocol.
ESAK [mGy/mAs]
26 kV
29 kV
32 kV
35 kV
38 kV
| 9 | Use 26 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 10 | Use 29 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 11 | Use 32 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 12 | Use 35 kV and measure the air kerma threetimes with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 13 | Use 38 kV and measure the air kerma three times with different mAs.
[8 empty] | mAs values | Used mAs | Measured air kerma (mGy) | Corrected air kerma (mGy) | Tube output (mGy/mAs)
[8 empty] | Min |  | 0 | #VALUE! | #VALUE!
[8 empty] | Intermediate |  | 0 | #VALUE! | #VALUE!
[8 empty] | Max |  | 0 | #VALUE! | #VALUE!
| 14 | Make sure that the cells fetch the data from the correct cells in the Xi data sheet. Measurements not used (e.g. test or wrong exposures) should be removed from the Xi data sheet manually after all data is recorded. | [5 empty] | N/A
Test result:
|  |  | Low collimator |  |  |  |  | High collimator
|  |  | Measured tube output (mGy/mAs): |  | Dev. from value in ST: |  |  |  | Measured tube output (mGy/mAs): | Dev. from value in ST:
|  | 26 kV | #VALUE! |  | #VALUE! |  |  | 26 kV | #VALUE! | #VALUE!
|  | 29 kV | #VALUE! |  | #VALUE! |  |  | 29 kV | #VALUE! | #VALUE!
|  | 32 kV | #VALUE! |  | #VALUE! |  |  | 32 kV | #VALUE! | #VALUE!
|  | 35 kV | #VALUE! |  | #VALUE! |  |  | 35 kV | #VALUE! | #VALUE!
|  | 38 kV | #VALUE! |  | #VALUE! |  |  | 38 kV | #VALUE! | #VALUE!
[6 empty] | Max abs. deviation from ST value: | #VALUE!
Expected result:
The measured values shall not differ by more than ±15% from the values estimated by the system.
The tube output at 32 kV shall be at least 0.11 mGy/mAs.
Outcome of control:
---
Comments, additional information
Figure 3
A
B
C
D
Figure 1
Figure 2
Figure 4

## Air Kerma reproducibility
Air kerma reproducibility (6.4)
Note: This test uses data from the X-ray tube output test and is automatically filled out
Material:
Dose meter calibrated for the correct beam quality.
Instructions:
Verify that the data has been fetched appropriately from the source.
Test data:
| Lowest mAs |  | Measured air kerma (mGy) |  | Highest mAs |  | Measured air kerma (mGy)
|  | 1 | 0 |  |  | 1 | 0
|  | 2 | 0 |  |  | 2 | 0
|  | 3 | 0 |  |  | 3 | 0
|  | 4 | 0 |  |  | 4 | 0
|  | 5 | 0 |  |  | 5 | 0
|  | 6 | 0 |  |  | 6 | 0
Test result:
| Relative standard deviation: |  | #DIV/0! |  |  |  | #DIV/0!
Expected result:
Relative standard deviation shall be less than 5%
Outcome of control:
---
Comments, additional information

## AEC_Breast thickness & Exposure
AEC system: Breast thickness and exposure (6.5)
Material:
• 0.1 mm or 0.2 mm thick aluminium filter. • PMMA plates. Thickness should be possible to vary between 20 and 70 mm in increments of 10 mm. (Note: If absolute CNR values are to be measured, the Al sheet must be a calibrated 0.1 mm filter.
Instructions:
Follow instructions in QCP manual.
Test data and results:
Dose level:
Optional
Low collimator
| PMMA (mm) / Breast (mm) | Minimum CNR rel. 5 cm PMMA | Max dose (achievable) (mGy) | kV | mAs | AGD (mGy) | Measured CNR | Relative CNR | Relative CNR OK? | Dose rel. achievable level | Abs. target CNR values | Rel. dev. from target
| 20 / 20 | 130% | 0.6 |  |  |  |  | --- | ---- | --- |  | ---
| 30 / 31 | 110% | 1 |  |  |  |  | --- | ---- | --- |  | ---
| 40 / 45 | 100% | 1.6 |  |  |  |  | --- | ---- | --- |  | ---
| 50 / 60 | 100% | 2.6 |  |  |  |  | 100% |  | --- |  | ---
| 60 / 75 | 90% | 4 |  |  |  |  | --- | ---- | --- |  | ---
| 70 / 88 | 80% | 5.8 |  |  |  |  | --- | ---- | --- |  | ---
High collimator
| 50 / 60 | 100% | na |  |  |  |  | --- |  |  |  | ---
| 80 / 102 | 75% | na |  |  |  |  | --- | ---- |  |  | ---
| 85 / 109 | 70% | na |  |  |  |  | --- | ---- |  |  | ---
[9 empty] | Dose rel. achievable level (maximum): | ---
Expected result:
For each thickness, AGD shall be less than the achievable value and CNR shall be higher than the required min rel. 5 cm PMMA.
Outcome of control:
---
Comments, additional information

## AEC_Density compensation
AEC system - Density compensation (7.1)
Material:
PMMA plates. Thickness should vary between 30 and 50 mm in increments of 5 mm.
Instructions:
Follow instructions in QCP
Test data:
|  |  | PMMA thickness | Used mAs | Mean (background: m) | Std dev (background: s) | SNR | Rel SNR dev from mean
| Expected mAs: |  | 30 |  |  |  | --- | ---
|  |  | 35 |  |  |  | --- | ---
| Used kV: |  | 40 |  |  |  | --- | ---
|  |  | 45 |  |  |  | --- | ---
|  |  | 50 |  |  |  | --- | ---
Test result:
|  | Max abs relative SNR deviation: |  | ---
Expected result:
Max abs relative SNR deviation < 15%
Outcome of control:
---
Comments, additional information

## Tube voltage
Tube voltage (7.2)
Note: If Unfors Xi dose meter is used for this test, data can be automatically written to this protocol (Xi Data sheet).
The use of Unfirs Xi is optional.
Material:
kVp-meter calibrated for the correct beam quality.
Instructions:
General test instructions
- Use the AWS with manual exposure parameters for all measurements.
- Compression paddle must be present for all measurements
- Compression paddle shall always be positioned at least 9 cm above the patient support.
- X-ray tube must be warm before measurements (If not warm make 3 exposures, using 35kV max mAs).
NOTE: When making many exposures in a row; wait 1 min between successive scans.
Instructions for Unfors Xi
| 1 | Press on
| 2 | Verify values:
|  | Trig delay |  | 0 ms
|  | kVp delay |  | 1000 ms
|  | calc. delay |  | 0.5 s
| 3 | Press select when "Mammography" is displayed
| 4 | Press select when "W/Al" is displayed
| 5 | Start XiView
| 6 | Select this excel Workbook from the dropdownlist (A) in Figure 3 and then select the Xi Data Worksheet from droplist (B), see Figure 3 and Figure 4.
| 7 | Verify that the choosen column and row is B and 45 respectively when you start, see (C) in Figure 3.
| 8 | Press start
Test procedure and test data:
| Subtest ID | Action by tester |  |  |  | Measured and calculated data
| 1 | Position the Xi detector 6 cm from the chest wall and laterally centered using the positioning jigg. Place the Unfors Al filter on the compresson paddle. For the exposures below, ensure that Al filter covers the detector area of the Unfors Xi, see Figure 1 and Figure 2.
| 2 | Use 26 kV. Make three exposures at about 12 mAs.
0
[7 empty] | Measured kV: |  | 0
0
| 3 | Use 29 kV. Make three exposures at about 12 mAs.
0
[7 empty] | Measured kV: |  | 0
0
| 4 | Use 32 kV. Make six exposures at about 12 mAs.
0
0
[7 empty] | Measured kV: |  | 0
0
0
0
| 5 | Use 35 kV. Make three exposures at about 12 mAs.
0
[7 empty] | Measured kV: |  | 0
0
| 6 | Use 38 kV. Make three exposures at about 12 mAs.
0
[7 empty] | Measured kV: |  | 0
0
Test result:
|  |  |  | Accuracy |  | Reproducibility at 32 kV
|  |  |  | Abs. (kV) | Rel.
|  |  | 26 kV | 26 | 100% | Abs. | 0 | kV
|  |  | 29 kV | 29 | 100% | Rel. | #DIV/0!
|  |  | 32 kV | 32 | 100%
|  |  | 35 kV | 35 | 100%
|  |  | 38 kV | 38 | 100%
Expected result:
For all kV, accuracy better than 1 kV.
For all kV, reproducibility better than 0.5 kV.
Outcome of control:
---
Comments, additional information
Figure 4
Figure 2
Figure 3
A
B
C
D
Figure 1

## HVL low coll
Half value layer for low collimator (7.3a)
Material:
(A) Unfors Xi or
(B) Dose meter, lead diaphragm, aluminum sheets, it shall be possible to vary thickness in 0.05 mm steps between 0.25 and 0.7 mm
Instructions:
If Unfors Xi is used in previous tests, this test is automatiocally filled out.
If not, follow instructions in QCP manual
Test data:
| Is Unfors Xi used? |  |  | (write "y" or "n")
| First measurement after tube exchange? |  |  |  |  | (write "y" or "n")
Test data and result:
|  | mm Al | ESAK | Rel. ESAK
| 26 kV |  |  | --- | HVL: | --- | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 29 kV |  |  | --- | HVL: | --- | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 32 kV |  |  | --- | HVL: | --- | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 35 kV |  |  | --- | HVL: | --- | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 38 kV |  |  | --- | HVL: | --- | mm Al
---
---
Expected result, first measurement after tube exchange:
For all kV, measured HVL at tube exchange shall be within system specification, with an additional 0.02 mm measurement error included.
| kV | Spec low | Spec high | Measured | OK/not OK
| 26 | 0.31 | 0.41 | --- | not OK
| 29 | 0.36 | 0.46 | --- | not OK
| 32 | 0.41 | 0.51 | --- | not OK
| 35 | 0.46 | 0.56 | --- | not OK
| 38 | 0.51 | 0.61 | --- | not OK
Expected result:
For all kV, measured HVL shall be within ±0.03 mm from reference values.
|  | kV | Ref value | Measured | OK/not OK
|  | 26 |  | --- | #VALUE!
|  | 29 |  | --- | #VALUE!
|  | 32 |  | --- | #VALUE!
|  | 35 |  | --- | #VALUE!
|  | 38 |  | --- | #VALUE!
Outcome of control:
---
Comments, additional information

## HVL high coll
Half value layer for high collimator (7.3b)
Material:
(A) Unfors Xi or
(B) Dose meter, lead diaphragm, aluminum sheets, it shall be possible to vary thickness in 0.05 mm steps between 0.25 and 0.7 mm
Instructions:
If Unfors Xi is used in previous tests, this test is automatiocally filled out.
If not, follow instructions in QCP manual
Test data:
| Is Unfors Xi used? |  | y | (write "y" or "n")
| First measurement after tube exchange? |  |  |  |  | (write "y" or "n")
Test data and result:
|  | mm Al | ESAK | Rel. ESAK
| 26 kV |  |  | --- | HVL: | #DIV/0! | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 29 kV |  |  | --- | HVL: | #DIV/0! | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 32 kV |  |  | --- | HVL: | #DIV/0! | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 35 kV |  |  | --- | HVL: | #DIV/0! | mm Al
---
---
|  | mm Al | ESAK | Rel. ESAK
| 38 kV |  |  | --- | HVL: | #DIV/0! | mm Al
---
---
Expected result, first measurement after tube exchange:
For all kV, measured HVL at tube exchange shall be within system specification, with an additional 0.02 mm measurement error included.
| kV | Spec low | Spec high | Measured | OK/not OK
| 26 | 0.31 | 0.41 | --- | not OK
| 29 | 0.36 | 0.46 | --- | not OK
| 32 | 0.41 | 0.51 | --- | not OK
| 35 | 0.46 | 0.56 | --- | not OK
| 38 | 0.51 | 0.61 | --- | not OK
Expected result:
For all kV, measured HVL shall be within ±0.03 mm from reference values.
|  | kV | Ref value | Measured | OK/not OK
|  | 26 |  | --- | #VALUE!
|  | 29 |  | --- | #VALUE!
|  | 32 |  | --- | #VALUE!
|  | 35 |  | --- | #VALUE!
|  | 38 |  | --- | #VALUE!
Outcome of control:
---
Comments, additional information

## Contrast Detail Resolution
Contrast-detail resolution (7.4)
Material:
• Contrast detail phantom CDMAM 3.4 including a total of 4 cm PMMA base. • Diagnostic workstation with calibrated 5 mega pixel monitors located in a room with clinical reading conditions.
Instructions:
Follow instructrions in QCP manual.
Test data and result:
|  |  |  | Dose level: |  |  | kV: |  | AGD:
mAs:
| Diam. (mm) | Threshold thickness (mm) | Thickest disc not seen (mm) | [6 empty] | Pass/Fail
|  |  | Im 1 | Im 2 | Im 3 | Im 4 | Im 5 | Im 6 | Mean
| 1.6 | 0.074 | [6 empty] | --- | ---
| 1 | 0.091 | [6 empty] | --- | ---
| 0.63 | 0.123 | [6 empty] | --- | ---
| 0.4 | 0.189 | [6 empty] | --- | ---
| 0.25 | 0.352 | [6 empty] | --- | ---
| 0.1 | 1.68 | [6 empty] | --- | ---
Expected result:
For at least 5 of the 6 diameters (4/5 for C70 dose level), the observed lowest thickness must be lower than the thickness threshold.
Outcome of control:
---
Comments, additional information
Alternative evaluation using the Cdcom program
Test data and result: (fill in exposure parameters above)
|  | D (mm) | Tthres (mm) | D*Tthres
|  | 1 |  | --- |  |  | Mean D*Tthres: | ---
|  | 0.8 |  | --- |  |  | IQI: | ---
|  | 0.63 |  | ---
|  | 0.5 |  | ---
|  | 0.4 |  | ---
|  | 0.31 |  | ---
|  | 0.25 |  | ---
|  | 0.2 |  | ---
|  | 0.16 |  | ---
|  | 0.13 |  | ---
|  | 0.1 |  | ---
Expected result:
|  | For C70: | IQI (0.16-1.0 mm disks) shall be greater than 15
|  | For C100: | IQI (0.10-1.0 mm disks) shall be greater than 17.5
|  | For C120: | IQI (0.10-1.0 mm disks) shall be greater than 20
Outcome of control:
---

## Detector linearity
Detector linearity (7.5)
Material:
DQC phantom
Instructions:
See QCP for instructions
Test data:
| mAs | Used mAs | Background: m | Background: s | SNR2 | SNR2/mAs
| min |  |  |  | --- | ---
| ≈ 11 |  |  |  | --- | ---
| ≈ 14 |  |  |  | --- | ---
| ≈ 18 |  |  |  | --- | ---
| ≈ 22 |  |  |  | --- | ---
| ≈ 26 |  |  |  | --- | ---
| max |  |  |  | --- | ---
Test result:
|  |  | Rel std. dev. of SNR2/mAs: | ---
|  |  | SNR2 - mAs correlation, R2: | ---
Expected result:
The squared correlation coefficient between mAs and SNR2 shall be at least 0.99
Outcome of control:
---
Comments, additional information

## Image and x-ray field
Image field and X-ray field agreement (7.6 & 7.7)
[8 empty] | Difference between real and phantom |  | Difference between indicated and phantom
Material:
1011954 Image Field Phantom
X-ray fluorescent paper
Part A: Agreement between real and indicated image field
Instructions:
Place image field phantom on patient support. Check that the phantom is aligned with the four corners (see Fig. 1). Acquire an image using image type Phantom. Read the 8 values in the displayed image (indiecated by A to the right). Round off to nearest 0.5 mm. For the stairs, give the lowest value for which you can still see the edge. Measure how much the painted image field differs from the phantom (indicated by B). Use 0.5 mm precision.
Test data and result:
| Expected result | Actual result |  |  |  | Pass? | Comments
|  |  | Diff. between real image field and edge of phantom (+ means that real image field is outside) Read this in image. | Diff. between indicated image field and edge of phantom (+ means indicated field is larger) Read this from breast support. | Diff. between indicated and real image field. (+ means indicated field is larger) |  |  |  | + means that indicated field is outside phantom
+ means that real field is outside phantom
|  |  | Do NOT fill out here but to the right | [7 empty] | Phantom
| Between 0 and 5 mm | Rear edge [mm]: | #DIV/0! | #DIV/0! | --- | ---- |  | Difference
| Between 0 and 5 mm | Left edge [mm]: | #DIV/0! | #DIV/0! | --- | ---- |  | between
| Between 0 and 5 mm | Front edge [mm]: | #DIV/0! | #DIV/0! | --- | ---- |  | indicated
| Between 0 and 5 mm | Right edge [mm]: | #DIV/0! | #DIV/0! | --- | ---- |  | image
| ≤ 5 mm | Missad tissue at chest wall [mm]: | --- |  |  | ---- |  | field
Outcome of Part A:
|  |  |  | --- | [5 empty] | Stairs |  | Stairs
Comments, additional information
Part B: Agreement between real image field and radiation field
Instructions:
Place the pieces of x-ray fluorescent paper parallel to the side and rear edges of the real image field edge (just measured in Part A) and 5 mm outward from the edge (see Fig. 2). Put a small strip of the intensifying screen in the middle of the field. (The strip in the center serves as a comparison source during exposure.) Turn off room lights. Select manual parameters of 35 kV and the maximal mAs. Make an exposure and observe the fluorescent material during the scan. Re-position the intensifying screen at the chest wall edge (see Fig. 3). Make a manual exposure at 35 kVp and the maximum mAs and observe the intensifying screen during the scan
Figure 1
Test result:
Does the screen glow? (yes / no)
Left:
Right:
Rear:
Is there radiation visible beyond the chest wall edge of the patient support? (yes / no)
Outcome of Part B:
---
Comments, additional information
A

## Xi Data
Test case | kV | ESAK | ESAK rate | Exp. Time | HVL
Low collimator
Output
26 kV
29 kV
35 kV
38 kV
Output
Linearity
Reproducibility
low mAs, 32 kV
intermed. mAs, 32 kV
high mAs, 32 kV
High collimator
Output
26 kV
29 kV
32 kV
35 kV
38 kV
Test case | kV | ESAK | ESAK rate | Exp. Time | HVL
Tube voltage
accuracy
26 kV
29 kV
accuracy
reproducibility
32 kV
35 kV
38 kV

## AGD calc tool
AGD calculator for Philips MicroDose L30
All data from Addendum on Digital Mammography to Chapter 3 of the European Guidelines for Quality Assurance in Mammography Screening, version 1.0, Nov. 2003
Fill out below (yellow fields):
PMMA thickness (mm):
HVL (mm Al):
ESAK/mAs (at 45 mm):
mAs:
|  | Breast thickness (mm): | ---
|  | g-factor: | #N/A
|  | c-factor: | #N/A
|  | s-factor: | 1.05
|  | height factor: | ---
|  | AGD (mGy): | ---

## CDMAM tool
Tool for visual evaluation of CDMAM 3.4 images
Figure 1: Correct image rotation.
Fill out table below:
T (mm) | D=2.00 | D=1.60 | D=1.25 | D=1.00 | D=0.80 | D=0.63 | D=0.50 | D=0.40 | D=0.31 | D=0.25 | D=0.20 | D=0.16 | D=0.13 | D=0.10 | D=0.08 | D=0.06 | T (mm)
2 | [16 empty] | 2
1.42 | [16 empty] | 1.42
1 | [16 empty] | 1
0.71 | [16 empty] | 0.71
0.5 | [16 empty] | 0.5
0.36 | [16 empty] | 0.36
0.25 | [16 empty] | 0.25
0.2 | [16 empty] | 0.2
0.16 | [16 empty] | 0.16
0.13 | [16 empty] | 0.13
0.1 | [16 empty] | 0.1
0.08 | [16 empty] | 0.08
0.06 | [16 empty] | 0.06
0.05 | [16 empty] | 0.05
0.04 | [16 empty] | 0.04
0.03 | [16 empty] | 0.03
Results: |  | 1 = true, 0 = false
T (mm) | D=2.00 | D=1.60 | D=1.25 | D=1.00 | D=0.80 | D=0.63 | D=0.50 | D=0.40 | D=0.31 | D=0.25 | D=0.20 | D=0.16 | D=0.13 | D=0.10 | D=0.08 | D=0.06 | T (mm)
2 | [6 empty] | - | - | - | - | - | - | - | - | - |  | 2
1.42 | [5 empty] | - | - | - | - | - | - | - | - | - | - | - | 1.42
1 |  |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | 1
0.71 |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | - | 0.71
0.5 |  |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 0.5
0.36 |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 0.36
0.25 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 0.25
0.2 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 0.2
0.16 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 0.16
0.13 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - |  | 0.13
0.1 | - | - | - | - | - | - | - | - | - | - | - | - | - | - |  |  | 0.1
0.08 | - | - | - | - | - | - | - | - | - | - | - | - | - |  |  |  | 0.08
0.06 | - | - | - | - | - | - | - | - | - | - | - | [5 empty] | 0.06
0.05 | - | - | - | - | - | - | - | - | - | - | - | [5 empty] | 0.05
0.04 | - | - | - | - | - | - | - | - | - | - | [6 empty] | 0.04
0.03 |  | - | - | - | - | - | - | - | - | [7 empty] | 0.03
Purpose
This document should aid and speed up the evaluation of CDMAM3.4 images.
Instructions
Study the CDMAM image on the diagnostic workstation. Make sure the viewing conditions are appropriate (e.g. light). The image should be rotated as shown in Fig. 1. Be sure to set the grey scale optimally. Zooming is allowed.
For each diameter that shall be evaluated, start with the thickest disc. Note in the "Protocol" sheet in which corner the disc is situated; "U" for up, "D" for down, "L" for left and "R" for right.
Continue with the next thinner disc, as shown by the red arrow in the figure.
Stop when you cannot determine the disc position.
When all diameters are done, the resulting contrast detail curve can be seen in the "Result" sheet.
U
R
D
L

## g-grid
g-factors
PMMA thickness (mm)
|  | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80
| 0.3 | 0.378 | 0.3663 | 0.3546 | 0.3429 | 0.3312 | 0.3195 | 0.3078 | 0.2961 | 0.2844 | 0.2727 | 0.261 | 0.2532 | 0.2454 | 0.2376 | 0.2298 | 0.222 | 0.2142 | 0.2064 | 0.1986 | 0.1908 | 0.183 | 0.1774 | 0.1718 | 0.1662 | 0.1606 | 0.155 | 0.151 | 0.147 | 0.143 | 0.139 | 0.135 | 0.1321 | 0.1292 | 0.1263 | 0.1234 | 0.1205 | 0.1176 | 0.1147 | 0.1118 | 0.1089 | 0.106 | 0.104 | 0.102 | 0.1 | 0.098 | 0.096 | 0.094 | 0.092 | 0.09 | 0.088 | 0.086 | 0.0848 | 0.0836 | 0.0824 | 0.0812 | 0.08 | 0.0788 | 0.0776 | 0.0764 | 0.0752 | 0.074
| 0.31 | 0.3866 | 0.3747 | 0.3628 | 0.3509 | 0.339 | 0.3271 | 0.3152 | 0.3033 | 0.2914 | 0.2795 | 0.2676 | 0.25964 | 0.25168 | 0.24372 | 0.23576 | 0.2278 | 0.21984 | 0.21188 | 0.20392 | 0.19596 | 0.188 | 0.18228 | 0.17656 | 0.17084 | 0.16512 | 0.1594 | 0.15528 | 0.15116 | 0.14704 | 0.14292 | 0.1388 | 0.13582 | 0.13284 | 0.12986 | 0.12688 | 0.1239 | 0.12092 | 0.11794 | 0.11496 | 0.11198 | 0.109 | 0.10694 | 0.10488 | 0.10282 | 0.10076 | 0.0987 | 0.09664 | 0.09458 | 0.09252 | 0.09046 | 0.0884 | 0.08718 | 0.08596 | 0.08474 | 0.08352 | 0.0823 | 0.08108 | 0.07986 | 0.07864 | 0.07742 | 0.0762
| 0.32 | 0.3952 | 0.3831 | 0.371 | 0.3589 | 0.3468 | 0.3347 | 0.3226 | 0.3105 | 0.2984 | 0.2863 | 0.2742 | 0.26608 | 0.25796 | 0.24984 | 0.24172 | 0.2336 | 0.22548 | 0.21736 | 0.20924 | 0.20112 | 0.193 | 0.18716 | 0.18132 | 0.17548 | 0.16964 | 0.1638 | 0.15956 | 0.15532 | 0.15108 | 0.14684 | 0.1426 | 0.13954 | 0.13648 | 0.13342 | 0.13036 | 0.1273 | 0.12424 | 0.12118 | 0.11812 | 0.11506 | 0.112 | 0.10988 | 0.10776 | 0.10564 | 0.10352 | 0.1014 | 0.09928 | 0.09716 | 0.09504 | 0.09292 | 0.0908 | 0.08956 | 0.08832 | 0.08708 | 0.08584 | 0.0846 | 0.08336 | 0.08212 | 0.08088 | 0.07964 | 0.0784
| 0.33 | 0.4038 | 0.3915 | 0.3792 | 0.3669 | 0.3546 | 0.3423 | 0.33 | 0.3177 | 0.3054 | 0.2931 | 0.2808 | 0.27252 | 0.26424 | 0.25596 | 0.24768 | 0.2394 | 0.23112 | 0.22284 | 0.21456 | 0.20628 | 0.198 | 0.19204 | 0.18608 | 0.18012 | 0.17416 | 0.1682 | 0.16384 | 0.15948 | 0.15512 | 0.15076 | 0.1464 | 0.14326 | 0.14012 | 0.13698 | 0.13384 | 0.1307 | 0.12756 | 0.12442 | 0.12128 | 0.11814 | 0.115 | 0.11282 | 0.11064 | 0.10846 | 0.10628 | 0.1041 | 0.10192 | 0.09974 | 0.09756 | 0.09538 | 0.0932 | 0.09194 | 0.09068 | 0.08942 | 0.08816 | 0.0869 | 0.08564 | 0.08438 | 0.08312 | 0.08186 | 0.0806
| 0.34 | 0.4124 | 0.3999 | 0.3874 | 0.3749 | 0.3624 | 0.3499 | 0.3374 | 0.3249 | 0.3124 | 0.2999 | 0.2874 | 0.27896 | 0.27052 | 0.26208 | 0.25364 | 0.2452 | 0.23676 | 0.22832 | 0.21988 | 0.21144 | 0.203 | 0.19692 | 0.19084 | 0.18476 | 0.17868 | 0.1726 | 0.16812 | 0.16364 | 0.15916 | 0.15468 | 0.1502 | 0.14698 | 0.14376 | 0.14054 | 0.13732 | 0.1341 | 0.13088 | 0.12766 | 0.12444 | 0.12122 | 0.118 | 0.11576 | 0.11352 | 0.11128 | 0.10904 | 0.1068 | 0.10456 | 0.10232 | 0.10008 | 0.09784 | 0.0956 | 0.09432 | 0.09304 | 0.09176 | 0.09048 | 0.0892 | 0.08792 | 0.08664 | 0.08536 | 0.08408 | 0.0828
| 0.35 | 0.421 | 0.4083 | 0.3956 | 0.3829 | 0.3702 | 0.3575 | 0.3448 | 0.3321 | 0.3194 | 0.3067 | 0.294 | 0.2854 | 0.2768 | 0.2682 | 0.2596 | 0.251 | 0.2424 | 0.2338 | 0.2252 | 0.2166 | 0.208 | 0.2018 | 0.1956 | 0.1894 | 0.1832 | 0.177 | 0.1724 | 0.1678 | 0.1632 | 0.1586 | 0.154 | 0.1507 | 0.1474 | 0.1441 | 0.1408 | 0.1375 | 0.1342 | 0.1309 | 0.1276 | 0.1243 | 0.121 | 0.1187 | 0.1164 | 0.1141 | 0.1118 | 0.1095 | 0.1072 | 0.1049 | 0.1026 | 0.1003 | 0.098 | 0.0967 | 0.0954 | 0.0941 | 0.0928 | 0.0915 | 0.0902 | 0.0889 | 0.0876 | 0.0863 | 0.085
| 0.36 | 0.4288 | 0.41596 | 0.40312 | 0.39028 | 0.37744 | 0.3646 | 0.35176 | 0.33892 | 0.32608 | 0.31324 | 0.3004 | 0.29164 | 0.28288 | 0.27412 | 0.26536 | 0.2566 | 0.24784 | 0.23908 | 0.23032 | 0.22156 | 0.2128 | 0.20648 | 0.20016 | 0.19384 | 0.18752 | 0.1812 | 0.17648 | 0.17176 | 0.16704 | 0.16232 | 0.1576 | 0.15424 | 0.15088 | 0.14752 | 0.14416 | 0.1408 | 0.13744 | 0.13408 | 0.13072 | 0.12736 | 0.124 | 0.12166 | 0.11932 | 0.11698 | 0.11464 | 0.1123 | 0.10996 | 0.10762 | 0.10528 | 0.10294 | 0.1006 | 0.09926 | 0.09792 | 0.09658 | 0.09524 | 0.0939 | 0.09256 | 0.09122 | 0.08988 | 0.08854 | 0.0872
| 0.37 | 0.4366 | 0.42362 | 0.41064 | 0.39766 | 0.38468 | 0.3717 | 0.35872 | 0.34574 | 0.33276 | 0.31978 | 0.3068 | 0.29788 | 0.28896 | 0.28004 | 0.27112 | 0.2622 | 0.25328 | 0.24436 | 0.23544 | 0.22652 | 0.2176 | 0.21116 | 0.20472 | 0.19828 | 0.19184 | 0.1854 | 0.18056 | 0.17572 | 0.17088 | 0.16604 | 0.1612 | 0.15778 | 0.15436 | 0.15094 | 0.14752 | 0.1441 | 0.14068 | 0.13726 | 0.13384 | 0.13042 | 0.127 | 0.12462 | 0.12224 | 0.11986 | 0.11748 | 0.1151 | 0.11272 | 0.11034 | 0.10796 | 0.10558 | 0.1032 | 0.10182 | 0.10044 | 0.09906 | 0.09768 | 0.0963 | 0.09492 | 0.09354 | 0.09216 | 0.09078 | 0.0894
| 0.38 | 0.4444 | 0.43128 | 0.41816 | 0.40504 | 0.39192 | 0.3788 | 0.36568 | 0.35256 | 0.33944 | 0.32632 | 0.3132 | 0.30412 | 0.29504 | 0.28596 | 0.27688 | 0.2678 | 0.25872 | 0.24964 | 0.24056 | 0.23148 | 0.2224 | 0.21584 | 0.20928 | 0.20272 | 0.19616 | 0.1896 | 0.18464 | 0.17968 | 0.17472 | 0.16976 | 0.1648 | 0.16132 | 0.15784 | 0.15436 | 0.15088 | 0.1474 | 0.14392 | 0.14044 | 0.13696 | 0.13348 | 0.13 | 0.12758 | 0.12516 | 0.12274 | 0.12032 | 0.1179 | 0.11548 | 0.11306 | 0.11064 | 0.10822 | 0.1058 | 0.10438 | 0.10296 | 0.10154 | 0.10012 | 0.0987 | 0.09728 | 0.09586 | 0.09444 | 0.09302 | 0.0916
| 0.39 | 0.4522 | 0.43894 | 0.42568 | 0.41242 | 0.39916 | 0.3859 | 0.37264 | 0.35938 | 0.34612 | 0.33286 | 0.3196 | 0.31036 | 0.30112 | 0.29188 | 0.28264 | 0.2734 | 0.26416 | 0.25492 | 0.24568 | 0.23644 | 0.2272 | 0.22052 | 0.21384 | 0.20716 | 0.20048 | 0.1938 | 0.18872 | 0.18364 | 0.17856 | 0.17348 | 0.1684 | 0.16486 | 0.16132 | 0.15778 | 0.15424 | 0.1507 | 0.14716 | 0.14362 | 0.14008 | 0.13654 | 0.133 | 0.13054 | 0.12808 | 0.12562 | 0.12316 | 0.1207 | 0.11824 | 0.11578 | 0.11332 | 0.11086 | 0.1084 | 0.10694 | 0.10548 | 0.10402 | 0.10256 | 0.1011 | 0.09964 | 0.09818 | 0.09672 | 0.09526 | 0.0938
| 0.4 | 0.46 | 0.4466 | 0.4332 | 0.4198 | 0.4064 | 0.393 | 0.3796 | 0.3662 | 0.3528 | 0.3394 | 0.326 | 0.3166 | 0.3072 | 0.2978 | 0.2884 | 0.279 | 0.2696 | 0.2602 | 0.2508 | 0.2414 | 0.232 | 0.2252 | 0.2184 | 0.2116 | 0.2048 | 0.198 | 0.1928 | 0.1876 | 0.1824 | 0.1772 | 0.172 | 0.1684 | 0.1648 | 0.1612 | 0.1576 | 0.154 | 0.1504 | 0.1468 | 0.1432 | 0.1396 | 0.136 | 0.1335 | 0.131 | 0.1285 | 0.126 | 0.1235 | 0.121 | 0.1185 | 0.116 | 0.1135 | 0.111 | 0.1095 | 0.108 | 0.1065 | 0.105 | 0.1035 | 0.102 | 0.1005 | 0.099 | 0.0975 | 0.096
| 0.41 | 0.4672 | 0.4537 | 0.4402 | 0.4267 | 0.4132 | 0.3997 | 0.3862 | 0.3727 | 0.3592 | 0.3457 | 0.3322 | 0.3227 | 0.3132 | 0.3037 | 0.2942 | 0.2847 | 0.2752 | 0.2657 | 0.2562 | 0.2467 | 0.2372 | 0.23024 | 0.22328 | 0.21632 | 0.20936 | 0.2024 | 0.19712 | 0.19184 | 0.18656 | 0.18128 | 0.176 | 0.17232 | 0.16864 | 0.16496 | 0.16128 | 0.1576 | 0.15392 | 0.15024 | 0.14656 | 0.14288 | 0.1392 | 0.13662 | 0.13404 | 0.13146 | 0.12888 | 0.1263 | 0.12372 | 0.12114 | 0.11856 | 0.11598 | 0.1134 | 0.11186 | 0.11032 | 0.10878 | 0.10724 | 0.1057 | 0.10416 | 0.10262 | 0.10108 | 0.09954 | 0.098
HVL | 0.42 | 0.4744 | 0.4608 | 0.4472 | 0.4336 | 0.42 | 0.4064 | 0.3928 | 0.3792 | 0.3656 | 0.352 | 0.3384 | 0.3288 | 0.3192 | 0.3096 | 0.3 | 0.2904 | 0.2808 | 0.2712 | 0.2616 | 0.252 | 0.2424 | 0.23528 | 0.22816 | 0.22104 | 0.21392 | 0.2068 | 0.20144 | 0.19608 | 0.19072 | 0.18536 | 0.18 | 0.17624 | 0.17248 | 0.16872 | 0.16496 | 0.1612 | 0.15744 | 0.15368 | 0.14992 | 0.14616 | 0.1424 | 0.13974 | 0.13708 | 0.13442 | 0.13176 | 0.1291 | 0.12644 | 0.12378 | 0.12112 | 0.11846 | 0.1158 | 0.11422 | 0.11264 | 0.11106 | 0.10948 | 0.1079 | 0.10632 | 0.10474 | 0.10316 | 0.10158 | 0.1
(mm Al) | 0.43 | 0.4816 | 0.4679 | 0.4542 | 0.4405 | 0.4268 | 0.4131 | 0.3994 | 0.3857 | 0.372 | 0.3583 | 0.3446 | 0.3349 | 0.3252 | 0.3155 | 0.3058 | 0.2961 | 0.2864 | 0.2767 | 0.267 | 0.2573 | 0.2476 | 0.24032 | 0.23304 | 0.22576 | 0.21848 | 0.2112 | 0.20576 | 0.20032 | 0.19488 | 0.18944 | 0.184 | 0.18016 | 0.17632 | 0.17248 | 0.16864 | 0.1648 | 0.16096 | 0.15712 | 0.15328 | 0.14944 | 0.1456 | 0.14286 | 0.14012 | 0.13738 | 0.13464 | 0.1319 | 0.12916 | 0.12642 | 0.12368 | 0.12094 | 0.1182 | 0.11658 | 0.11496 | 0.11334 | 0.11172 | 0.1101 | 0.10848 | 0.10686 | 0.10524 | 0.10362 | 0.102
| 0.44 | 0.4888 | 0.475 | 0.4612 | 0.4474 | 0.4336 | 0.4198 | 0.406 | 0.3922 | 0.3784 | 0.3646 | 0.3508 | 0.341 | 0.3312 | 0.3214 | 0.3116 | 0.3018 | 0.292 | 0.2822 | 0.2724 | 0.2626 | 0.2528 | 0.24536 | 0.23792 | 0.23048 | 0.22304 | 0.2156 | 0.21008 | 0.20456 | 0.19904 | 0.19352 | 0.188 | 0.18408 | 0.18016 | 0.17624 | 0.17232 | 0.1684 | 0.16448 | 0.16056 | 0.15664 | 0.15272 | 0.1488 | 0.14598 | 0.14316 | 0.14034 | 0.13752 | 0.1347 | 0.13188 | 0.12906 | 0.12624 | 0.12342 | 0.1206 | 0.11894 | 0.11728 | 0.11562 | 0.11396 | 0.1123 | 0.11064 | 0.10898 | 0.10732 | 0.10566 | 0.104
| 0.45 | 0.496 | 0.4821 | 0.4682 | 0.4543 | 0.4404 | 0.4265 | 0.4126 | 0.3987 | 0.3848 | 0.3709 | 0.357 | 0.3471 | 0.3372 | 0.3273 | 0.3174 | 0.3075 | 0.2976 | 0.2877 | 0.2778 | 0.2679 | 0.258 | 0.2504 | 0.2428 | 0.2352 | 0.2276 | 0.22 | 0.2144 | 0.2088 | 0.2032 | 0.1976 | 0.192 | 0.188 | 0.184 | 0.18 | 0.176 | 0.172 | 0.168 | 0.164 | 0.16 | 0.156 | 0.152 | 0.1491 | 0.1462 | 0.1433 | 0.1404 | 0.1375 | 0.1346 | 0.1317 | 0.1288 | 0.1259 | 0.123 | 0.1213 | 0.1196 | 0.1179 | 0.1162 | 0.1145 | 0.1128 | 0.1111 | 0.1094 | 0.1077 | 0.106
| 0.46 | 0.5026 | 0.48866 | 0.47472 | 0.46078 | 0.44684 | 0.4329 | 0.41896 | 0.40502 | 0.39108 | 0.37714 | 0.3632 | 0.35322 | 0.34324 | 0.33326 | 0.32328 | 0.3133 | 0.30332 | 0.29334 | 0.28336 | 0.27338 | 0.2634 | 0.25572 | 0.24804 | 0.24036 | 0.23268 | 0.225 | 0.21928 | 0.21356 | 0.20784 | 0.20212 | 0.1964 | 0.19224 | 0.18808 | 0.18392 | 0.17976 | 0.1756 | 0.17144 | 0.16728 | 0.16312 | 0.15896 | 0.1548 | 0.15188 | 0.14896 | 0.14604 | 0.14312 | 0.1402 | 0.13728 | 0.13436 | 0.13144 | 0.12852 | 0.1256 | 0.12386 | 0.12212 | 0.12038 | 0.11864 | 0.1169 | 0.11516 | 0.11342 | 0.11168 | 0.10994 | 0.1082
| 0.47 | 0.5092 | 0.49522 | 0.48124 | 0.46726 | 0.45328 | 0.4393 | 0.42532 | 0.41134 | 0.39736 | 0.38338 | 0.3694 | 0.35934 | 0.34928 | 0.33922 | 0.32916 | 0.3191 | 0.30904 | 0.29898 | 0.28892 | 0.27886 | 0.2688 | 0.26104 | 0.25328 | 0.24552 | 0.23776 | 0.23 | 0.22416 | 0.21832 | 0.21248 | 0.20664 | 0.2008 | 0.19648 | 0.19216 | 0.18784 | 0.18352 | 0.1792 | 0.17488 | 0.17056 | 0.16624 | 0.16192 | 0.1576 | 0.15466 | 0.15172 | 0.14878 | 0.14584 | 0.1429 | 0.13996 | 0.13702 | 0.13408 | 0.13114 | 0.1282 | 0.12642 | 0.12464 | 0.12286 | 0.12108 | 0.1193 | 0.11752 | 0.11574 | 0.11396 | 0.11218 | 0.1104
| 0.48 | 0.5158 | 0.50178 | 0.48776 | 0.47374 | 0.45972 | 0.4457 | 0.43168 | 0.41766 | 0.40364 | 0.38962 | 0.3756 | 0.36546 | 0.35532 | 0.34518 | 0.33504 | 0.3249 | 0.31476 | 0.30462 | 0.29448 | 0.28434 | 0.2742 | 0.26636 | 0.25852 | 0.25068 | 0.24284 | 0.235 | 0.22904 | 0.22308 | 0.21712 | 0.21116 | 0.2052 | 0.20072 | 0.19624 | 0.19176 | 0.18728 | 0.1828 | 0.17832 | 0.17384 | 0.16936 | 0.16488 | 0.1604 | 0.15744 | 0.15448 | 0.15152 | 0.14856 | 0.1456 | 0.14264 | 0.13968 | 0.13672 | 0.13376 | 0.1308 | 0.12898 | 0.12716 | 0.12534 | 0.12352 | 0.1217 | 0.11988 | 0.11806 | 0.11624 | 0.11442 | 0.1126
| 0.49 | 0.5224 | 0.50834 | 0.49428 | 0.48022 | 0.46616 | 0.4521 | 0.43804 | 0.42398 | 0.40992 | 0.39586 | 0.3818 | 0.37158 | 0.36136 | 0.35114 | 0.34092 | 0.3307 | 0.32048 | 0.31026 | 0.30004 | 0.28982 | 0.2796 | 0.27168 | 0.26376 | 0.25584 | 0.24792 | 0.24 | 0.23392 | 0.22784 | 0.22176 | 0.21568 | 0.2096 | 0.20496 | 0.20032 | 0.19568 | 0.19104 | 0.1864 | 0.18176 | 0.17712 | 0.17248 | 0.16784 | 0.1632 | 0.16022 | 0.15724 | 0.15426 | 0.15128 | 0.1483 | 0.14532 | 0.14234 | 0.13936 | 0.13638 | 0.1334 | 0.13154 | 0.12968 | 0.12782 | 0.12596 | 0.1241 | 0.12224 | 0.12038 | 0.11852 | 0.11666 | 0.1148
| 0.5 | 0.529 | 0.5149 | 0.5008 | 0.4867 | 0.4726 | 0.4585 | 0.4444 | 0.4303 | 0.4162 | 0.4021 | 0.388 | 0.3777 | 0.3674 | 0.3571 | 0.3468 | 0.3365 | 0.3262 | 0.3159 | 0.3056 | 0.2953 | 0.285 | 0.277 | 0.269 | 0.261 | 0.253 | 0.245 | 0.2388 | 0.2326 | 0.2264 | 0.2202 | 0.214 | 0.2092 | 0.2044 | 0.1996 | 0.1948 | 0.19 | 0.1852 | 0.1804 | 0.1756 | 0.1708 | 0.166 | 0.163 | 0.16 | 0.157 | 0.154 | 0.151 | 0.148 | 0.145 | 0.142 | 0.139 | 0.136 | 0.1341 | 0.1322 | 0.1303 | 0.1284 | 0.1265 | 0.1246 | 0.1227 | 0.1208 | 0.1189 | 0.117
| 0.51 | 0.535 | 0.52092 | 0.50684 | 0.49276 | 0.47868 | 0.4646 | 0.45052 | 0.43644 | 0.42236 | 0.40828 | 0.3942 | 0.3838 | 0.3734 | 0.363 | 0.3526 | 0.3422 | 0.3318 | 0.3214 | 0.311 | 0.3006 | 0.2902 | 0.28224 | 0.27428 | 0.26632 | 0.25836 | 0.2504 | 0.244 | 0.2376 | 0.2312 | 0.2248 | 0.2184 | 0.21362 | 0.20884 | 0.20406 | 0.19928 | 0.1945 | 0.18972 | 0.18494 | 0.18016 | 0.17538 | 0.1706 | 0.1675 | 0.1644 | 0.1613 | 0.1582 | 0.1551 | 0.152 | 0.1489 | 0.1458 | 0.1427 | 0.1396 | 0.13766 | 0.13572 | 0.13378 | 0.13184 | 0.1299 | 0.12796 | 0.12602 | 0.12408 | 0.12214 | 0.1202
| 0.52 | 0.541 | 0.52694 | 0.51288 | 0.49882 | 0.48476 | 0.4707 | 0.45664 | 0.44258 | 0.42852 | 0.41446 | 0.4004 | 0.3899 | 0.3794 | 0.3689 | 0.3584 | 0.3479 | 0.3374 | 0.3269 | 0.3164 | 0.3059 | 0.2954 | 0.28748 | 0.27956 | 0.27164 | 0.26372 | 0.2558 | 0.2492 | 0.2426 | 0.236 | 0.2294 | 0.2228 | 0.21804 | 0.21328 | 0.20852 | 0.20376 | 0.199 | 0.19424 | 0.18948 | 0.18472 | 0.17996 | 0.1752 | 0.172 | 0.1688 | 0.1656 | 0.1624 | 0.1592 | 0.156 | 0.1528 | 0.1496 | 0.1464 | 0.1432 | 0.14122 | 0.13924 | 0.13726 | 0.13528 | 0.1333 | 0.13132 | 0.12934 | 0.12736 | 0.12538 | 0.1234
| 0.53 | 0.547 | 0.53296 | 0.51892 | 0.50488 | 0.49084 | 0.4768 | 0.46276 | 0.44872 | 0.43468 | 0.42064 | 0.4066 | 0.396 | 0.3854 | 0.3748 | 0.3642 | 0.3536 | 0.343 | 0.3324 | 0.3218 | 0.3112 | 0.3006 | 0.29272 | 0.28484 | 0.27696 | 0.26908 | 0.2612 | 0.2544 | 0.2476 | 0.2408 | 0.234 | 0.2272 | 0.22246 | 0.21772 | 0.21298 | 0.20824 | 0.2035 | 0.19876 | 0.19402 | 0.18928 | 0.18454 | 0.1798 | 0.1765 | 0.1732 | 0.1699 | 0.1666 | 0.1633 | 0.16 | 0.1567 | 0.1534 | 0.1501 | 0.1468 | 0.14478 | 0.14276 | 0.14074 | 0.13872 | 0.1367 | 0.13468 | 0.13266 | 0.13064 | 0.12862 | 0.1266
| 0.54 | 0.553 | 0.53898 | 0.52496 | 0.51094 | 0.49692 | 0.4829 | 0.46888 | 0.45486 | 0.44084 | 0.42682 | 0.4128 | 0.4021 | 0.3914 | 0.3807 | 0.37 | 0.3593 | 0.3486 | 0.3379 | 0.3272 | 0.3165 | 0.3058 | 0.29796 | 0.29012 | 0.28228 | 0.27444 | 0.2666 | 0.2596 | 0.2526 | 0.2456 | 0.2386 | 0.2316 | 0.22688 | 0.22216 | 0.21744 | 0.21272 | 0.208 | 0.20328 | 0.19856 | 0.19384 | 0.18912 | 0.1844 | 0.181 | 0.1776 | 0.1742 | 0.1708 | 0.1674 | 0.164 | 0.1606 | 0.1572 | 0.1538 | 0.1504 | 0.14834 | 0.14628 | 0.14422 | 0.14216 | 0.1401 | 0.13804 | 0.13598 | 0.13392 | 0.13186 | 0.1298
| 0.55 | 0.559 | 0.545 | 0.531 | 0.517 | 0.503 | 0.489 | 0.475 | 0.461 | 0.447 | 0.433 | 0.419 | 0.4082 | 0.3974 | 0.3866 | 0.3758 | 0.365 | 0.3542 | 0.3434 | 0.3326 | 0.3218 | 0.311 | 0.3032 | 0.2954 | 0.2876 | 0.2798 | 0.272 | 0.2648 | 0.2576 | 0.2504 | 0.2432 | 0.236 | 0.2313 | 0.2266 | 0.2219 | 0.2172 | 0.2125 | 0.2078 | 0.2031 | 0.1984 | 0.1937 | 0.189 | 0.1855 | 0.182 | 0.1785 | 0.175 | 0.1715 | 0.168 | 0.1645 | 0.161 | 0.1575 | 0.154 | 0.1519 | 0.1498 | 0.1477 | 0.1456 | 0.1435 | 0.1414 | 0.1393 | 0.1372 | 0.1351 | 0.133
| 0.56 | 0.5642 | 0.55026 | 0.53632 | 0.52238 | 0.50844 | 0.4945 | 0.48056 | 0.46662 | 0.45268 | 0.43874 | 0.4248 | 0.41398 | 0.40316 | 0.39234 | 0.38152 | 0.3707 | 0.35988 | 0.34906 | 0.33824 | 0.32742 | 0.3166 | 0.3086 | 0.3006 | 0.2926 | 0.2846 | 0.2766 | 0.26948 | 0.26236 | 0.25524 | 0.24812 | 0.241 | 0.23622 | 0.23144 | 0.22666 | 0.22188 | 0.2171 | 0.21232 | 0.20754 | 0.20276 | 0.19798 | 0.1932 | 0.18964 | 0.18608 | 0.18252 | 0.17896 | 0.1754 | 0.17184 | 0.16828 | 0.16472 | 0.16116 | 0.1576 | 0.15546 | 0.15332 | 0.15118 | 0.14904 | 0.1469 | 0.14476 | 0.14262 | 0.14048 | 0.13834 | 0.1362
| 0.57 | 0.5694 | 0.55552 | 0.54164 | 0.52776 | 0.51388 | 0.5 | 0.48612 | 0.47224 | 0.45836 | 0.44448 | 0.4306 | 0.41976 | 0.40892 | 0.39808 | 0.38724 | 0.3764 | 0.36556 | 0.35472 | 0.34388 | 0.33304 | 0.3222 | 0.314 | 0.3058 | 0.2976 | 0.2894 | 0.2812 | 0.27416 | 0.26712 | 0.26008 | 0.25304 | 0.246 | 0.24114 | 0.23628 | 0.23142 | 0.22656 | 0.2217 | 0.21684 | 0.21198 | 0.20712 | 0.20226 | 0.1974 | 0.19378 | 0.19016 | 0.18654 | 0.18292 | 0.1793 | 0.17568 | 0.17206 | 0.16844 | 0.16482 | 0.1612 | 0.15902 | 0.15684 | 0.15466 | 0.15248 | 0.1503 | 0.14812 | 0.14594 | 0.14376 | 0.14158 | 0.1394
| 0.58 | 0.5746 | 0.56078 | 0.54696 | 0.53314 | 0.51932 | 0.5055 | 0.49168 | 0.47786 | 0.46404 | 0.45022 | 0.4364 | 0.42554 | 0.41468 | 0.40382 | 0.39296 | 0.3821 | 0.37124 | 0.36038 | 0.34952 | 0.33866 | 0.3278 | 0.3194 | 0.311 | 0.3026 | 0.2942 | 0.2858 | 0.27884 | 0.27188 | 0.26492 | 0.25796 | 0.251 | 0.24606 | 0.24112 | 0.23618 | 0.23124 | 0.2263 | 0.22136 | 0.21642 | 0.21148 | 0.20654 | 0.2016 | 0.19792 | 0.19424 | 0.19056 | 0.18688 | 0.1832 | 0.17952 | 0.17584 | 0.17216 | 0.16848 | 0.1648 | 0.16258 | 0.16036 | 0.15814 | 0.15592 | 0.1537 | 0.15148 | 0.14926 | 0.14704 | 0.14482 | 0.1426
| 0.59 | 0.5798 | 0.56604 | 0.55228 | 0.53852 | 0.52476 | 0.511 | 0.49724 | 0.48348 | 0.46972 | 0.45596 | 0.4422 | 0.43132 | 0.42044 | 0.40956 | 0.39868 | 0.3878 | 0.37692 | 0.36604 | 0.35516 | 0.34428 | 0.3334 | 0.3248 | 0.3162 | 0.3076 | 0.299 | 0.2904 | 0.28352 | 0.27664 | 0.26976 | 0.26288 | 0.256 | 0.25098 | 0.24596 | 0.24094 | 0.23592 | 0.2309 | 0.22588 | 0.22086 | 0.21584 | 0.21082 | 0.2058 | 0.20206 | 0.19832 | 0.19458 | 0.19084 | 0.1871 | 0.18336 | 0.17962 | 0.17588 | 0.17214 | 0.1684 | 0.16614 | 0.16388 | 0.16162 | 0.15936 | 0.1571 | 0.15484 | 0.15258 | 0.15032 | 0.14806 | 0.1458
| 0.6 | 0.585 | 0.5713 | 0.5576 | 0.5439 | 0.5302 | 0.5165 | 0.5028 | 0.4891 | 0.4754 | 0.4617 | 0.448 | 0.4371 | 0.4262 | 0.4153 | 0.4044 | 0.3935 | 0.3826 | 0.3717 | 0.3608 | 0.3499 | 0.339 | 0.3302 | 0.3214 | 0.3126 | 0.3038 | 0.295 | 0.2882 | 0.2814 | 0.2746 | 0.2678 | 0.261 | 0.2559 | 0.2508 | 0.2457 | 0.2406 | 0.2355 | 0.2304 | 0.2253 | 0.2202 | 0.2151 | 0.21 | 0.2062 | 0.2024 | 0.1986 | 0.1948 | 0.191 | 0.1872 | 0.1834 | 0.1796 | 0.1758 | 0.172 | 0.1697 | 0.1674 | 0.1651 | 0.1628 | 0.1605 | 0.1582 | 0.1559 | 0.1536 | 0.1513 | 0.149

## c-grid
c-factors
PMMA thickness (mm)
|  | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80
| 0.3 | 0.889 | 0.8941 | 0.8992 | 0.9043 | 0.9094 | 0.9145 | 0.9196 | 0.9247 | 0.9298 | 0.9349 | 0.94 | 0.9503 | 0.9606 | 0.9709 | 0.9812 | 0.9915 | 1.0018 | 1.0121 | 1.0224 | 1.0327 | 1.043 | 1.0562 | 1.0694 | 1.0826 | 1.0958 | 1.109 | 1.12 | 1.131 | 1.142 | 1.153 | 1.164 | 1.173 | 1.182 | 1.191 | 1.2 | 1.209 | 1.218 | 1.227 | 1.236 | 1.245 | 1.254 | 1.2585 | 1.263 | 1.2675 | 1.272 | 1.2765 | 1.281 | 1.2855 | 1.29 | 1.2945 | 1.299 | 1.2998 | 1.3006 | 1.3014 | 1.3022 | 1.303 | 1.3038 | 1.3046 | 1.3054 | 1.3062 | 1.307
| 0.31 | 0.8902 | 0.89524 | 0.90028 | 0.90532 | 0.91036 | 0.9154 | 0.92044 | 0.92548 | 0.93052 | 0.93556 | 0.9406 | 0.9508 | 0.961 | 0.9712 | 0.9814 | 0.9916 | 1.0018 | 1.012 | 1.0222 | 1.0324 | 1.0426 | 1.0557 | 1.0688 | 1.082 | 1.0951 | 1.1082 | 1.1192 | 1.1302 | 1.1412 | 1.1522 | 1.1632 | 1.1721 | 1.181 | 1.1899 | 1.1988 | 1.2077 | 1.2166 | 1.2255 | 1.2344 | 1.2433 | 1.2522 | 1.2567 | 1.2613 | 1.2658 | 1.2704 | 1.2749 | 1.2794 | 1.284 | 1.2885 | 1.2931 | 1.2976 | 1.2984 | 1.2992 | 1.2999 | 1.3007 | 1.3015 | 1.3023 | 1.3031 | 1.3038 | 1.3046 | 1.3054
| 0.32 | 0.8914 | 0.89638 | 0.90136 | 0.90634 | 0.91132 | 0.9163 | 0.92128 | 0.92626 | 0.93124 | 0.93622 | 0.9412 | 0.9513 | 0.9614 | 0.9715 | 0.9816 | 0.9917 | 1.0018 | 1.0119 | 1.022 | 1.0321 | 1.0422 | 1.0552 | 1.0683 | 1.0813 | 1.0944 | 1.1074 | 1.1184 | 1.1294 | 1.1404 | 1.1514 | 1.1624 | 1.1712 | 1.18 | 1.1888 | 1.1976 | 1.2064 | 1.2152 | 1.224 | 1.2328 | 1.2416 | 1.2504 | 1.255 | 1.2596 | 1.2641 | 1.2687 | 1.2733 | 1.2779 | 1.2825 | 1.287 | 1.2916 | 1.2962 | 1.297 | 1.2977 | 1.2985 | 1.2992 | 1.3 | 1.3008 | 1.3015 | 1.3023 | 1.303 | 1.3038
| 0.33 | 0.8926 | 0.89752 | 0.90244 | 0.90736 | 0.91228 | 0.9172 | 0.92212 | 0.92704 | 0.93196 | 0.93688 | 0.9418 | 0.9518 | 0.9618 | 0.9718 | 0.9818 | 0.9918 | 1.0018 | 1.0118 | 1.0218 | 1.0318 | 1.0418 | 1.0548 | 1.0677 | 1.0807 | 1.0936 | 1.1066 | 1.1176 | 1.1286 | 1.1396 | 1.1506 | 1.1616 | 1.1703 | 1.179 | 1.1877 | 1.1964 | 1.2051 | 1.2138 | 1.2225 | 1.2312 | 1.2399 | 1.2486 | 1.2532 | 1.2578 | 1.2625 | 1.2671 | 1.2717 | 1.2763 | 1.2809 | 1.2856 | 1.2902 | 1.2948 | 1.2955 | 1.2963 | 1.297 | 1.2978 | 1.2985 | 1.2992 | 1.3 | 1.3007 | 1.3015 | 1.3022
| 0.34 | 0.8938 | 0.89866 | 0.90352 | 0.90838 | 0.91324 | 0.9181 | 0.92296 | 0.92782 | 0.93268 | 0.93754 | 0.9424 | 0.9523 | 0.9622 | 0.9721 | 0.982 | 0.9919 | 1.0018 | 1.0117 | 1.0216 | 1.0315 | 1.0414 | 1.0543 | 1.0672 | 1.08 | 1.0929 | 1.1058 | 1.1168 | 1.1278 | 1.1388 | 1.1498 | 1.1608 | 1.1694 | 1.178 | 1.1866 | 1.1952 | 1.2038 | 1.2124 | 1.221 | 1.2296 | 1.2382 | 1.2468 | 1.2515 | 1.2561 | 1.2608 | 1.2654 | 1.2701 | 1.2748 | 1.2794 | 1.2841 | 1.2887 | 1.2934 | 1.2941 | 1.2948 | 1.2956 | 1.2963 | 1.297 | 1.2977 | 1.2984 | 1.2992 | 1.2999 | 1.3006
| 0.35 | 0.895 | 0.8998 | 0.9046 | 0.9094 | 0.9142 | 0.919 | 0.9238 | 0.9286 | 0.9334 | 0.9382 | 0.943 | 0.9528 | 0.9626 | 0.9724 | 0.9822 | 0.992 | 1.0018 | 1.0116 | 1.0214 | 1.0312 | 1.041 | 1.0538 | 1.0666 | 1.0794 | 1.0922 | 1.105 | 1.116 | 1.127 | 1.138 | 1.149 | 1.16 | 1.1685 | 1.177 | 1.1855 | 1.194 | 1.2025 | 1.211 | 1.2195 | 1.228 | 1.2365 | 1.245 | 1.2497 | 1.2544 | 1.2591 | 1.2638 | 1.2685 | 1.2732 | 1.2779 | 1.2826 | 1.2873 | 1.292 | 1.2927 | 1.2934 | 1.2941 | 1.2948 | 1.2955 | 1.2962 | 1.2969 | 1.2976 | 1.2983 | 1.299
| 0.36 | 0.8966 | 0.90128 | 0.90596 | 0.91064 | 0.91532 | 0.92 | 0.92468 | 0.92936 | 0.93404 | 0.93872 | 0.9434 | 0.95314 | 0.96288 | 0.97262 | 0.98236 | 0.9921 | 1.0018 | 1.0116 | 1.0213 | 1.0311 | 1.0408 | 1.0535 | 1.0662 | 1.079 | 1.0917 | 1.1044 | 1.1152 | 1.1259 | 1.1367 | 1.1474 | 1.1582 | 1.1667 | 1.1752 | 1.1836 | 1.1921 | 1.2006 | 1.2091 | 1.2176 | 1.226 | 1.2345 | 1.243 | 1.2477 | 1.2524 | 1.2571 | 1.2618 | 1.2665 | 1.2712 | 1.2759 | 1.2806 | 1.2853 | 1.29 | 1.2908 | 1.2915 | 1.2923 | 1.293 | 1.2938 | 1.2946 | 1.2953 | 1.2961 | 1.2968 | 1.2976
| 0.37 | 0.8982 | 0.90276 | 0.90732 | 0.91188 | 0.91644 | 0.921 | 0.92556 | 0.93012 | 0.93468 | 0.93924 | 0.9438 | 0.95348 | 0.96316 | 0.97284 | 0.98252 | 0.9922 | 1.0019 | 1.0116 | 1.0212 | 1.0309 | 1.0406 | 1.0532 | 1.0659 | 1.0785 | 1.0912 | 1.1038 | 1.1143 | 1.1248 | 1.1354 | 1.1459 | 1.1564 | 1.1649 | 1.1733 | 1.1818 | 1.1902 | 1.1987 | 1.2072 | 1.2156 | 1.2241 | 1.2325 | 1.241 | 1.2457 | 1.2504 | 1.2551 | 1.2598 | 1.2645 | 1.2692 | 1.2739 | 1.2786 | 1.2833 | 1.288 | 1.2888 | 1.2896 | 1.2905 | 1.2913 | 1.2921 | 1.2929 | 1.2937 | 1.2946 | 1.2954 | 1.2962
| 0.38 | 0.8998 | 0.90424 | 0.90868 | 0.91312 | 0.91756 | 0.922 | 0.92644 | 0.93088 | 0.93532 | 0.93976 | 0.9442 | 0.95382 | 0.96344 | 0.97306 | 0.98268 | 0.9923 | 1.0019 | 1.0115 | 1.0212 | 1.0308 | 1.0404 | 1.053 | 1.0655 | 1.0781 | 1.0906 | 1.1032 | 1.1135 | 1.1238 | 1.134 | 1.1443 | 1.1546 | 1.163 | 1.1715 | 1.1799 | 1.1884 | 1.1968 | 1.2052 | 1.2137 | 1.2221 | 1.2306 | 1.239 | 1.2437 | 1.2484 | 1.2531 | 1.2578 | 1.2625 | 1.2672 | 1.2719 | 1.2766 | 1.2813 | 1.286 | 1.2869 | 1.2878 | 1.2886 | 1.2895 | 1.2904 | 1.2913 | 1.2922 | 1.293 | 1.2939 | 1.2948
| 0.39 | 0.9014 | 0.90572 | 0.91004 | 0.91436 | 0.91868 | 0.923 | 0.92732 | 0.93164 | 0.93596 | 0.94028 | 0.9446 | 0.95416 | 0.96372 | 0.97328 | 0.98284 | 0.9924 | 1.002 | 1.0115 | 1.0211 | 1.0306 | 1.0402 | 1.0527 | 1.0652 | 1.0776 | 1.0901 | 1.1026 | 1.1126 | 1.1227 | 1.1327 | 1.1428 | 1.1528 | 1.1612 | 1.1696 | 1.1781 | 1.1865 | 1.1949 | 1.2033 | 1.2117 | 1.2202 | 1.2286 | 1.237 | 1.2417 | 1.2464 | 1.2511 | 1.2558 | 1.2605 | 1.2652 | 1.2699 | 1.2746 | 1.2793 | 1.284 | 1.2849 | 1.2859 | 1.2868 | 1.2878 | 1.2887 | 1.2896 | 1.2906 | 1.2915 | 1.2925 | 1.2934
| 0.4 | 0.903 | 0.9072 | 0.9114 | 0.9156 | 0.9198 | 0.924 | 0.9282 | 0.9324 | 0.9366 | 0.9408 | 0.945 | 0.9545 | 0.964 | 0.9735 | 0.983 | 0.9925 | 1.002 | 1.0115 | 1.021 | 1.0305 | 1.04 | 1.0524 | 1.0648 | 1.0772 | 1.0896 | 1.102 | 1.1118 | 1.1216 | 1.1314 | 1.1412 | 1.151 | 1.1594 | 1.1678 | 1.1762 | 1.1846 | 1.193 | 1.2014 | 1.2098 | 1.2182 | 1.2266 | 1.235 | 1.2397 | 1.2444 | 1.2491 | 1.2538 | 1.2585 | 1.2632 | 1.2679 | 1.2726 | 1.2773 | 1.282 | 1.283 | 1.284 | 1.285 | 1.286 | 1.287 | 1.288 | 1.289 | 1.29 | 1.291 | 1.292
| 0.41 | 0.904 | 0.90812 | 0.91224 | 0.91636 | 0.92048 | 0.9246 | 0.92872 | 0.93284 | 0.93696 | 0.94108 | 0.9452 | 0.95466 | 0.96412 | 0.97358 | 0.98304 | 0.9925 | 1.002 | 1.0114 | 1.0209 | 1.0303 | 1.0398 | 1.0521 | 1.0644 | 1.0768 | 1.0891 | 1.1014 | 1.1113 | 1.1212 | 1.131 | 1.1409 | 1.1508 | 1.1591 | 1.1675 | 1.1758 | 1.1842 | 1.1925 | 1.2008 | 1.2092 | 1.2175 | 1.2259 | 1.2342 | 1.2388 | 1.2435 | 1.2481 | 1.2528 | 1.2574 | 1.262 | 1.2667 | 1.2713 | 1.276 | 1.2806 | 1.2816 | 1.2827 | 1.2837 | 1.2848 | 1.2858 | 1.2868 | 1.2879 | 1.2889 | 1.29 | 1.291
HVL | 0.42 | 0.905 | 0.90904 | 0.91308 | 0.91712 | 0.92116 | 0.9252 | 0.92924 | 0.93328 | 0.93732 | 0.94136 | 0.9454 | 0.95482 | 0.96424 | 0.97366 | 0.98308 | 0.9925 | 1.0019 | 1.0113 | 1.0208 | 1.0302 | 1.0396 | 1.0518 | 1.0641 | 1.0763 | 1.0886 | 1.1008 | 1.1108 | 1.1207 | 1.1307 | 1.1406 | 1.1506 | 1.1589 | 1.1672 | 1.1754 | 1.1837 | 1.192 | 1.2003 | 1.2086 | 1.2168 | 1.2251 | 1.2334 | 1.238 | 1.2426 | 1.2471 | 1.2517 | 1.2563 | 1.2609 | 1.2655 | 1.27 | 1.2746 | 1.2792 | 1.2803 | 1.2814 | 1.2824 | 1.2835 | 1.2846 | 1.2857 | 1.2868 | 1.2878 | 1.2889 | 1.29
(mm Al) | 0.43 | 0.906 | 0.90996 | 0.91392 | 0.91788 | 0.92184 | 0.9258 | 0.92976 | 0.93372 | 0.93768 | 0.94164 | 0.9456 | 0.95498 | 0.96436 | 0.97374 | 0.98312 | 0.9925 | 1.0019 | 1.0113 | 1.0206 | 1.03 | 1.0394 | 1.0516 | 1.0637 | 1.0759 | 1.088 | 1.1002 | 1.1102 | 1.1203 | 1.1303 | 1.1404 | 1.1504 | 1.1586 | 1.1668 | 1.1751 | 1.1833 | 1.1915 | 1.1997 | 1.2079 | 1.2162 | 1.2244 | 1.2326 | 1.2371 | 1.2416 | 1.2462 | 1.2507 | 1.2552 | 1.2597 | 1.2642 | 1.2688 | 1.2733 | 1.2778 | 1.2789 | 1.28 | 1.2812 | 1.2823 | 1.2834 | 1.2845 | 1.2856 | 1.2868 | 1.2879 | 1.289
| 0.44 | 0.907 | 0.91088 | 0.91476 | 0.91864 | 0.92252 | 0.9264 | 0.93028 | 0.93416 | 0.93804 | 0.94192 | 0.9458 | 0.95514 | 0.96448 | 0.97382 | 0.98316 | 0.9925 | 1.0018 | 1.0112 | 1.0205 | 1.0299 | 1.0392 | 1.0513 | 1.0634 | 1.0754 | 1.0875 | 1.0996 | 1.1097 | 1.1198 | 1.13 | 1.1401 | 1.1502 | 1.1584 | 1.1665 | 1.1747 | 1.1828 | 1.191 | 1.1992 | 1.2073 | 1.2155 | 1.2236 | 1.2318 | 1.2363 | 1.2407 | 1.2452 | 1.2496 | 1.2541 | 1.2586 | 1.263 | 1.2675 | 1.2719 | 1.2764 | 1.2776 | 1.2787 | 1.2799 | 1.281 | 1.2822 | 1.2834 | 1.2845 | 1.2857 | 1.2868 | 1.288
| 0.45 | 0.908 | 0.9118 | 0.9156 | 0.9194 | 0.9232 | 0.927 | 0.9308 | 0.9346 | 0.9384 | 0.9422 | 0.946 | 0.9553 | 0.9646 | 0.9739 | 0.9832 | 0.9925 | 1.0018 | 1.0111 | 1.0204 | 1.0297 | 1.039 | 1.051 | 1.063 | 1.075 | 1.087 | 1.099 | 1.1092 | 1.1194 | 1.1296 | 1.1398 | 1.15 | 1.1581 | 1.1662 | 1.1743 | 1.1824 | 1.1905 | 1.1986 | 1.2067 | 1.2148 | 1.2229 | 1.231 | 1.2354 | 1.2398 | 1.2442 | 1.2486 | 1.253 | 1.2574 | 1.2618 | 1.2662 | 1.2706 | 1.275 | 1.2762 | 1.2774 | 1.2786 | 1.2798 | 1.281 | 1.2822 | 1.2834 | 1.2846 | 1.2858 | 1.287
| 0.46 | 0.9088 | 0.91258 | 0.91636 | 0.92014 | 0.92392 | 0.9277 | 0.93148 | 0.93526 | 0.93904 | 0.94282 | 0.9466 | 0.9558 | 0.965 | 0.9742 | 0.9834 | 0.9926 | 1.0018 | 1.011 | 1.0202 | 1.0294 | 1.0386 | 1.0506 | 1.0625 | 1.0745 | 1.0864 | 1.0984 | 1.1085 | 1.1186 | 1.1286 | 1.1387 | 1.1488 | 1.1569 | 1.165 | 1.1731 | 1.1812 | 1.1893 | 1.1974 | 1.2055 | 1.2136 | 1.2217 | 1.2298 | 1.2342 | 1.2386 | 1.2431 | 1.2475 | 1.2519 | 1.2563 | 1.2607 | 1.2652 | 1.2696 | 1.274 | 1.2752 | 1.2764 | 1.2777 | 1.2789 | 1.2801 | 1.2813 | 1.2825 | 1.2838 | 1.285 | 1.2862
| 0.47 | 0.9096 | 0.91336 | 0.91712 | 0.92088 | 0.92464 | 0.9284 | 0.93216 | 0.93592 | 0.93968 | 0.94344 | 0.9472 | 0.9563 | 0.9654 | 0.9745 | 0.9836 | 0.9927 | 1.0018 | 1.0109 | 1.02 | 1.0291 | 1.0382 | 1.0501 | 1.062 | 1.074 | 1.0859 | 1.0978 | 1.1078 | 1.1177 | 1.1277 | 1.1376 | 1.1476 | 1.1557 | 1.1638 | 1.1719 | 1.18 | 1.1881 | 1.1962 | 1.2043 | 1.2124 | 1.2205 | 1.2286 | 1.233 | 1.2375 | 1.2419 | 1.2464 | 1.2508 | 1.2552 | 1.2597 | 1.2641 | 1.2686 | 1.273 | 1.2742 | 1.2755 | 1.2767 | 1.278 | 1.2792 | 1.2804 | 1.2817 | 1.2829 | 1.2842 | 1.2854
| 0.48 | 0.9104 | 0.91414 | 0.91788 | 0.92162 | 0.92536 | 0.9291 | 0.93284 | 0.93658 | 0.94032 | 0.94406 | 0.9478 | 0.9568 | 0.9658 | 0.9748 | 0.9838 | 0.9928 | 1.0018 | 1.0108 | 1.0198 | 1.0288 | 1.0378 | 1.0497 | 1.0616 | 1.0734 | 1.0853 | 1.0972 | 1.107 | 1.1169 | 1.1267 | 1.1366 | 1.1464 | 1.1545 | 1.1626 | 1.1707 | 1.1788 | 1.1869 | 1.195 | 1.2031 | 1.2112 | 1.2193 | 1.2274 | 1.2319 | 1.2363 | 1.2408 | 1.2452 | 1.2497 | 1.2542 | 1.2586 | 1.2631 | 1.2675 | 1.272 | 1.2733 | 1.2745 | 1.2758 | 1.277 | 1.2783 | 1.2796 | 1.2808 | 1.2821 | 1.2833 | 1.2846
| 0.49 | 0.9112 | 0.91492 | 0.91864 | 0.92236 | 0.92608 | 0.9298 | 0.93352 | 0.93724 | 0.94096 | 0.94468 | 0.9484 | 0.9573 | 0.9662 | 0.9751 | 0.984 | 0.9929 | 1.0018 | 1.0107 | 1.0196 | 1.0285 | 1.0374 | 1.0492 | 1.0611 | 1.0729 | 1.0848 | 1.0966 | 1.1063 | 1.116 | 1.1258 | 1.1355 | 1.1452 | 1.1533 | 1.1614 | 1.1695 | 1.1776 | 1.1857 | 1.1938 | 1.2019 | 1.21 | 1.2181 | 1.2262 | 1.2307 | 1.2352 | 1.2396 | 1.2441 | 1.2486 | 1.2531 | 1.2576 | 1.262 | 1.2665 | 1.271 | 1.2723 | 1.2736 | 1.2748 | 1.2761 | 1.2774 | 1.2787 | 1.28 | 1.2812 | 1.2825 | 1.2838
| 0.5 | 0.912 | 0.9157 | 0.9194 | 0.9231 | 0.9268 | 0.9305 | 0.9342 | 0.9379 | 0.9416 | 0.9453 | 0.949 | 0.9578 | 0.9666 | 0.9754 | 0.9842 | 0.993 | 1.0018 | 1.0106 | 1.0194 | 1.0282 | 1.037 | 1.0488 | 1.0606 | 1.0724 | 1.0842 | 1.096 | 1.1056 | 1.1152 | 1.1248 | 1.1344 | 1.144 | 1.1521 | 1.1602 | 1.1683 | 1.1764 | 1.1845 | 1.1926 | 1.2007 | 1.2088 | 1.2169 | 1.225 | 1.2295 | 1.234 | 1.2385 | 1.243 | 1.2475 | 1.252 | 1.2565 | 1.261 | 1.2655 | 1.27 | 1.2713 | 1.2726 | 1.2739 | 1.2752 | 1.2765 | 1.2778 | 1.2791 | 1.2804 | 1.2817 | 1.283
| 0.51 | 0.913 | 0.91666 | 0.92032 | 0.92398 | 0.92764 | 0.9313 | 0.93496 | 0.93862 | 0.94228 | 0.94594 | 0.9496 | 0.9583 | 0.967 | 0.9757 | 0.9844 | 0.9931 | 1.0018 | 1.0105 | 1.0192 | 1.0279 | 1.0366 | 1.0483 | 1.06 | 1.0716 | 1.0833 | 1.095 | 1.1046 | 1.1142 | 1.1238 | 1.1334 | 1.143 | 1.151 | 1.1591 | 1.1671 | 1.1752 | 1.1832 | 1.1912 | 1.1993 | 1.2073 | 1.2154 | 1.2234 | 1.2279 | 1.2323 | 1.2368 | 1.2412 | 1.2457 | 1.2502 | 1.2546 | 1.2591 | 1.2635 | 1.268 | 1.2693 | 1.2706 | 1.2719 | 1.2732 | 1.2745 | 1.2758 | 1.2771 | 1.2784 | 1.2797 | 1.281
| 0.52 | 0.914 | 0.91762 | 0.92124 | 0.92486 | 0.92848 | 0.9321 | 0.93572 | 0.93934 | 0.94296 | 0.94658 | 0.9502 | 0.9588 | 0.9674 | 0.976 | 0.9846 | 0.9932 | 1.0018 | 1.0104 | 1.019 | 1.0276 | 1.0362 | 1.0478 | 1.0593 | 1.0709 | 1.0824 | 1.094 | 1.1036 | 1.1132 | 1.1228 | 1.1324 | 1.142 | 1.15 | 1.158 | 1.1659 | 1.1739 | 1.1819 | 1.1899 | 1.1979 | 1.2058 | 1.2138 | 1.2218 | 1.2262 | 1.2306 | 1.2351 | 1.2395 | 1.2439 | 1.2483 | 1.2527 | 1.2572 | 1.2616 | 1.266 | 1.2673 | 1.2686 | 1.2699 | 1.2712 | 1.2725 | 1.2738 | 1.2751 | 1.2764 | 1.2777 | 1.279
| 0.53 | 0.915 | 0.91858 | 0.92216 | 0.92574 | 0.92932 | 0.9329 | 0.93648 | 0.94006 | 0.94364 | 0.94722 | 0.9508 | 0.9593 | 0.9678 | 0.9763 | 0.9848 | 0.9933 | 1.0018 | 1.0103 | 1.0188 | 1.0273 | 1.0358 | 1.0472 | 1.0587 | 1.0701 | 1.0816 | 1.093 | 1.1026 | 1.1122 | 1.1218 | 1.1314 | 1.141 | 1.1489 | 1.1568 | 1.1648 | 1.1727 | 1.1806 | 1.1885 | 1.1964 | 1.2044 | 1.2123 | 1.2202 | 1.2246 | 1.229 | 1.2333 | 1.2377 | 1.2421 | 1.2465 | 1.2509 | 1.2552 | 1.2596 | 1.264 | 1.2653 | 1.2666 | 1.2679 | 1.2692 | 1.2705 | 1.2718 | 1.2731 | 1.2744 | 1.2757 | 1.277
| 0.54 | 0.916 | 0.91954 | 0.92308 | 0.92662 | 0.93016 | 0.9337 | 0.93724 | 0.94078 | 0.94432 | 0.94786 | 0.9514 | 0.9598 | 0.9682 | 0.9766 | 0.985 | 0.9934 | 1.0018 | 1.0102 | 1.0186 | 1.027 | 1.0354 | 1.0467 | 1.058 | 1.0694 | 1.0807 | 1.092 | 1.1016 | 1.1112 | 1.1208 | 1.1304 | 1.14 | 1.1479 | 1.1557 | 1.1636 | 1.1714 | 1.1793 | 1.1872 | 1.195 | 1.2029 | 1.2107 | 1.2186 | 1.2229 | 1.2273 | 1.2316 | 1.236 | 1.2403 | 1.2446 | 1.249 | 1.2533 | 1.2577 | 1.262 | 1.2633 | 1.2646 | 1.2659 | 1.2672 | 1.2685 | 1.2698 | 1.2711 | 1.2724 | 1.2737 | 1.275
| 0.55 | 0.917 | 0.9205 | 0.924 | 0.9275 | 0.931 | 0.9345 | 0.938 | 0.9415 | 0.945 | 0.9485 | 0.952 | 0.9603 | 0.9686 | 0.9769 | 0.9852 | 0.9935 | 1.0018 | 1.0101 | 1.0184 | 1.0267 | 1.035 | 1.0462 | 1.0574 | 1.0686 | 1.0798 | 1.091 | 1.1006 | 1.1102 | 1.1198 | 1.1294 | 1.139 | 1.1468 | 1.1546 | 1.1624 | 1.1702 | 1.178 | 1.1858 | 1.1936 | 1.2014 | 1.2092 | 1.217 | 1.2213 | 1.2256 | 1.2299 | 1.2342 | 1.2385 | 1.2428 | 1.2471 | 1.2514 | 1.2557 | 1.26 | 1.2613 | 1.2626 | 1.2639 | 1.2652 | 1.2665 | 1.2678 | 1.2691 | 1.2704 | 1.2717 | 1.273
| 0.56 | 0.9178 | 0.92124 | 0.92468 | 0.92812 | 0.93156 | 0.935 | 0.93844 | 0.94188 | 0.94532 | 0.94876 | 0.9522 | 0.96046 | 0.96872 | 0.97698 | 0.98524 | 0.9935 | 1.0018 | 1.01 | 1.0183 | 1.0265 | 1.0348 | 1.0459 | 1.057 | 1.0682 | 1.0793 | 1.0904 | 1.0999 | 1.1094 | 1.119 | 1.1285 | 1.138 | 1.1457 | 1.1534 | 1.1611 | 1.1688 | 1.1765 | 1.1842 | 1.1919 | 1.1996 | 1.2073 | 1.215 | 1.2193 | 1.2236 | 1.2278 | 1.2321 | 1.2364 | 1.2407 | 1.245 | 1.2492 | 1.2535 | 1.2578 | 1.2591 | 1.2604 | 1.2617 | 1.263 | 1.2643 | 1.2656 | 1.2669 | 1.2682 | 1.2695 | 1.2708
| 0.57 | 0.9186 | 0.92198 | 0.92536 | 0.92874 | 0.93212 | 0.9355 | 0.93888 | 0.94226 | 0.94564 | 0.94902 | 0.9524 | 0.96062 | 0.96884 | 0.97706 | 0.98528 | 0.9935 | 1.0017 | 1.0099 | 1.0182 | 1.0264 | 1.0346 | 1.0456 | 1.0567 | 1.0677 | 1.0788 | 1.0898 | 1.0992 | 1.1087 | 1.1181 | 1.1276 | 1.137 | 1.1446 | 1.1522 | 1.1598 | 1.1674 | 1.175 | 1.1826 | 1.1902 | 1.1978 | 1.2054 | 1.213 | 1.2173 | 1.2215 | 1.2258 | 1.23 | 1.2343 | 1.2386 | 1.2428 | 1.2471 | 1.2513 | 1.2556 | 1.2569 | 1.2582 | 1.2595 | 1.2608 | 1.2621 | 1.2634 | 1.2647 | 1.266 | 1.2673 | 1.2686
| 0.58 | 0.9194 | 0.92272 | 0.92604 | 0.92936 | 0.93268 | 0.936 | 0.93932 | 0.94264 | 0.94596 | 0.94928 | 0.9526 | 0.96078 | 0.96896 | 0.97714 | 0.98532 | 0.9935 | 1.0017 | 1.0099 | 1.018 | 1.0262 | 1.0344 | 1.0454 | 1.0563 | 1.0673 | 1.0782 | 1.0892 | 1.0986 | 1.1079 | 1.1173 | 1.1266 | 1.136 | 1.1435 | 1.151 | 1.1585 | 1.166 | 1.1735 | 1.181 | 1.1885 | 1.196 | 1.2035 | 1.211 | 1.2152 | 1.2195 | 1.2237 | 1.228 | 1.2322 | 1.2364 | 1.2407 | 1.2449 | 1.2492 | 1.2534 | 1.2547 | 1.256 | 1.2573 | 1.2586 | 1.2599 | 1.2612 | 1.2625 | 1.2638 | 1.2651 | 1.2664
| 0.59 | 0.9202 | 0.92346 | 0.92672 | 0.92998 | 0.93324 | 0.9365 | 0.93976 | 0.94302 | 0.94628 | 0.94954 | 0.9528 | 0.96094 | 0.96908 | 0.97722 | 0.98536 | 0.9935 | 1.0016 | 1.0098 | 1.0179 | 1.0261 | 1.0342 | 1.0451 | 1.056 | 1.0668 | 1.0777 | 1.0886 | 1.0979 | 1.1072 | 1.1164 | 1.1257 | 1.135 | 1.1424 | 1.1498 | 1.1572 | 1.1646 | 1.172 | 1.1794 | 1.1868 | 1.1942 | 1.2016 | 1.209 | 1.2132 | 1.2174 | 1.2217 | 1.2259 | 1.2301 | 1.2343 | 1.2385 | 1.2428 | 1.247 | 1.2512 | 1.2525 | 1.2538 | 1.2551 | 1.2564 | 1.2577 | 1.259 | 1.2603 | 1.2616 | 1.2629 | 1.2642
| 0.6 | 0.921 | 0.9242 | 0.9274 | 0.9306 | 0.9338 | 0.937 | 0.9402 | 0.9434 | 0.9466 | 0.9498 | 0.953 | 0.9611 | 0.9692 | 0.9773 | 0.9854 | 0.9935 | 1.0016 | 1.0097 | 1.0178 | 1.0259 | 1.034 | 1.0448 | 1.0556 | 1.0664 | 1.0772 | 1.088 | 1.0972 | 1.1064 | 1.1156 | 1.1248 | 1.134 | 1.1413 | 1.1486 | 1.1559 | 1.1632 | 1.1705 | 1.1778 | 1.1851 | 1.1924 | 1.1997 | 1.207 | 1.2112 | 1.2154 | 1.2196 | 1.2238 | 1.228 | 1.2322 | 1.2364 | 1.2406 | 1.2448 | 1.249 | 1.2503 | 1.2516 | 1.2529 | 1.2542 | 1.2555 | 1.2568 | 1.2581 | 1.2594 | 1.2607 | 1.262

## Thickness
PMMA (mm) | Breast (mm)
0 | 0
1 | 1.05
2 | 2.1
3 | 3.15
4 | 4.2
5 | 5.25
6 | 6.3
7 | 7.35
8 | 8.4
9 | 9.45
10 | 10.5
11 | 11.55
12 | 12.6
13 | 13.65
14 | 14.7
15 | 15.75
16 | 16.8
17 | 17.85
18 | 18.9
19 | 19.95
20 | 21
21 | 22.1
22 | 23.2
23 | 24.3
24 | 25.4
25 | 26.5
26 | 27.6
27 | 28.7
28 | 29.8
29 | 30.9
30 | 32
31 | 33.3
32 | 34.6
33 | 35.9
34 | 37.2
35 | 38.5
36 | 39.8
37 | 41.1
38 | 42.4
39 | 43.7
40 | 45
41 | 46.6
42 | 48.2
43 | 49.8
44 | 51.4
45 | 53
46 | 54.4
47 | 55.8
48 | 57.2
49 | 58.6
50 | 60
51 | 61.5
52 | 63
53 | 64.5
54 | 66
55 | 67.5
56 | 69
57 | 70.5
58 | 72
59 | 73.5
60 | 75
61 | 76.5
62 | 78
63 | 79.5
64 | 81
65 | 82.5
66 | 84
67 | 85.5
68 | 87
69 | 88.5
70 | 90
71 | 91.3
72 | 92.6
73 | 93.9
74 | 95.2
75 | 96.5
76 | 97.8
77 | 99.1
78 | 100.4
79 | 101.7
80 | 103

## CDMAM truth
T (mm) | D=2.00 | D=1.60 | D=1.25 | D=1.00 | D=0.8 | D=0.63 | D=0.50 | D=0.40 | D=0.31 | D=0.25 | D=0.20 | D=0.16 | D=0.13 | D=0.10 | D=0.08 | D=0.06 | T (mm)
2 | [6 empty] | U | L | U | L | D | L | L | D | U |  | 2
1.42 | [5 empty] | L | D | D | L | L | D | R | U | R | R | D | 1.42
1 |  |  |  |  | L | D | L | L | R | D | R | L | D | D | R | R | 1
0.71 |  |  |  | U | R | U | R | U | U | D | R | U | U | R | D | U | 0.71
0.5 |  |  | U | D | R | U | L | L | R | R | D | D | L | D | L | U | 0.5
0.36 |  | D | D | L | U | L | R | R | D | R | R | U | R | R | L | D | 0.36
0.25 | R | L | R | R | U | D | U | D | U | D | U | D | U | U | R | U | 0.25
0.2 | U | U | R | U | D | R | U | D | R | L | U | D | D | R | R | U | 0.2
0.16 | L | L | D | R | U | L | R | L | U | L | R | U | R | U | U | D | 0.16
0.13 | U | R | L | L | D | D | L | D | U | R | D | U | D | L | R |  | 0.13
0.1 | R | R | D | U | L | L | U | R | D | R | U | D | D | U |  |  | 0.1
0.08 | D | U | D | R | U | U | D | R | L | L | D | L | U |  |  |  | 0.08
0.06 | L | R | L | D | L | D | D | U | D | U | R | D |  |  |  |  | 0.06
0.05 | U | U | L | R | R | L | U | L | L | D | L | [5 empty] | 0.05
0.04 | R | L | R | D | L | U | D | L | R | U | [6 empty] | 0.04
0.03 |  | L | L | U | L | D | L | U | R | [7 empty] | 0.03
Truth