## Summary
Medical Physicist CT Survey Report | [5 empty] | Pass
Fail
This report summarizes the results of tests performed in accordance with the American College of
Radiology CT QC Manual. | [5 empty] | Pass
Fail
Facility Name |  |  | Unit ID |  |  | NA
Address 1 |  |  | Manufacturer
Address 2 |  |  | Model
City, State, ZIP |  |  | Serial Number
Date of Manufacture
CTAP # (if applicable)
|  |  | Survey Date | 0
Medical Physicist | Eugene Mah |  | Report Date
Signature
Medical Physicist Tests |  | Pass/Fail |  | Technologist QC Evaluation | Pass/Fail/NA
| Review of CT Protocols |  |  | Water CT Number and SD (Daily)
| Scout Prescription Accuracy |  |  | Artifact Evaluation (Daily)
| Alignment Light Accuracy |  |  | Wet Laser QC (Weekly)
| Table Travel Accuracy |  |  | Visual Checklist (Monthly)
| Radiation Beam Width |  |  | Dry Laser QC (Monthly)
| Low-Contrast Performance |  |  | Acquisition Display QC (Monthly)
Spatial Resolution
CT Number Accuracy
Artifact Evaluation
Dosimetry
CT Number Uniformity
Acquisition Display Calibration
Comments

## Sheet1
1 | [11 empty] | Revision 1.1-20190606 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | First,CTDI,ACRPhantom
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [22 empty] | All: | First,CTDI,ACRPhantom,Comments
5 | [5 empty] | CT System Compliance Inspection | [10 empty] | CT System Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  | Change | First,CTDI,ACRPhantom
8 | [5 empty] | System Information | [5 empty] | Previous Date: | [10 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility:
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department:
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division:
13 |  |  | Survey ID: |  |  |  |  | Accession Number: | [5 empty] | Survey ID: |  |  |  | Accession Number: |  |  |  |  | Survey ID:
14 | [13 empty] | Room Number: | [8 empty] | Room Number:
15 |  | CT Scanner | [20 empty] | Site Number:
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | CT Scanner | [9 empty] | Authorized Use:
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation:
18 | [7 empty] | SoftwareVersion: | [5 empty] | Model: |  |  |  | Serial Number: |  |  |  |  | Accesssion Number:
19 |  | X-Ray Tube | [15 empty] | Software Version:
20 |  |  | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube | [9 empty] | X-Ray Generator
21 |  |  | Manufacture Date: |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Manufacturer:
22 |  | Insert | [5 empty] | Small: | [5 empty] | Manufacture Date: |  |  |  | Large: |  |  |  |  | Model:
23 |  |  | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Manufacture Date:
24 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Micro: |  |  |  |  | Serial Number:
25 |  |  | Serial Number: |  |  |  |  | Filtration | [5 empty] | Model: | [8 empty] | Software Version:
26 |  | Housing | [5 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | Filtration
27 |  |  | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
28 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Added: |  |  |  |  | Tube Designation/Use:
29 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | Manufacture Date:
30 | [13 empty] | Serial Number: | [8 empty] | Manufacturer:
31 | [22 empty] | Model:
32 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance | [12 empty] | Serial Number:
33 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Inspection Results | [5 empty] | Manufacturer:
34 | DHEC RHB 10.2.1 |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Model:
35 |  |  | Pregnancy Warning sign is posted | [6 empty] | TBD | [12 empty] | Serial Number:
36 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD | [6 empty] | Labels, Notices, Postings | [5 empty] | Large:
37 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | DHEC Registration sticker is present, clearly visible and legible | [9 empty] | Small:
38 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [9 empty] | Micro:
39 | [12 empty] | Pregnancy Warning sign is posted | [9 empty] | Inherent:
40 |  |  |  | Radiation Safety | [8 empty] | Radiation warning label posted on the generator control panel | [9 empty] | Added:
41 |  |  | Radiation warning light at room entrance is functional with selected tube | [6 empty] | TBD |  |  | Operator manuals are available. | [9 empty] | Measurement Protocols
42 |  |  | Lead aprons available. | [6 empty] | TBD |  |  | Monthly radiation monitoring reports are posted. | [9 empty] | Adult Abdomen
43 |  |  | A properly designed and installed apron rack is present. | [6 empty] | TBD | [12 empty] | kVp
44 |  |  | Documentation of annual protective apparel integrity inspection is available. | [6 empty] | TBD | [6 empty] | Radiation Safety | [5 empty] | mA
45 |  |  | Patient restraint devices available. | [6 empty] | TBD |  |  | Radiation warning light at room entrance is functional with selected tube | [9 empty] | mAs
46 | [12 empty] | Lead aprons available. | [9 empty] | Eff mAs
47 |  |  | Reconstructed images are free of artifacts | [6 empty] | TBD |  |  | A properly designed and installed apron rack is present. | [9 empty] | Time/rotation (s)
48 |  |  | Acquisition station display performance is adequate | [6 empty] | TBD |  |  | Documentation of annual protective apparel integrity inspection is available. | [9 empty] | Scan FOV (cm)
49 |  |  | Daily QC performed by technologist | [6 empty] | TBD |  |  | Patient restraint devices available. | [9 empty] | Display FOV (cm)
50 |  |  | Interlock systems function to prevent scanning while door is open | [6 empty] | TBD | [12 empty] | Reconstruction Algorithm
51 | DHEC RHB 4.11.1.1 |  | Tomographic plane is indicated by a light source | [6 empty] | TBD |  |  | Reconstructed images are free of artifacts | [9 empty] | Axial/Helical
52 | DHEC RHB 4.11.1.3.1 |  | Visual indication of x-ray production is present and functional | [6 empty] | TBD |  |  | Acquisition station display performance is adequate | [9 empty] | Z-Axis collimation
53 | DHEC RHB 4.11.1.6.1 |  | Positioning with the alignment light is accurate to within 2mm | [6 empty] | TBD |  |  | Daily QC performed by technologist | [9 empty] | # of channels
54 | DHEC RHB 4.11.1.6.3 |  | Table travel is accurate to within 1mm | [6 empty] | TBD |  |  | Interlock systems function to prevent scanning while door is open | [9 empty] | Table Increment/speed
55 | DHEC RHB 4.11.2.4.1 |  | Technologist has a clear view of the patient during the scan | [6 empty] | TBD |  |  | Tomographic plane is indicated by a light source | [9 empty] | Pitch
56 | DHEC RHB 4.11.2.1 |  | Patient communication system is operational | [6 empty] | TBD |  |  | Visual indication of x-ray production is present and functional | [9 empty] | Recon scan width (mm)
57 | DHEC RHB 4.11.2.1 |  | Exposure switch is secured at a safe location | [6 empty] | TBD |  |  | Positioning with the alignment light is accurate to within 2mm | [9 empty] | Recon scan interval (mm)
58 | [12 empty] | Table travel is accurate to within 1mm | [9 empty] | Dose reduction techniques
59 |  |  |  | Radiation Dose Measurements | [8 empty] | Technologist has a clear view of the patient during the scan | [9 empty] | Indiated CTDIvol (mGy)
60 |  |  | Routine head protocol CTDI is below ACR reference level of 75 mGy | [6 empty] | NA |  |  | Patient communication system is operational | [9 empty] | Adult Chest
61 |  |  | Routine body protocol CTDI is below ACR reference level of 25 mGy | [6 empty] | NA |  |  | Exposure switch is secured at a safe location | [9 empty] | kVp
62 |  |  | Pediatric body protocol CTDI is below ACR reference level of 15 mGy | [6 empty] | NA | [6 empty] | Radiation Dose Measurements | [5 empty] | mA
63 |  |  | Pediatric head protocol CTDI is below ACR reference level of 35 mGy | [6 empty] | NA |  | 3 | Routine head protocol CTDI is below ACR reference level of 75 mGy | [9 empty] | mAs
64 | [11 empty] | 3 | Routine body protocol CTDI is below ACR reference level of 25 mGy | [9 empty] | Eff mAs
65 | Date: | [8 empty] | Inspector: | Eugene Mah | 3 | Pediatric body protocol CTDI is below ACR reference level of 15 mGy | [9 empty] | Time/rotation (s)
66 | Room Number | [8 empty] | Survey ID: |  | 3 | Pediatric head protocol CTDI is below ACR reference level of 35 mGy | [9 empty] | Scan FOV (cm)
1 | [10 empty] | Medical University of South Carolina | [11 empty] | Display FOV (cm)
2 | [5 empty] | Inspection Results |  |  |  |  | CT System Compliance Inspection | [11 empty] | Reconstruction Algorithm
3 | [5 empty] | Measurement Protocols | [10 empty] | Measurement Protocols | [5 empty] | Axial/Helical
4 |  |  |  | Adult |  |  | Pediatric | [6 empty] | Adult |  |  | Pediatric | [5 empty] | Z-Axis collimation
5 |  |  |  | Abdomen | Chest | Head | Abdomen | Head | [5 empty] | Abdomen | Chest | Head | Abdomen | Head |  |  |  |  | # of channels
6 |  |  | kVp | [9 empty] | kVp | [9 empty] | Table Increment/speed
7 |  |  | mA | [9 empty] | mA | [9 empty] | Pitch
8 |  |  | mAs | [9 empty] | mAs | [9 empty] | Recon scan width (mm)
9 |  |  | Eff mAs | [9 empty] | Eff mAs | [9 empty] | Recon scan interval (mm)
10 |  |  | Time/rotation (s) | [9 empty] | Time/rotation (s) | [9 empty] | Dose reduction techniques
11 |  |  | Scan FOV (cm) | [9 empty] | Scan FOV (cm) | [9 empty] | Indiated CTDIvol (mGy)
12 |  |  | Display FOV (cm) | [9 empty] | Display FOV (cm) | [9 empty] | Adult Head
13 |  |  | Reconstruction Algorithm | [9 empty] | Reconstruction Algorithm | [9 empty] | kVp
14 |  |  | Axial/Helical | [9 empty] | Axial/Helical | [9 empty] | mA
15 |  |  | Z-Axis collimation | [9 empty] | Z-Axis collimation | [9 empty] | mAs
16 |  |  | # of channels | [9 empty] | # of channels | [9 empty] | Eff mAs
17 |  |  | Table Increment/speed | [9 empty] | Table Increment/speed | [9 empty] | Time/rotation (s)
18 |  |  | Pitch | [9 empty] | Pitch | [9 empty] | Scan FOV (cm)
19 |  |  | Recon scan width (mm) | [9 empty] | Recon scan width (mm) | [9 empty] | Display FOV (cm)
20 |  |  | Recon scan interval (mm) | [9 empty] | Recon scan interval (mm) | [9 empty] | Reconstruction Algorithm
21 |  |  | Dose reduction techniques | [9 empty] | Dose reduction techniques | [9 empty] | Axial/Helical
22 |  |  | Indiated CTDIvol (mGy) | [9 empty] | Indiated CTDIvol (mGy) | [9 empty] | Z-Axis collimation
23 | [13 empty] | Adult |  |  | Pediatric | [5 empty] | # of channels
24 | CTDI Measurements | [12 empty] | Abdomen | Chest | Head | Abdomen | Head |  |  |  |  | Table Increment/speed
25 | Meter: | Piranha |  |  | Calibration Date: | [7 empty] | kVp | [9 empty] | Pitch
26 | Serial #: |  |  |  | Calibration Due: | [7 empty] | mA | [9 empty] | Recon scan width (mm)
27 | [12 empty] | mAs | [9 empty] | Recon scan interval (mm)
28 |  |  |  | Adult |  |  | Pediatric | [5 empty] | Eff mAs | [9 empty] | Dose reduction techniques
29 |  |  |  | Abdomen | Chest | Head | Abdomen | Head | Criteria: |  |  |  | Time/rotation (s) | [9 empty] | Indiated CTDIvol (mGy)
30 |  |  | Indicated CTDIvol (mGy): | [5 empty] | CTDIvol is less than ACR Pass/Fail dose |  |  |  | Scan FOV (cm) | [9 empty] | Ped Abdomen
31 |  | CTDIvol Measured (mGy) | [7 empty] | Reference | Pass/Fail |  | Display FOV (cm) | [9 empty] | kVp
32 | [9 empty] | (mGy) | (mGy) |  | Reconstruction Algorithm | [9 empty] | mA
33 | [8 empty] | Adult Abdomen | 25 | 30 |  | Axial/Helical | [9 empty] | mAs
34 |  |  | Average CTDIvol: | [5 empty] | Adult Head: | 75 | 80 |  | Z-Axis collimation | [9 empty] | Eff mAs
35 |  |  | Phantom size indicated (cm): | [5 empty] | Ped Abd (16cm) | 15 | 20 |  | # of channels | [9 empty] | Time/rotation (s)
36 |  |  | Phantom size used (cm): | [5 empty] | Ped Abd (32cm) | 7.5 | 10 |  | Table Increment/speed | [9 empty] | Scan FOV (cm)
37 |  |  | Variation from previous year: | [5 empty] | Pediatric Head: | 35 | 40 |  | Pitch | [9 empty] | Display FOV (cm)
38 |  |  | Variation from indicated CTDI: | [9 empty] | Recon scan width (mm) | [9 empty] | Reconstruction Algorithm
39 |  |  | Acceptable: | [9 empty] | Recon scan interval (mm) | [9 empty] | Axial/Helical
40 |  |  | CTDIvol/mAs: | [9 empty] | Dose reduction techniques | [9 empty] | Z-Axis collimation
41 |  |  | Criteria: | CTDI is within 5% of previous year. | [8 empty] | Indiated CTDIvol (mGy) | [9 empty] | # of channels
42 |  |  |  | Measured CTDI is within 20% of indicated CTDI | [18 empty] | Table Increment/speed
43 | [22 empty] | Pitch
44 | Beam Profile | [15 empty] | Inspection Results | [5 empty] | Recon scan width (mm)
45 |  | Channels (N) | [9 empty] | CTDI Measurements | [10 empty] | Recon scan interval (mm)
46 |  | Detector width T (mm) | [9 empty] | Meter: | Piranha |  | Calibration Date: | [7 empty] | Dose reduction techniques
47 |  | NxT Beam width (mm) | [9 empty] | Serial #: |  |  | Calibration Date: | [7 empty] | Indiated CTDIvol (mGy)
48 |  | Nominal beam width (mm) | [20 empty] | Ped Head
49 |  | Average beam width (mm) | [11 empty] | Adult |  |  | Pediatric |  |  | Criteria: |  |  | kVp
50 | [13 empty] | Abdomen | Chest | Head | Abdomen | Head |  | CTDIvol is less than ACR Pass/Fail dose |  |  | mA
51 |  | Acceptable: | [10 empty] | Indicated CTDIvol (mGy) | [7 empty] | Reference | Pass/Fail | mAs
52 |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value | [7 empty] | CTDIvol Measured (mGy) | [8 empty] | (mGy) | (mGy) | Eff mAs
53 | [19 empty] | Adult Abdomen | 25 | 30 | Time/rotation (s)
54 | [19 empty] | Adult Head: | 75 | 80 | Scan FOV (cm)
55 | Artifact Evaluation | [11 empty] | Average CTDIvol: | [6 empty] | Ped Abd (16cm) | 15 | 20 | Display FOV (cm)
56 |  | Water phantom is free from rings, streaks, lines, cupping artifacts | [10 empty] | Phantom size indicated (cm) | [6 empty] | Ped Abd (32cm) | 7.5 | 10 | Reconstruction Algorithm
57 | [12 empty] | Phantom size used (cm) | [6 empty] | Pediatric Head: | 35 | 40 | Axial/Helical
58 | Acquisition Console Monitor | [11 empty] | Previous year: | [9 empty] | Z-Axis collimation
59 |  | SMPTE 5%/95% visible |  |  | Lmin (cd/m^2) | [7 empty] | Variation from prev year: | [9 empty] | # of channels
60 |  | Display free of artifacts |  |  | Lmax (cd/m^2) | [7 empty] | Variation from indicated: | [9 empty] | Table Increment/speed
61 |  | Display uniformity |  |  | % Non-uniformity | [7 empty] | Acceptable: | [9 empty] | Pitch
62 | [12 empty] | CTDIvol/100 mAs: | [9 empty] | Recon scan width (mm)
63 | [12 empty] | Prev CTDIvol/100 mAs: | [9 empty] | Recon scan interval (mm)
64 | [12 empty] | Criteria: | CTDI is within 5% of previous year. | [8 empty] | Dose reduction techniques
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  |  | Measured CTDI is within 20% of indicated CTDI | [8 empty] | Indiated CTDIvol (mGy)
66 | Room Number | [8 empty] | Survey ID: |  | Beam Profile | [10 empty] | CTDI Measurements
1 | [10 empty] | Medical University of South Carolina |  | Channels (N) | [9 empty] | Adult Abdomen
2 | [5 empty] | Inspection Results |  |  |  |  | CT System Compliance Inspection |  | Detector width T (mm) |  |  |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value |  |  |  | Adult Chest
3 | Alignment Accuracy | [11 empty] | NxT Beam width (mm) | [9 empty] | Adult Head
4 |  | Landmark position: | 0 | mm |  | Slice thickness: | 1 | mm |  |  |  |  | Nominal beam width (mm) | [9 empty] | Ped Abdomen
5 | [12 empty] | Measured beam width (mm) | [9 empty] | Ped Head
6 |  |  |  | Foot | Head |  |  |  | Scout prescription accuracy: | [13 empty] | CTDIvol/100 mAs
7 |  |  | Location of alignment slice: |  |  | mm |  |  | Central wire visible: | [13 empty] | Adult Abdomen
8 |  |  | All markers visible | [5 empty] | Table movement accuracy: |  |  |  | Average (mm): | [9 empty] | Adult Chest
9 | [8 empty] | Alignment light accuracy: |  |  |  | Acceptable: | [9 empty] | Adult Head
10 | [22 empty] | Ped Abdomen
11 | CT Number Accuracy | [10 empty] | Artifact Evaluation | [10 empty] | Ped Head
12 |  | kVp | 80 | 100 | 110 | 120 | 130 | 140 | 150 |  |  |  | Water phantom is free from rings, streaks, lines, cupping artifacts
13 |  | mAs | [20 empty] | Comment Page 1
14 |  | Polyethylene
15 |  | Water | [14 empty] | Inspection Results
16 |  | Water Std Dev | [9 empty] | Alignment Accuracy
17 |  | Acrylic | [10 empty] | Landmark position: | 0 | mm | Slice thickness | 1 | mm
18 |  | Bone
19 |  | Air | [11 empty] | Foot | Head |  |  |  | Enter 1 for YES, 2 for NO
20 |  |  |  |  | Acceptable: | [7 empty] | Location of alignment slice: |  |  | mm |  | Scout prescription accuracy:
21 |  |  | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp | [8 empty] | All markers visible: |  |  |  |  | Central wire visible:
22 |  |  |  | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp | [13 empty] | Table movement accuracy:
23 | [17 empty] | Alignment light accuracy:
24 | CT Slice Thickness Accuracy | [10 empty] | CT Number Accuracy
25 |  |  | Slice thickness (mm): | [8 empty] | kVp | 80 | 100 | 110 | 120 | 130 | 140 | 150 |  | Acceptable HU
26 |  |  | Top (mm): | [8 empty] | mAs | [8 empty] | Low | High
27 |  |  | Bottom (mm): | [8 empty] | Polyethylene | [7 empty] | Polyethylene | -107 | -84
28 |  |  | Average (mm): | [8 empty] | Water | [7 empty] | Water | -7 | 7
29 |  |  | Water HU: | [8 empty] | Water Std Dev | [7 empty] | Acrylic | 110 | 135
30 |  |  | Water HU Std Dev: | [8 empty] | Acrylic | [7 empty] | Bone | 850 | 970
31 |  |  | Acceptable: | [8 empty] | Bone | [7 empty] | Air | -1005 | -970
32 |  |  |  | Criteria: | Slice thickness is within 1.5mm of the set thickness | [6 empty] | Air
33 |  |  |  |  | CT number for water is 0 +/- 5 | [9 empty] | Acceptable: |  | Enter 1 for YES, 2 for NO
34 | [12 empty] | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
35 | Low Contrast Performance | [12 empty] | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
36 |  |  | Adult |  | Pediatric
37 |  |  | Abdomen | Head | Abdomen | Head | [5 empty] | CT Slice Thickness Accuracy
38 |  | Inside | [5 empty] | 6mm rods visible: |  |  |  |  | Slice thickness (mm):
39 |  | Adjacent | [5 empty] | Acceptable: |  |  |  |  | Top (mm):
40 |  | Adjacent StDev | [10 empty] | Bottom (mm):
41 |  | Contrast | [10 empty] | Average (mm):
42 |  | CNR | [10 empty] | Water HU:
43 |  |  |  | Criteria: | All four 6mm rods are visible | [7 empty] | Water HU Std Dev:
44 |  |  |  |  | CNR > 1.0 for adult abdomen and head protocols | [7 empty] | Acceptable:
45 |  |  |  |  | CNR > 0.4 for pediatric abdomen protocol | [8 empty] | Criteria: | Slice thickness is within 1.5mm of the set thickness
46 |  |  |  |  | CNR > 1.0 for pediatric head protocol | [9 empty] | CT number for water is 0 +/- 5
47 | Uniformity
48 |  | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable | [5 empty] | Low Contrast Performance
49 | Center: | [11 empty] | Adult |  | Pediatric
50 | Top: | [11 empty] | Abdomen | Head | Abdomen | Head |  |  | Enter 1 for YES, 2 for NO, 3 for Not Applicable
51 | Bottom: | [10 empty] | Inside | [5 empty] | 6mm rods visible: |  | Acceptable:
52 | Right: | [10 empty] | Adjacent | [5 empty] | Criteria: | All four 6mm rods are visible
53 | Left: | [10 empty] | Adj StDev | [6 empty] | CNR > 1.0 for adult abdomen and head protocols
54 |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region | [8 empty] | Contrast | [6 empty] | CNR > 0.4 for pediatric abdomen protocol
55 |  | Uniformity image is free from artifacts | [9 empty] | CNR | [6 empty] | CNR > 1.0 for pediatric head protocol
56
57 | High Contrast Resolution | [10 empty] | Uniformity
58 |  |  | lp/cm | Acceptable | [8 empty] | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable
59 |  | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen |  |  |  |  | Center:
60 |  | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution |  |  |  |  | Top:
61 | [11 empty] | Bottom:
62 | [11 empty] | Right:
63 | [11 empty] | Left:
64 | [12 empty] | Uniformity image is free from artifacts |  |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region
65 | Date: | [8 empty] | Inspector: | Eugene Mah | High Contrast Resolution
66 | Room Number | [8 empty] | Survey ID: |  |  | lp/cm | Acceptable
1 | [10 empty] | Medical University of South Carolina | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen
2 | [5 empty] | Comments Page 1 |  |  |  |  | CT System Compliance Inspection | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution
3
4 | [11 empty] | Acquisition Console Monitor
5 | [15 empty] | Display uniformity
6 | [12 empty] | SMPTE 5%/95% visible |  |  | Top left:
7 | [12 empty] | Display free of artifacts |  |  | Top right:
8 | [12 empty] | Display uniformity |  |  | Center:
9 | [12 empty] | Lmin (cd/m^2) |  |  | Bottom left:
10 | [12 empty] | Lmax (cd/m^2) |  |  | Bottom right:
11 | [11 empty] | Critiera: | Lmin <= 1.2 cd/m^2, Lmax >= 90 cd/m^2 |  |  | Average:
12 | [12 empty] | Display uniformity is +/-15% |  |  | SD:
13 | [15 empty] | % Non-uniformity:
14
15 | [16 empty] | Comments Page 1
16 | [15 empty] | Previous Year Comments
17 | [12 empty] | Comments:
18 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
19 | [12 empty] | Additional Comments:
20 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
21 | [12 empty] | Additional Comments:
22 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
23 | [12 empty] | Additional Comments:
24 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
25 | [12 empty] | Additional Comments:
26 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
27 | [12 empty] | Additional Comments:
28 | [12 empty] | Do not exceed 390 chars |  | 0
29 | [12 empty] | Additional Comments:
30 | [12 empty] | Do not exceed 390 chars |  | 0
31 | [12 empty] | Additional Comments:
32 | [12 empty] | Do not exceed 390 chars |  | 0
33 | [12 empty] | Additional Comments:
34 | [12 empty] | Do not exceed 390 chars |  | 0
35 | [12 empty] | Additional Comments:
36 | [12 empty] | Do not exceed 390 chars |  | 0
37 | [12 empty] | Additional Comments:
38 | [12 empty] | Do not exceed 390 chars |  | 0
39 | [12 empty] | Additional Comments:
40 | [12 empty] | Do not exceed 390 chars |  | 0
41 | [12 empty] | Additional Comments:
42 | [12 empty] | Do not exceed 390 chars |  | 0
43 | [12 empty] | Additional Comments:
44 | [12 empty] | Do not exceed 390 chars |  | 0
45 | [12 empty] | Additional Comments:
46 | [12 empty] | Do not exceed 390 chars |  | 0
47 | [12 empty] | Additional Comments:
48 | [12 empty] | Do not exceed 390 chars |  | 0
49 | [12 empty] | Additional Comments:
50 | [12 empty] | Do not exceed 390 chars |  | 0
51 | [12 empty] | Additional Comments:
52 | [12 empty] | Do not exceed 390 chars |  | 0
53 | [12 empty] | Additional Comments:
54 | [12 empty] | Do not exceed 390 chars |  | 0
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:

## Sheet2
Siemens CT Beam Profiles http://128.23.56.214/MPwiki/index.php?title=Siemens_CT_Beam_Profiles
CT Beam Profile – Siemens
NxT | Definition Flash
| Z-sharp | Nominal (mm) | Tolerance)
2x1 | Off | 3.5 | 1
1x5 | Off | 6 | 1
10x0.6 | On | 10.5 | 1.5
8x0.6 | On | 11.5 | 1.5
1x10 | Off | 11.1 | 1.5
20x0.6 | On | 18.8 | 1.5
12x1.2 | Off | 19.4 | 1.5
32x0.6 | Off | 24.4 | 2
32x0.6 | On | 27.4 | 2
32x1.2 | Off | 44 | 2
64x0.6 | On | 45.5 | 2
NxT | Definition AS
| Z-sharp | Nominal (mm) | Tolerance)
2x1 | Off | 4.7 | 1
1x5 | Off | 6.6 | 1
10x0.6 | On | 11.2 | 1
8x0.6 | Off | 8.5 | 1
1x10 | Off | 11.6 | 1
8x0.6 | On | 11.5 | 1
6x1.2 | Off | 11.8 | 1
12x1.2 | Off | 19.8 | 1
60x0.6 | Off | 41 | 1.5
64x0.6 | On | 45.5 | 1.5
32x1.2 | Off | 43.5 | 1.5
30x0.6 | Off | 23.5 | 1
32x0.6 | On | 28 | 1
20x0.6 | Off | 18 | 1
16x1.2 | Off | 25 | 1
20x0.6 | On | 20 | 1
NxT | Sensation 64
| Nominal (mm) | Tolerance)
30x0.6 | 22.5 | 1.5
24x1.2 | 33.5 | 1.5
12x1.2 | 19.5 | 1.5
1x10 | 11.5 | 1.5
1x5 | 6.5 | 1.5
2x1 | 4 | 1
6x0.6 | 7 | 1
NxT | Sensation 16
| Nominal (mm) | Tolerance)
16x1.5 | 30.5 | 1
16x0.75 | 18.5 | 1
12x1.5 | 24.5 | 1
12x0.75 | 15.5 | 1
2x5 | 12.5 | 1
2x1 | 3.5 | 1
2x0.6 | 3 | 1
NxT | Emotion 16
| Nominal (mm) | Tolerance)
4x0.6 | 4.9 | 1
12x0.6 | 12.6 | 1
16x0.6 | 14.5 | 1
2x5 | 12.6 | 1
12x1.2 | 18.5 | 1
2x8 | 18.5 | 1
16x1.2 | 24.2 | 1
NxT | Somatom Force
| Nominal (mm) | Tolerance)
2x1 | 3 | 1
6x0.5 | 7 | 1
1x5 | 7.5 | 1
1x10 | 12 | 1.5
12x1.2 | 24 | 2
32x0.6 (UHR) | 22.5 | 2
32x0.6 (sUHR) | 24.5 | 2
32x0.6 (StdHR) | 28.5 | 2
64x0.6 | 47.5 | 2
96x0.6 | 66 | 2
48x1.2 | 67 | 2
NxT | Somatom Confidence
| Nominal (mm) | Tolerance)
20x0.6 | 18.8 | 1.5
2x1 | 3.5 | 1
12x1.2 | 19.8 | 1.5
16x1.2 | 25 | 2