## Summary
Medical Physicist CT Survey Report

This report summarizes the results of tests performed in accordance with the American College of
Radiology CT QC Manual.

Facility Name			Unit ID
Address 1			Manufacturer
Address 2			Model
City, State, ZIP			Serial Number
			Date of Manufacture
CTAP # (if applicable)
			Survey Date	30-Dec-1899
Medical Physicist	Eugene Mah		Report Date
Signature

Medical Physicist Tests		Pass/Fail		Technologist QC Evaluation	Pass/Fail/NA

	Review of CT Protocols			Water CT Number and SD (Daily)
	Scout Prescription Accuracy			Artifact Evaluation (Daily)
	Alignment Light Accuracy			Wet Laser QC (Weekly)
	Table Travel Accuracy			Visual Checklist (Monthly)
	Radiation Beam Width			Dry Laser QC (Monthly)
	Low-Contrast Performance			Acquisition Display QC (Monthly)
	Spatial Resolution
	CT Number Accuracy
	Artifact Evaluation
	Dosimetry
	CT Number Uniformity
	Acquisition Display Calibration

Comments

## Sheet1
1					Revision 1.0-20140806					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					CT System Compliance Inspection					CT System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:	Eugene Mah	Eugene Mah		Print Area		Change	Page1
8					System Information					Previous Date:					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector		Change	Eugene Mah
10				Facility:					Site Number:					Facility:				Site Number:					Facility:
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:
13				Survey ID:					Accession Number:					Survey ID:				Accession Number:					Survey ID:
14					Room Number:					Room Number:
15			CT Scanner					Site Number:
16				Manufacturer:					Manufacture Date:					CT Scanner					Authorized Use:
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:
18					SoftwareVersion:					Model:				Serial Number:					Accesssion Number:
19			X-Ray Tube					Software Version:
20				Tube Designation/Use:				Focal Spot Sizes (mm)					X-Ray Tube					X-Ray Generator
21				Manufacture Date:					Large:					Tube Designation/Use:				Focal Spot Sizes (mm)					Manufacturer:
22			Insert					Small:					Manufacture Date:				Large:					Model:
23				Manufacturer:					Micro:					Insert					Small:					Manufacture Date:
24				Model:					Manufacturer:				Micro:					Serial Number:
25				Serial Number:					Filtration					Model:					Software Version:
26			Housing					Inherent:					Serial Number:				Filtration
27				Manufacturer:					Added:					Housing					Inherent:					X-Ray Tube 1
28				Model:					Manufacturer:				Added:					Tube Designation/Use:
29				Serial Number:					Model:					Manufacture Date:
30					Serial Number:					Manufacturer:
31					Model:
32		Rule Number			Labels, Notices, Postings					Compliance					Serial Number:
33		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Inspection Results					Manufacturer:
34		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD			Enter 1 for YES, 2 for NO, 3 for NA					Model:
35				Pregnancy Warning sign is posted					TBD					Serial Number:
36		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD					Labels, Notices, Postings					Large:
37				Operator manuals are available.					TBD				DHEC Registration sticker is present, clearly visible and legible					Small:
38				Monthly radiation monitoring reports are posted.					TBD				DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Micro:
39					Pregnancy Warning sign is posted					Inherent:
40					Radiation Safety					Radiation warning label posted on the generator control panel					Added:
41				Radiation warning light at room entrance is functional with selected tube					TBD				Operator manuals are available.					Measurement Protocols
42				Lead aprons available.					TBD				Monthly radiation monitoring reports are posted.					Adult Abdomen
43				A properly designed and installed apron rack is present.					TBD					kVp
44				Documentation of annual protective apparel integrity inspection is available.					TBD					Radiation Safety					mA
45				Patient restraint devices available.					TBD				Radiation warning light at room entrance is functional with selected tube					mAs
46					Lead aprons available.					Eff mAs
47				Reconstructed images are free of artifacts					TBD				A properly designed and installed apron rack is present.					Time/rotation (s)
48				Acquisition station display performance is adequate					TBD				Documentation of annual protective apparel integrity inspection is available.					Scan FOV (cm)
49				Daily QC performed by technologist					TBD				Patient restraint devices available.					Display FOV (cm)
50				Interlock systems function to prevent scanning while door is open					TBD					Reconstruction Algorithm
51		DHEC RHB 4.11.1.1		Tomographic plane is indicated by a light source					TBD				Reconstructed images are free of artifacts					Axial/Helical
52		DHEC RHB 4.11.1.3.1		Visual indication of x-ray production is present and functional					TBD				Acquisition station display performance is adequate					Z-Axis collimation
53		DHEC RHB 4.11.1.6.1		Positioning with the alignment light is accurate to within 2mm					TBD				Daily QC performed by technologist					# of channels
54		DHEC RHB 4.11.1.6.3		Table travel is accurate to within 1mm					TBD				Interlock systems function to prevent scanning while door is open					Table Increment/speed
55		DHEC RHB 4.11.2.5.1		Technologist has a clear view of the patient during the scan					TBD				Tomographic plane is indicated by a light source					Pitch
56				Patient communication system is operational					TBD				Visual indication of x-ray production is present and functional					Recon scan width (mm)
57					Positioning with the alignment light is accurate to within 2mm					Recon scan interval (mm)
58					Radiation Dose Measurements					Table travel is accurate to within 1mm					Dose reduction techniques
59				Routine head protocol CTDI is below ACR reference level of 75 mGy					NA				Technologist has a clear view of the patient during the scan					Indiated CTDIvol (mGy)
60				Routine body protocol CTDI is below ACR reference level of 25 mGy					NA				Patient communication system is operational					Adult Chest
61				Pediatric body protocol CTDI is below ACR reference level of 15 mGy					NA					kVp
62				Pediatric head protocol CTDI is below ACR reference level of 35 mGy					NA					Radiation Dose Measurements					mA
63					3	Routine head protocol CTDI is below ACR reference level of 75 mGy					mAs
64					3	Routine body protocol CTDI is below ACR reference level of 25 mGy					Eff mAs
65		Date:					Inspector:	Eugene Mah		3	Pediatric body protocol CTDI is below ACR reference level of 15 mGy					Time/rotation (s)
66		Room Number					Survey ID:			3	Pediatric head protocol CTDI is below ACR reference level of 35 mGy					Scan FOV (cm)
1					Medical University of South Carolina					Display FOV (cm)
2					Inspection Results					CT System Compliance Inspection					Reconstruction Algorithm
3					Measurement Protocols					Measurement Protocols					Axial/Helical
4					Adult			Pediatric					Adult			Pediatric					Z-Axis collimation
5					Abdomen	Chest	Head	Abdomen	Head					Abdomen	Chest	Head	Abdomen	Head					# of channels
6				kVp					kVp					Table Increment/speed
7				mA					mA					Pitch
8				mAs					mAs					Recon scan width (mm)
9				Eff mAs					Eff mAs					Recon scan interval (mm)
10				Time/rotation (s)					Time/rotation (s)					Dose reduction techniques
11				Scan FOV (cm)					Scan FOV (cm)					Indiated CTDIvol (mGy)
12				Display FOV (cm)					Display FOV (cm)					Adult Head
13				Reconstruction Algorithm					Reconstruction Algorithm					kVp
14				Axial/Helical					Axial/Helical					mA
15				Z-Axis collimation					Z-Axis collimation					mAs
16				# of channels					# of channels					Eff mAs
17				Table Increment/speed					Table Increment/speed					Time/rotation (s)
18				Pitch					Pitch					Scan FOV (cm)
19				Recon scan width (mm)					Recon scan width (mm)					Display FOV (cm)
20				Recon scan interval (mm)					Recon scan interval (mm)					Reconstruction Algorithm
21				Dose reduction techniques					Dose reduction techniques					Axial/Helical
22				Indiated CTDIvol (mGy)					Indiated CTDIvol (mGy)					Z-Axis collimation
23					Adult			Pediatric					# of channels
24		CTDI Measurements					Abdomen	Chest	Head	Abdomen	Head					Table Increment/speed
25					Adult			Pediatric					kVp					Pitch
26					Abdomen	Chest	Head	Abdomen	Head	Criteria:					mA					Recon scan width (mm)
27				Indicated CTDIvol (mGy):					CTDIvol is less than ACR Pass/Fail dose					mAs					Recon scan interval (mm)
28			CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					Reference	Pass/Fail			Eff mAs					Dose reduction techniques
29					(mGy)	(mGy)			Time/rotation (s)					Indiated CTDIvol (mGy)
30					Adult Abdomen	25	30			Scan FOV (cm)					Ped Abdomen
31				Average CTDIvol:					Adult Head:	75	80			Display FOV (cm)					kVp
32				Phantom size indicated (cm):					Ped Abd (16cm)	15	20			Reconstruction Algorithm					mA
33				Phantom size used (cm):					Ped Abd (32cm)	7.5	10			Axial/Helical					mAs
34				Variation from previous year:					Pediatric Head:	35	40			Z-Axis collimation					Eff mAs
35				Variation from indicated CTDI:					# of channels					Time/rotation (s)
36				Acceptable:					Table Increment/speed					Scan FOV (cm)
37				CTDIvol/mAs:					Pitch					Display FOV (cm)
38				Criteria:	CTDI is within 5% of previous year.					Recon scan width (mm)					Reconstruction Algorithm
39					Measured CTDI is within 20% of indicated CTDI					Recon scan interval (mm)					Axial/Helical
40					Dose reduction techniques					Z-Axis collimation
41		Beam Profile					Indiated CTDIvol (mGy)					# of channels
42			Channels (N)					Table Increment/speed
43			Detector width T (mm)					Pitch
44			NxT Beam width (mm)	0	0	0	0					Inspection Results					Recon scan width (mm)
45			Nominal beam width (mm)					CTDI Measurements					Recon scan interval (mm)
46			Measured beam width (mm)					Adult			Pediatric			Criteria:				Dose reduction techniques
47					Abdomen	Chest	Head	Abdomen	Head		CTDIvol is less than ACR Pass/Fail dose				Indiated CTDIvol (mGy)
48					Indicated CTDIvol (mGy)					Reference	Pass/Fail		Ped Head
49					CTDIvol Measured (mGy)	CTDIvol Measured (mGy)					(mGy)	(mGy)		kVp
50			Acceptable:					Adult Abdomen	25	30		mA
51				Criteria:	Beam profile is within 3mm or 30% of specified value					Adult Head:	75	80		mAs
52					Average CTDIvol:					Ped Abd (16cm)	15	20		Eff mAs
53					Phantom size indicated (cm)					Ped Abd (32cm)	7.5	10		Time/rotation (s)
54					Phantom size used (cm)					Pediatric Head:	35	40		Scan FOV (cm)
55		Artifact Evaluation					Previous year:					Display FOV (cm)
56			Water phantom is free from rings, streaks, lines, cupping artifacts					Variation from prev year:					Reconstruction Algorithm
57					Variation from indicated:					Axial/Helical
58		Acquisition Console Monitor					Acceptable:					Z-Axis collimation
59			SMPTE 5%/95% visible			Lmin (cd/m^2)					CTDIvol/100 mAs:					# of channels
60			Display free of artifacts			Lmax (cd/m^2)					Prev CTDIvol/100 mAs:					Table Increment/speed
61			Display uniformity		#DIV/0!	% Non-uniformity					Criteria:	CTDI is within 5% of previous year.					Pitch
62					Measured CTDI is within 20% of indicated CTDI					Recon scan width (mm)
63					Beam Profile					Recon scan interval (mm)
64					Channels (N)					Dose reduction techniques
65		Date:					Inspector:	Eugene Mah			Detector width T (mm)					Criteria:	Beam profile is within 3mm or 30% of specified value					Indiated CTDIvol (mGy)
66		Room Number					Survey ID:				NxT Beam width (mm)	0	0	0	0					CTDI Measurements
1					Medical University of South Carolina			Nominal beam width (mm)					Adult Abdomen
2					Inspection Results					CT System Compliance Inspection			Measured beam width (mm)					Adult Chest
3		Alignment Accuracy					Adult Head
4			Landmark position:		mm		Slice thickness:		mm					Ped Abdomen
5					Average (mm):					Ped Head
6					Foot	Head				Scout prescription accuracy:					Acceptable:					CTDIvol/100 mAs
7				Location of alignment slice:			mm			Central wire visible:					Adult Abdomen
8				All markers visible					Table movement accuracy:				Artifact Evaluation					Adult Chest
9					Alignment light accuracy:					Water phantom is free from rings, streaks, lines, cupping artifacts					Adult Head
10					Ped Abdomen
11		CT Number Accuracy					Ped Head
12			kVp	80	100	110	120	130	140	150					Inspection Results
13			mAs					Alignment Accuracy					Comment Page 1
14			Polyethylene					Landmark position:		mm	Slice thickness		mm
15			Water
16			Water Std Dev					Foot	Head				Enter 1 for YES, 2 for NO
17			Acrylic					Location of alignment slice:			mm		Scout prescription accuracy:
18			Bone					All markers visible:					Central wire visible:
19			Air					Table movement accuracy:
20					Acceptable:					Alignment light accuracy:
21				Critiera:	Water HU is 0 +/- 5, air HU is within acceptable range at all kVp					CT Number Accuracy
22					CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp					kVp	80	100	110	120	130	140	150		Acceptable HU
23					mAs					Low	High
24		CT Slice Thickness Accuracy					Polyethylene					Polyethylene	-107	-84
25				Slice thickness (mm):					Water					Water	-7	7
26				Top (mm):					Water Std Dev					Acrylic	110	135
27				Bottom (mm):					Acrylic					Bone	850	970
28				Average (mm):					Bone					Air	-1005	-970
29				Water HU:					Air
30				Water HU Std Dev:					Acceptable:		Enter 1 for YES, 2 for NO
31				Acceptable:					Critiera:	Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
32					Criteria:	Slice thickness is within 1.5mm of the set thickness					CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
33					CT number for water is 0 +/- 5
34					CT Slice Thickness Accuracy
35		Low Contrast Performance					Slice thickness (mm):
36				Adult		Pediatric					Top (mm):
37				Abdomen	Head	Abdomen	Head					Bottom (mm):
38			Inside					6mm rods visible:					Average (mm):
39			Adjacent					Acceptable:					Water HU:
40			Adjacent StDev					Water HU Std Dev:
41			Contrast					Acceptable:
42			CNR					Criteria:	Slice thickness is within 1.5mm of the set thickness
43					Criteria:	All four 6mm rods are visible					CT number for water is 0 +/- 5
44					CNR > 1.0 for adult abdomen and head protocols
45					CNR > 0.4 for pediatric abdomen protocol					Low Contrast Performance
46					CNR > 1.0 for pediatric head protocol					Adult		Pediatric
47		Uniformity					Abdomen	Head	Abdomen	Head			Enter 1 for YES, 2 for NO, 3 for Not Applicable
48			Mean	St Dev	Diff (Center)	Diff (Mean)	Acceptable					Inside					6mm rods visible:		Acceptable:
49		Center:					Adjacent					Criteria:	All four 6mm rods are visible
50		Top:					Adj StDev					CNR > 1.0 for adult abdomen and head protocols
51		Bottom:					Contrast					CNR > 0.4 for pediatric abdomen protocol
52		Right:					CNR					CNR > 1.0 for pediatric head protocol
53		Left:
54			Criteria:	CT number of all ROIs is within +/- 5 HU of the central region					Uniformity
55			Uniformity image is free from artifacts					Mean	St Dev	Diff (Center)	Diff (Mean)	Acceptable
56					Center:
57		High Contrast Resolution					Top:
58				lp/cm	Acceptable					Bottom:
59			Abdomen				Critiera:	6 lp/cm for abdomen					Right:
60			Hi-res Chest					8 lp/cm for high resolution					Left:
61					Uniformity image is free from artifacts			Criteria:	CT number of all ROIs is within +/- 5 HU of the central region
62					High Contrast Resolution
63					lp/cm	Acceptable
64					Abdomen				Critiera:	6 lp/cm for abdomen
65		Date:					Inspector:	Eugene Mah		Hi-res Chest					8 lp/cm for high resolution
66		Room Number					Survey ID:
1					Medical University of South Carolina		Acquisition Console Monitor
2					Comments Page 1					CT System Compliance Inspection					Display uniformity
3					SMPTE 5%/95% visible			Top left:
4					Display free of artifacts			Top right:
5					Display uniformity			Center:
6					Lmin (cd/m^2)			Bottom left:
7					Lmax (cd/m^2)			Bottom right:
8					Critiera:	Lmin <= 1.2 cd/m^2, Lmax >= 90 cd/m^2			Average:	#DIV/0!
9					Display uniformity is +/-15%			SD:	#DIV/0!
10					% Non-uniformity:	#DIV/0!
11
12					Comments Page 1
13					Previous Year Comments
14					Comments:
15					Do not exceed 390 chars		0			**Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
16					Additional Comments:
17					Do not exceed 390 chars		0			**Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
18					Additional Comments:
19					Do not exceed 390 chars		0			A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
20					Additional Comments:
21					Do not exceed 390 chars		0			All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
22					Additional Comments:
23					Do not exceed 390 chars		0			**The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
24					Additional Comments:
25					Do not exceed 390 chars		0
26					Additional Comments:
27					Do not exceed 390 chars		0
28					Additional Comments:
29					Do not exceed 390 chars		0
30					Additional Comments:
31					Do not exceed 390 chars		0
32					Additional Comments:
33					Do not exceed 390 chars		0
34					Additional Comments:
35					Do not exceed 390 chars		0
36					Additional Comments:
37					Do not exceed 390 chars		0
38					Additional Comments:
39					Do not exceed 390 chars		0
40					Additional Comments:
41					Do not exceed 390 chars		0
42					Additional Comments:
43					Do not exceed 390 chars		0
44					Additional Comments:
45					Do not exceed 390 chars		0
46					Additional Comments:
47					Do not exceed 390 chars		0
48					Additional Comments:
49					Do not exceed 390 chars		0
50					Additional Comments:
51					Do not exceed 390 chars		0
52
53
54
55
56
57
58
59
60
61
62
63
64
65		Date:					Inspector:	Eugene Mah
66		Room Number					Survey ID:

## Sheet2
CT Beam Profile – Siemens
NxT	Definition Flash
	Z-sharp	Nominal (mm)	Tolerance)
2x1	Off	3.5	1.0
1x5	Off	6.0	1.0
10x0.6	On	10.5	1.5
8x0.6	On	11.5	1.5
1x10	Off	11.1	1.5
20x0.6	On	18.8	1.5
12x1.2	Off	19.4	1.5
32x0.6	Off	24.4	2.0
32x0.6	On	27.4	2.0
32x1.2	Off	44.0	2.0
64x0.6	On	45.5	2.0

NxT	Definition AS
	Z-sharp	Nominal (mm)	Tolerance)
2x1	Off	4.7	1.0
1x5	Off	6.6	1.0
10x0.6	On	11.2	1.0
8x0.6	Off	8.5	1.0
1x10	Off	11.6	1.0
8x0.6	On	11.5	1.0
6x1.2	Off	11.8	1.0
12x1.2	Off	19.8	1.0
60x0.6	Off	41.0	1.5
64x0.6	On	45.5	1.5
32x1.2	Off	43.5	1.5
30x0.6	Off	23.5	1.0
32x0.6	On	28.0	1.0
20x0.6	Off	18.0	1.0
16x1.2	Off	25.0	1.0
20x0.6	On	20.0	1.0

NxT	Sensation 64
	Nominal (mm)	Tolerance)
30x0.6	22.5	1.5
24x1.2	33.5	1.5
12x1.2	19.5	1.5
1x10	11.5	1.5
1x5	6.5	1.5
2x1	4.0	1.0
6x0.6	7.0	1.0

NxT	Sensation 16
	Nominal (mm)	Tolerance)
16x1.5	30.5	1.0
16x0.75	18.5	1.0
12x1.5	24.5	1.0
12x0.75	15.5	1.0
2x5	12.5	1.0
2x1	3.5	1.0
2x0.6	3.0	1.0

NxT	Emotion 16
	Nominal (mm)	Tolerance)
4x0.6	4.9	1.0
12x0.6	12.6	1.0
16x0.6	14.5	1.0
2x5	12.6	1.0
12x1.2	18.5	1.0
2x8	18.5	1.0
16x1.2	24.2	1.0

NxT	Somatom Force
	Nominal (mm)	Tolerance)
2x1	3.0	1.0
6x0.5	7.0	1.0
1x5	7.5	1.0
1x10	12.0	1.5
12x1.2	24.0	2.0
32x0.6 (UHR)	22.5	2.0
32x0.6 (sUHR)	24.5	2.0
32x0.6 (StdHR)	28.5	2.0
64x0.6	47.5	2.0
96x0.6	66.0	2.0
48x1.2	67.0	2.0