## Summary
Medical Physicist CT Survey Report | [5 empty] | Pass
Fail
This report summarizes the results of tests performed in accordance with the American College of
Radiology CT QC Manual. | [5 empty] | Pass
Fail
Facility Name |  |  | Unit ID |  |  | NA
Address 1 |  |  | Manufacturer
Address 2 |  |  | Model
City, State, ZIP |  |  | Serial Number
Date of Manufacture
CTAP # (if applicable)
|  |  | Survey Date | 0
Medical Physicist | Eugene Mah |  | Report Date
Signature
Medical Physicist Tests |  | Pass/Fail |  | Technologist QC Evaluation | Pass/Fail/NA
| Review of CT Protocols |  |  | Water CT Number and SD (Daily)
| Scout Prescription Accuracy |  |  | Artifact Evaluation (Daily)
| Alignment Light Accuracy |  |  | Wet Laser QC (Weekly)
| Table Travel Accuracy |  |  | Visual Checklist (Monthly)
| Radiation Beam Width |  |  | Dry Laser QC (Monthly)
| Low-Contrast Performance |  |  | Acquisition Display QC (Monthly)
Spatial Resolution
CT Number Accuracy
Artifact Evaluation
Dosimetry
CT Number Uniformity
Acquisition Display Calibration
Comments

## Sheet1
1 | [11 empty] | Revision 1.1-20190606 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | Page1
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [22 empty] | All: | Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5 | [5 empty] | CT System Compliance Inspection | [10 empty] | CT System Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  | Change | Page1
8 | [5 empty] | System Information | [5 empty] | Previous Date: | [10 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility:
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department:
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division:
13 |  |  | Survey ID: |  |  |  |  | Accession Number: | [5 empty] | Survey ID: |  |  |  | Accession Number: |  |  |  |  | Survey ID:
14 | [13 empty] | Room Number: | [8 empty] | Room Number:
15 |  | CT Scanner | [20 empty] | Site Number:
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | CT Scanner | [9 empty] | Authorized Use:
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation:
18 | [7 empty] | SoftwareVersion: | [5 empty] | Model: |  |  |  | Serial Number: |  |  |  |  | Accesssion Number:
19 |  | X-Ray Tube | [15 empty] | Software Version:
20 |  |  | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube | [9 empty] | X-Ray Generator
21 |  |  | Manufacture Date: |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Manufacturer:
22 |  | Insert | [5 empty] | Small: | [5 empty] | Manufacture Date: |  |  |  | Large: |  |  |  |  | Model:
23 |  |  | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Manufacture Date:
24 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Micro: |  |  |  |  | Serial Number:
25 |  |  | Serial Number: |  |  |  |  | Filtration | [5 empty] | Model: | [8 empty] | Software Version:
26 |  | Housing | [5 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | Filtration
27 |  |  | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
28 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Added: |  |  |  |  | Tube Designation/Use:
29 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | Manufacture Date:
30 | [13 empty] | Serial Number: | [8 empty] | Manufacturer:
31 | [22 empty] | Model:
32 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance | [12 empty] | Serial Number:
33 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Inspection Results | [5 empty] | Manufacturer:
34 | DHEC RHB 10.2.1 |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Model:
35 |  |  | Pregnancy Warning sign is posted | [6 empty] | TBD | [12 empty] | Serial Number:
36 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD | [6 empty] | Labels, Notices, Postings | [5 empty] | Large:
37 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | DHEC Registration sticker is present, clearly visible and legible | [9 empty] | Small:
38 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [9 empty] | Micro:
39 | [12 empty] | Pregnancy Warning sign is posted | [9 empty] | Inherent:
40 |  |  |  | Radiation Safety | [8 empty] | Radiation warning label posted on the generator control panel | [9 empty] | Added:
41 |  |  | Radiation warning light at room entrance is functional with selected tube | [6 empty] | TBD |  |  | Operator manuals are available. | [9 empty] | Measurement Protocols
42 |  |  | Lead aprons available. | [6 empty] | TBD |  |  | Monthly radiation monitoring reports are posted. | [9 empty] | Adult Abdomen
43 |  |  | A properly designed and installed apron rack is present. | [6 empty] | TBD | [12 empty] | kVp
44 |  |  | Documentation of annual protective apparel integrity inspection is available. | [6 empty] | TBD | [6 empty] | Radiation Safety | [5 empty] | mA
45 |  |  | Patient restraint devices available. | [6 empty] | TBD |  |  | Radiation warning light at room entrance is functional with selected tube | [9 empty] | mAs
46 | [12 empty] | Lead aprons available. | [9 empty] | Eff mAs
47 |  |  | Reconstructed images are free of artifacts | [6 empty] | TBD |  |  | A properly designed and installed apron rack is present. | [9 empty] | Time/rotation (s)
48 |  |  | Acquisition station display performance is adequate | [6 empty] | TBD |  |  | Documentation of annual protective apparel integrity inspection is available. | [9 empty] | Scan FOV (cm)
49 |  |  | Daily QC performed by technologist | [6 empty] | TBD |  |  | Patient restraint devices available. | [9 empty] | Display FOV (cm)
50 |  |  | Interlock systems function to prevent scanning while door is open | [6 empty] | TBD | [12 empty] | Reconstruction Algorithm
51 | DHEC RHB 4.11.1.1 |  | Tomographic plane is indicated by a light source | [6 empty] | TBD |  |  | Reconstructed images are free of artifacts | [9 empty] | Axial/Helical
52 | DHEC RHB 4.11.1.3.1 |  | Visual indication of x-ray production is present and functional | [6 empty] | TBD |  |  | Acquisition station display performance is adequate | [9 empty] | Z-Axis collimation
53 | DHEC RHB 4.11.1.6.1 |  | Positioning with the alignment light is accurate to within 2mm | [6 empty] | TBD |  |  | Daily QC performed by technologist | [9 empty] | # of channels
54 | DHEC RHB 4.11.1.6.3 |  | Table travel is accurate to within 1mm | [6 empty] | TBD |  |  | Interlock systems function to prevent scanning while door is open | [9 empty] | Table Increment/speed
55 | DHEC RHB 4.11.2.5.1 |  | Technologist has a clear view of the patient during the scan | [6 empty] | TBD |  |  | Tomographic plane is indicated by a light source | [9 empty] | Pitch
56 |  |  | Patient communication system is operational | [6 empty] | TBD |  |  | Visual indication of x-ray production is present and functional | [9 empty] | Recon scan width (mm)
57 | [12 empty] | Positioning with the alignment light is accurate to within 2mm | [9 empty] | Recon scan interval (mm)
58 |  |  |  | Radiation Dose Measurements | [8 empty] | Table travel is accurate to within 1mm | [9 empty] | Dose reduction techniques
59 |  |  | Routine head protocol CTDI is below ACR reference level of 75 mGy | [6 empty] | NA |  |  | Technologist has a clear view of the patient during the scan | [9 empty] | Indiated CTDIvol (mGy)
60 |  |  | Routine body protocol CTDI is below ACR reference level of 25 mGy | [6 empty] | NA |  |  | Patient communication system is operational | [9 empty] | Adult Chest
61 |  |  | Pediatric body protocol CTDI is below ACR reference level of 15 mGy | [6 empty] | NA | [12 empty] | kVp
62 |  |  | Pediatric head protocol CTDI is below ACR reference level of 35 mGy | [6 empty] | NA | [6 empty] | Radiation Dose Measurements | [5 empty] | mA
63 | [11 empty] | 3 | Routine head protocol CTDI is below ACR reference level of 75 mGy | [9 empty] | mAs
64 | [11 empty] | 3 | Routine body protocol CTDI is below ACR reference level of 25 mGy | [9 empty] | Eff mAs
65 | Date: | [8 empty] | Inspector: | Eugene Mah | 3 | Pediatric body protocol CTDI is below ACR reference level of 15 mGy | [9 empty] | Time/rotation (s)
66 | Room Number | [8 empty] | Survey ID: |  | 3 | Pediatric head protocol CTDI is below ACR reference level of 35 mGy | [9 empty] | Scan FOV (cm)
1 | [10 empty] | Medical University of South Carolina | [11 empty] | Display FOV (cm)
2 | [5 empty] | Inspection Results |  |  |  |  | CT System Compliance Inspection | [11 empty] | Reconstruction Algorithm
3 | [5 empty] | Measurement Protocols | [10 empty] | Measurement Protocols | [5 empty] | Axial/Helical
4 |  |  |  | Adult |  |  | Pediatric | [6 empty] | Adult |  |  | Pediatric | [5 empty] | Z-Axis collimation
5 |  |  |  | Abdomen | Chest | Head | Abdomen | Head | [5 empty] | Abdomen | Chest | Head | Abdomen | Head |  |  |  |  | # of channels
6 |  |  | kVp | [9 empty] | kVp | [9 empty] | Table Increment/speed
7 |  |  | mA | [9 empty] | mA | [9 empty] | Pitch
8 |  |  | mAs | [9 empty] | mAs | [9 empty] | Recon scan width (mm)
9 |  |  | Eff mAs | [9 empty] | Eff mAs | [9 empty] | Recon scan interval (mm)
10 |  |  | Time/rotation (s) | [9 empty] | Time/rotation (s) | [9 empty] | Dose reduction techniques
11 |  |  | Scan FOV (cm) | [9 empty] | Scan FOV (cm) | [9 empty] | Indiated CTDIvol (mGy)
12 |  |  | Display FOV (cm) | [9 empty] | Display FOV (cm) | [9 empty] | Adult Head
13 |  |  | Reconstruction Algorithm | [9 empty] | Reconstruction Algorithm | [9 empty] | kVp
14 |  |  | Axial/Helical | [9 empty] | Axial/Helical | [9 empty] | mA
15 |  |  | Z-Axis collimation | [9 empty] | Z-Axis collimation | [9 empty] | mAs
16 |  |  | # of channels | [9 empty] | # of channels | [9 empty] | Eff mAs
17 |  |  | Table Increment/speed | [9 empty] | Table Increment/speed | [9 empty] | Time/rotation (s)
18 |  |  | Pitch | [9 empty] | Pitch | [9 empty] | Scan FOV (cm)
19 |  |  | Recon scan width (mm) | [9 empty] | Recon scan width (mm) | [9 empty] | Display FOV (cm)
20 |  |  | Recon scan interval (mm) | [9 empty] | Recon scan interval (mm) | [9 empty] | Reconstruction Algorithm
21 |  |  | Dose reduction techniques | [9 empty] | Dose reduction techniques | [9 empty] | Axial/Helical
22 |  |  | Indiated CTDIvol (mGy) | [9 empty] | Indiated CTDIvol (mGy) | [9 empty] | Z-Axis collimation
23 | [13 empty] | Adult |  |  | Pediatric | [5 empty] | # of channels
24 | CTDI Measurements | [12 empty] | Abdomen | Chest | Head | Abdomen | Head |  |  |  |  | Table Increment/speed
25 |  |  |  | Adult |  |  | Pediatric | [5 empty] | kVp | [9 empty] | Pitch
26 |  |  |  | Abdomen | Chest | Head | Abdomen | Head | Criteria: |  |  |  | mA | [9 empty] | Recon scan width (mm)
27 |  |  | Indicated CTDIvol (mGy): | [5 empty] | CTDIvol is less than ACR Pass/Fail dose |  |  |  | mAs | [9 empty] | Recon scan interval (mm)
28 |  | CTDIvol Measured (mGy) | CTDIvol Measured (mGy) | [6 empty] | Reference | Pass/Fail |  | Eff mAs | [9 empty] | Dose reduction techniques
29 | [9 empty] | (mGy) | (mGy) |  | Time/rotation (s) | [9 empty] | Indiated CTDIvol (mGy)
30 | [8 empty] | Adult Abdomen | 25 | 30 |  | Scan FOV (cm) | [9 empty] | Ped Abdomen
31 |  |  | Average CTDIvol: | [5 empty] | Adult Head: | 75 | 80 |  | Display FOV (cm) | [9 empty] | kVp
32 |  |  | Phantom size indicated (cm): | [5 empty] | Ped Abd (16cm) | 15 | 20 |  | Reconstruction Algorithm | [9 empty] | mA
33 |  |  | Phantom size used (cm): | [5 empty] | Ped Abd (32cm) | 7.5 | 10 |  | Axial/Helical | [9 empty] | mAs
34 |  |  | Variation from previous year: | [5 empty] | Pediatric Head: | 35 | 40 |  | Z-Axis collimation | [9 empty] | Eff mAs
35 |  |  | Variation from indicated CTDI: | [9 empty] | # of channels | [9 empty] | Time/rotation (s)
36 |  |  | Acceptable: | [9 empty] | Table Increment/speed | [9 empty] | Scan FOV (cm)
37 |  |  | CTDIvol/mAs: | [9 empty] | Pitch | [9 empty] | Display FOV (cm)
38 |  |  | Criteria: | CTDI is within 5% of previous year. | [8 empty] | Recon scan width (mm) | [9 empty] | Reconstruction Algorithm
39 |  |  |  | Measured CTDI is within 20% of indicated CTDI | [8 empty] | Recon scan interval (mm) | [9 empty] | Axial/Helical
40 | [12 empty] | Dose reduction techniques | [9 empty] | Z-Axis collimation
41 | Beam Profile | [11 empty] | Indiated CTDIvol (mGy) | [9 empty] | # of channels
42 |  | Channels (N) | [20 empty] | Table Increment/speed
43 |  | Detector width T (mm) | [20 empty] | Pitch
44 |  | NxT Beam width (mm) | [14 empty] | Inspection Results | [5 empty] | Recon scan width (mm)
45 |  | Nominal beam width (mm) | [9 empty] | CTDI Measurements | [10 empty] | Recon scan interval (mm)
46 |  | Average beam width (mm) | [11 empty] | Adult |  |  | Pediatric |  |  | Criteria: |  |  | Dose reduction techniques
47 | [13 empty] | Abdomen | Chest | Head | Abdomen | Head |  | CTDIvol is less than ACR Pass/Fail dose |  |  | Indiated CTDIvol (mGy)
48 |  | Acceptable: | [10 empty] | Indicated CTDIvol (mGy) | [7 empty] | Reference | Pass/Fail | Ped Head
49 |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value | [7 empty] | CTDIvol Measured (mGy) | CTDIvol Measured (mGy) | [7 empty] | (mGy) | (mGy) | kVp
50 | [19 empty] | Adult Abdomen | 25 | 30 | mA
51 | [19 empty] | Adult Head: | 75 | 80 | mAs
52 | [12 empty] | Average CTDIvol: | [6 empty] | Ped Abd (16cm) | 15 | 20 | Eff mAs
53 | [12 empty] | Phantom size indicated (cm) | [6 empty] | Ped Abd (32cm) | 7.5 | 10 | Time/rotation (s)
54 | [12 empty] | Phantom size used (cm) | [6 empty] | Pediatric Head: | 35 | 40 | Scan FOV (cm)
55 | Artifact Evaluation | [11 empty] | Previous year: | [9 empty] | Display FOV (cm)
56 |  | Water phantom is free from rings, streaks, lines, cupping artifacts | [10 empty] | Variation from prev year: | [9 empty] | Reconstruction Algorithm
57 | [12 empty] | Variation from indicated: | [9 empty] | Axial/Helical
58 | Acquisition Console Monitor | [11 empty] | Acceptable: | [9 empty] | Z-Axis collimation
59 |  | SMPTE 5%/95% visible |  |  | Lmin (cd/m^2) | [7 empty] | CTDIvol/100 mAs: | [9 empty] | # of channels
60 |  | Display free of artifacts |  |  | Lmax (cd/m^2) | [7 empty] | Prev CTDIvol/100 mAs: | [9 empty] | Table Increment/speed
61 |  | Display uniformity |  |  | % Non-uniformity | [7 empty] | Criteria: | CTDI is within 5% of previous year. | [8 empty] | Pitch
62 | [13 empty] | Measured CTDI is within 20% of indicated CTDI | [8 empty] | Recon scan width (mm)
63 | [11 empty] | Beam Profile | [10 empty] | Recon scan interval (mm)
64 | [12 empty] | Channels (N) | [9 empty] | Dose reduction techniques
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Detector width T (mm) |  |  |  |  | Criteria: | Beam profile is within 3mm or 30% of specified value |  |  |  | Indiated CTDIvol (mGy)
66 | Room Number | [8 empty] | Survey ID: |  |  | NxT Beam width (mm) | [9 empty] | CTDI Measurements
1 | [10 empty] | Medical University of South Carolina |  | Nominal beam width (mm) | [9 empty] | Adult Abdomen
2 | [5 empty] | Inspection Results |  |  |  |  | CT System Compliance Inspection |  | Measured beam width (mm) | [9 empty] | Adult Chest
3 | Alignment Accuracy | [21 empty] | Adult Head
4 |  | Landmark position: | 0 | mm |  | Slice thickness: | 1 | mm | [14 empty] | Ped Abdomen
5 | [12 empty] | Average (mm): | [9 empty] | Ped Head
6 |  |  |  | Foot | Head |  |  |  | Scout prescription accuracy: |  |  |  | Acceptable: | [9 empty] | CTDIvol/100 mAs
7 |  |  | Location of alignment slice: |  |  | mm |  |  | Central wire visible: | [13 empty] | Adult Abdomen
8 |  |  | All markers visible | [5 empty] | Table movement accuracy: |  |  | Artifact Evaluation | [10 empty] | Adult Chest
9 | [8 empty] | Alignment light accuracy: |  |  |  | Water phantom is free from rings, streaks, lines, cupping artifacts | [9 empty] | Adult Head
10 | [22 empty] | Ped Abdomen
11 | CT Number Accuracy | [21 empty] | Ped Head
12 |  | kVp | 80 | 100 | 110 | 120 | 130 | 140 | 150 | [7 empty] | Inspection Results
13 |  | mAs | [9 empty] | Alignment Accuracy | [10 empty] | Comment Page 1
14 |  | Polyethylene | [10 empty] | Landmark position: | 0 | mm | Slice thickness | 1 | mm
15 |  | Water
16 |  | Water Std Dev | [11 empty] | Foot | Head |  |  |  | Enter 1 for YES, 2 for NO
17 |  | Acrylic | [10 empty] | Location of alignment slice: |  |  | mm |  | Scout prescription accuracy:
18 |  | Bone | [10 empty] | All markers visible: |  |  |  |  | Central wire visible:
19 |  | Air | [15 empty] | Table movement accuracy:
20 |  |  |  |  | Acceptable: | [12 empty] | Alignment light accuracy:
21 |  |  | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp | [7 empty] | CT Number Accuracy
22 |  |  |  | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp | [7 empty] | kVp | 80 | 100 | 110 | 120 | 130 | 140 | 150 |  | Acceptable HU
23 | [11 empty] | mAs | [8 empty] | Low | High
24 | CT Slice Thickness Accuracy | [10 empty] | Polyethylene | [7 empty] | Polyethylene | -107 | -84
25 |  |  | Slice thickness (mm): | [8 empty] | Water | [7 empty] | Water | -7 | 7
26 |  |  | Top (mm): | [8 empty] | Water Std Dev | [7 empty] | Acrylic | 110 | 135
27 |  |  | Bottom (mm): | [8 empty] | Acrylic | [7 empty] | Bone | 850 | 970
28 |  |  | Average (mm): | [8 empty] | Bone | [7 empty] | Air | -1005 | -970
29 |  |  | Water HU: | [8 empty] | Air
30 |  |  | Water HU Std Dev: | [11 empty] | Acceptable: |  | Enter 1 for YES, 2 for NO
31 |  |  | Acceptable: | [9 empty] | Critiera: | Water HU is 0 +/- 5, air HU is within acceptable range at all kVp
32 |  |  |  | Criteria: | Slice thickness is within 1.5mm of the set thickness | [8 empty] | CT number for inserts using the adult abdomen protocol is within the acceptable range at 120 or 130 kVp
33 |  |  |  |  | CT number for water is 0 +/- 5
34 | [11 empty] | CT Slice Thickness Accuracy
35 | Low Contrast Performance | [11 empty] | Slice thickness (mm):
36 |  |  | Adult |  | Pediatric | [7 empty] | Top (mm):
37 |  |  | Abdomen | Head | Abdomen | Head | [6 empty] | Bottom (mm):
38 |  | Inside | [5 empty] | 6mm rods visible: |  |  |  |  | Average (mm):
39 |  | Adjacent | [5 empty] | Acceptable: |  |  |  |  | Water HU:
40 |  | Adjacent StDev | [10 empty] | Water HU Std Dev:
41 |  | Contrast | [10 empty] | Acceptable:
42 |  | CNR | [11 empty] | Criteria: | Slice thickness is within 1.5mm of the set thickness
43 |  |  |  | Criteria: | All four 6mm rods are visible | [9 empty] | CT number for water is 0 +/- 5
44 |  |  |  |  | CNR > 1.0 for adult abdomen and head protocols
45 |  |  |  |  | CNR > 0.4 for pediatric abdomen protocol | [6 empty] | Low Contrast Performance
46 |  |  |  |  | CNR > 1.0 for pediatric head protocol | [7 empty] | Adult |  | Pediatric
47 | Uniformity | [11 empty] | Abdomen | Head | Abdomen | Head |  |  | Enter 1 for YES, 2 for NO, 3 for Not Applicable
48 |  | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable | [5 empty] | Inside | [5 empty] | 6mm rods visible: |  | Acceptable:
49 | Center: | [10 empty] | Adjacent | [5 empty] | Criteria: | All four 6mm rods are visible
50 | Top: | [10 empty] | Adj StDev | [6 empty] | CNR > 1.0 for adult abdomen and head protocols
51 | Bottom: | [10 empty] | Contrast | [6 empty] | CNR > 0.4 for pediatric abdomen protocol
52 | Right: | [10 empty] | CNR | [6 empty] | CNR > 1.0 for pediatric head protocol
53 | Left:
54 |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region | [8 empty] | Uniformity
55 |  | Uniformity image is free from artifacts | [10 empty] | Mean | St Dev | Diff (Center) | Diff (Mean) | Acceptable
56 | [11 empty] | Center:
57 | High Contrast Resolution | [10 empty] | Top:
58 |  |  | lp/cm | Acceptable | [7 empty] | Bottom:
59 |  | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen |  |  |  |  | Right:
60 |  | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution |  |  |  |  | Left:
61 | [12 empty] | Uniformity image is free from artifacts |  |  | Criteria: | CT number of all ROIs is within +/- 5 HU of the central region
62 | [11 empty] | High Contrast Resolution
63 | [12 empty] | lp/cm | Acceptable
64 | [11 empty] | Abdomen |  |  |  | Critiera: | 6 lp/cm for abdomen
65 | Date: | [8 empty] | Inspector: | Eugene Mah | Hi-res Chest |  |  |  |  | 8 lp/cm for high resolution
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina | Acquisition Console Monitor
2 | [5 empty] | Comments Page 1 |  |  |  |  | CT System Compliance Inspection |  |  |  |  | Display uniformity
3 | [12 empty] | SMPTE 5%/95% visible |  |  | Top left:
4 | [12 empty] | Display free of artifacts |  |  | Top right:
5 | [12 empty] | Display uniformity |  |  | Center:
6 | [12 empty] | Lmin (cd/m^2) |  |  | Bottom left:
7 | [12 empty] | Lmax (cd/m^2) |  |  | Bottom right:
8 | [11 empty] | Critiera: | Lmin <= 1.2 cd/m^2, Lmax >= 90 cd/m^2 |  |  | Average:
9 | [12 empty] | Display uniformity is +/-15% |  |  | SD:
10 | [15 empty] | % Non-uniformity:
11
12 | [16 empty] | Comments Page 1
13 | [15 empty] | Previous Year Comments
14 | [12 empty] | Comments:
15 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
16 | [12 empty] | Additional Comments:
17 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
18 | [12 empty] | Additional Comments:
19 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
20 | [12 empty] | Additional Comments:
21 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
22 | [12 empty] | Additional Comments:
23 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
24 | [12 empty] | Additional Comments:
25 | [12 empty] | Do not exceed 390 chars |  | 0
26 | [12 empty] | Additional Comments:
27 | [12 empty] | Do not exceed 390 chars |  | 0
28 | [12 empty] | Additional Comments:
29 | [12 empty] | Do not exceed 390 chars |  | 0
30 | [12 empty] | Additional Comments:
31 | [12 empty] | Do not exceed 390 chars |  | 0
32 | [12 empty] | Additional Comments:
33 | [12 empty] | Do not exceed 390 chars |  | 0
34 | [12 empty] | Additional Comments:
35 | [12 empty] | Do not exceed 390 chars |  | 0
36 | [12 empty] | Additional Comments:
37 | [12 empty] | Do not exceed 390 chars |  | 0
38 | [12 empty] | Additional Comments:
39 | [12 empty] | Do not exceed 390 chars |  | 0
40 | [12 empty] | Additional Comments:
41 | [12 empty] | Do not exceed 390 chars |  | 0
42 | [12 empty] | Additional Comments:
43 | [12 empty] | Do not exceed 390 chars |  | 0
44 | [12 empty] | Additional Comments:
45 | [12 empty] | Do not exceed 390 chars |  | 0
46 | [12 empty] | Additional Comments:
47 | [12 empty] | Do not exceed 390 chars |  | 0
48 | [12 empty] | Additional Comments:
49 | [12 empty] | Do not exceed 390 chars |  | 0
50 | [12 empty] | Additional Comments:
51 | [12 empty] | Do not exceed 390 chars |  | 0
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:

## Sheet2
CT Beam Profile – Siemens
NxT | Definition Flash
| Z-sharp | Nominal (mm) | Tolerance)
2x1 | Off | 3.5 | 1
1x5 | Off | 6 | 1
10x0.6 | On | 10.5 | 1.5
8x0.6 | On | 11.5 | 1.5
1x10 | Off | 11.1 | 1.5
20x0.6 | On | 18.8 | 1.5
12x1.2 | Off | 19.4 | 1.5
32x0.6 | Off | 24.4 | 2
32x0.6 | On | 27.4 | 2
32x1.2 | Off | 44 | 2
64x0.6 | On | 45.5 | 2
NxT | Definition AS
| Z-sharp | Nominal (mm) | Tolerance)
2x1 | Off | 4.7 | 1
1x5 | Off | 6.6 | 1
10x0.6 | On | 11.2 | 1
8x0.6 | Off | 8.5 | 1
1x10 | Off | 11.6 | 1
8x0.6 | On | 11.5 | 1
6x1.2 | Off | 11.8 | 1
12x1.2 | Off | 19.8 | 1
60x0.6 | Off | 41 | 1.5
64x0.6 | On | 45.5 | 1.5
32x1.2 | Off | 43.5 | 1.5
30x0.6 | Off | 23.5 | 1
32x0.6 | On | 28 | 1
20x0.6 | Off | 18 | 1
16x1.2 | Off | 25 | 1
20x0.6 | On | 20 | 1
NxT | Sensation 64
| Nominal (mm) | Tolerance)
30x0.6 | 22.5 | 1.5
24x1.2 | 33.5 | 1.5
12x1.2 | 19.5 | 1.5
1x10 | 11.5 | 1.5
1x5 | 6.5 | 1.5
2x1 | 4 | 1
6x0.6 | 7 | 1
NxT | Sensation 16
| Nominal (mm) | Tolerance)
16x1.5 | 30.5 | 1
16x0.75 | 18.5 | 1
12x1.5 | 24.5 | 1
12x0.75 | 15.5 | 1
2x5 | 12.5 | 1
2x1 | 3.5 | 1
2x0.6 | 3 | 1
NxT | Emotion 16
| Nominal (mm) | Tolerance)
4x0.6 | 4.9 | 1
12x0.6 | 12.6 | 1
16x0.6 | 14.5 | 1
2x5 | 12.6 | 1
12x1.2 | 18.5 | 1
2x8 | 18.5 | 1
16x1.2 | 24.2 | 1
NxT | Somatom Force
| Nominal (mm) | Tolerance)
2x1 | 3 | 1
6x0.5 | 7 | 1
1x5 | 7.5 | 1
1x10 | 12 | 1.5
12x1.2 | 24 | 2
32x0.6 (UHR) | 22.5 | 2
32x0.6 (sUHR) | 24.5 | 2
32x0.6 (StdHR) | 28.5 | 2
64x0.6 | 47.5 | 2
96x0.6 | 66 | 2
48x1.2 | 67 | 2