## QC Test Summary-Lorad
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Lorad
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date | 0
Medical Physicist's Name |  | Eugene Mah | [5 empty] | Signature
X-Ray Unit Manufacturer |  | Lorad/Hologic | [5 empty] | Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # |  | MAN-00093 (2007) |  |  |  |  | (use any version applicable to model; contact mfr if questions)
Accessory Equipment |  | Manufacturer |  |  | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [9 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Mammographic Unit Assembly Evaluation
2. Collimation Assessment
3. Artifact Evaluation
4. kVp Accuracy and Reproducibility
5. Beam Quality Assessment - HVL Measurement
6. Evaluation of System Resolution
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mGy
9. Radiation Output Rate
10. Phantom Image Quality Evaluation
| Phantom image scores: |  |  | Fibers | 0 | Specks | 0 | Masses | 0
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
SNR (value)
| CNR (value) |  | (Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
CNR should not vary by more than ±15% (NA for Equipment Evaluation)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
13. DICOM Printer QC (Mammography Equipment Evaluations only)
14. Detector Flat Field Calibration (Mammography Equipment Evaluations only)
15. Compression Thickness Indicator (Mammography Equipment Evaluations only)
16. Compression (Mammography Equipment Evaluations only)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on G32: Click in boxes to use drop-down lists

## QC Test Summary-Tomo
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Hologic
Site Name |  | 0 | [5 empty] | Report Date | 0
Address |  | 0 | [5 empty] | Survey Date | 0
Medical Physicist's Name |  |  | Eugene Mah |  |  |  |  | Signature
X-Ray Unit Manufacturer |  |  | Lorad/Hologic |  |  |  |  | Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # |  |  | MAN-00093 (2007) |  |  |  | (use any version applicable to model; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  | Model |  | Location |  | QC Manual Version #
|  | Review Workstation* | 0 |  | 0 |  |  |  | 0
|  | Film Printer* | 0 |  | 0 |  |  |  | 0
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type |  |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [8 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Mammographic Unit Assembly Evaluation
Autodecompression can be overridden to maintain compression (and status displayed)
Manual emergency compression release can be activated in the event of a power failure
2. Collimation Assessment
Deviation between x-ray field and light field is less than 2% SID
X-ray field dose not extend beyond any side of the image receptor by more than 2% SID
3. Artifact Evaluation
Artifacts were not apparent or not significant
4. kVp Accuracy and Reproducibility
Measured average kVp within +/-5% of indicated kV
kVp coefficient of variation <= 0.02
5. Beam Quality Assessment - HVL Measurement
Half value layer is within acceptable lower limit at all kVp values tested
6. Evaluation of System Resolution
Measured performance within acceptable limits
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
Measured performance within acceptable limits
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
Average glandular dose for average breast is ≤3 mGy (300 mrad)
|  | Average glandular dose to a 4.2 cm thick breast | [5 empty] | Conventional |  | mGy
[8 empty] | Tomosynthesis | NA | mGy
[8 empty] | Combo | NA | mGy
9. Radiation Output Rate
| Radiation output rate >= 2 mGy/s | [8 empty] | mGy/s
10. Phantom Image Quality Evaluation | [5 empty] | Fibers |  | Specks |  | Masses
|  | Phantom image scores (Conventional): |  |  |  | 0 |  | 0 |  | 0
|  | Phantom image scores (Tomosynthesis): |  |  |  | 0 |  | 0 |  | 0
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
|  | SNR (value) |  | SNR >= 40
|  | CNR (value) |  | (Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
CNR should not vary by more than ±15% (NA for Equipment Evaluation)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
White level performance
Black level performance (CRT displays only)
Quality level performance (GSDF compliance)
Uniformity performance (CRT displays only)
13. DICOM Printer QC (Mammography Equipment Evaluations only)
14. Detector Flat Field Calibration (Mammography Equipment Evaluations only)
15. Compression Thickness Indicator (Mammography Equipment Evaluations only)
16. Compression (Mammography Equipment Evaluations only)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M21: Click in boxes to use drop-down lists
Note on H46: Click in boxes to use drop-down lists

## Tech QC Eval-Tomo
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Lorad, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | DICOM Printer Quality Control | Weekly
2. | Viewboxes and Viewing Conditions | Weekly
3. | Artifact Evaluation | Weekly
4. | Signal-To-Noise and Contrast-To-Noise Measurements | Weekly
5. | Phantom Image Quality Evaluation | Weekly
6. | Detector Flat-Field Calibration | Weekly
7. | Compression Thickness Indicator | Bi-weekly
8. | Visual Checklist | Monthly
9. | Repeat/Reject Analysis | Quarterly
10. | Compression | Semi-annually
11. | Geometry Calibration (Tomosynthesis Option) | Semi-annually
12. | Diagnostic Review Workstation QC (NA if only hardcopy read) | See Lorad QC Manual
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use any QC manual version that is applicable to the Lorad model surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for Lorad FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## MQSA Requirements
MEDICAL PHYSICIST'S CHECKLIST MQSA REQUIREMENTS FOR MAMMOGRAPHY EQUIPMENT
Facility Name: | 0
Unit Manufacturer: |  |  | Model:
Serial number: |  |  | Year Mfr: | #VALUE!
Medical Physicist: | Eugene Mah |  | Room ID:
Signature: |  |  | Survey Date: | 0
Feature | FDA Rule Section | Requirement | Applies to | Meets FDA Requirements? (if NA, please explain)
Motion of tube-image receptor assembly | 3(i) | The assembly shall be capable of being fixed in any position where it is designed to operate. Once fixed in any such position, it shall not undergo unintended motion. | S-F & FFDM
| 3(ii) | This mechanism shall not fail in the event of power interruption. | S-F & FFDM
Image receptor sizes | 4(i) | Systems using screen-film image receptors shall provide, at a minimum, for operation with image receptors of 18 x 24 cm and 24 x 30 cm. | S-F
| 4(ii) | Systems using screen-film image receptors shall be equipped with moving grids matched to all image receptor sizes provided. | S-F
| 4(iii) | Systems used for magnification procedures shall be capable of operation with the grid removed from between the source and image receptor. | S-F & FFDM
Beam limitation and light fields | 5(i) | All systems shall have beam limiting devices that allow the useful beam to extend to or beyond the chest wall edge of the image receptor. | S-F & FFDM
| 5(ii) | For any mammography system with a light beam that passes through the X-ray beam-limiting device, the light shall provide an average illumination of not less than 160 lux (15 ft-candles) at 100 cm or the maximum source-image receptor distance (SID), whichever is less. | S-F & FFDM (except Fischer)
Magnification | 6(i) | Systems used to perform noninterventional problem-solving procedures shall have radiographic magnification capability available for use by the operator. | S-F & FFDM
| 6(ii) | Systems used for magnification procedures shall provide, at a minimum, at least one magnification value within the range of 1.4 to 2.0. | S-F & FFDM
Focal spot selection | 7(i) | When more than one focal spot is provided, the system shall indicate, prior to exposure, which focal spot is selected. | S-F & FFDM
| 7(ii) | When more than one target material is provided, the system shall indicate, prior to exposure, the preselected target material. | S-F & FFDM
| 7(iii) | When the target material and/or focal spot is selected by a system algorithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure. | S-F & FFDM
Application of compression | 8(i)(A) | Each system shall provide an initial power-driven compression activated by hands-free controls operable from both sides of the patient. | S-F & FFDM
| 8(i)(B) | Each system shall provide fine adjustment compression controls operable from both sides of the patient. | S-F & FFDM
Compression paddle | 8(ii)(A) | Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system. | S-F & FFDM
| 8(ii)(B) | The compression paddle shall be flat and parallel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when compression is applied. | S-F & FFDM (except Fischer)
| 8(ii)(C) | Equipment intended by the manufacturer’s design to not be flat and parallel to the breast support table during compression shall meet the manufacturer’s design specifications and maintenance requirements. | S-F & FFDM
| 8(ii)(D) | The chest wall edge of the compression paddle shall be straight and parallel to the edge of the image receptor. | S-F & FFDM
| 8(ii)(E) | The chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image. | S-F & FFDM
Technique factor selection and display | 9(i) | Manual selection of mAs or at least one of its component parts (mA and/or time) shall be available. | S-F & FFDM
| 9(ii) | The technique factors (kVp and either mA and seconds or mAs) to be used during an exposure shall be indicated before the exposure begins, except when AEC is used, in which case the technique factors that are set prior to the exposure shall be indicated. | S-F & FFDM
| 9(iii) | Following AEC mode use, the system shall indicate the actual kVp and mAs (or mA and time) used during the exposure. | S-F & FFDM
Automatic exposure control | 10(i) | Each screen-film system shall provide an AEC mode that is operable in all combinations of equipment configuration provided, e.g., grid, non-grid; magnification, nonmagnification; and various target-filter combinations. | S-F
| 10(ii) | The positioning or selection of the detector shall permit flexibility in the placement of the detector under the target tissue. The size and the available positions of the detector shall be clearly indicated at the X-ray input surface of the breast compression paddle. The selected position of the detector shall be clearly indicated. | S-F
| 10(iii) | The system shall provide means for the operator to vary the selected optical density from the normal (zero) setting. | S-F
X-ray film* | 11 | The facility shall use X-ray film for mammography that has been designated by the film manufacturer as appropriate for mammography. | S-F
Intensifying screens* | 12 | The facility shall use intensifying screens for mammography that have been designated by the screen manufacturer as appropriate for mammography and shall use film that is matched to the screen’s spectral output as specified by the manufacturer. | S-F
Film processing solutions* | 13 | For processing mammography films, the facility shall use chemical solutions that are capable of developing the films used by the facility in a manner equivalent to the minimum requirements specified by the film manufacturer. | S-F
Lighting* | 14 | The facility shall make special lights for film illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the interpreting physicians. | S-F & FFDM (for hardcopy comparison)
Film masking devices* | 15 | Facilities shall ensure that film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are available to all interpreting physicians interpreting for the facility. | S-F & FFDM (for hardcopy comparison)
* NA is acceptable for new units at existing facilities if these were previously evaluated and have not changed

## Sheet1
1 | [11 empty] | Revision 2.1-20171024 | [11 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [6 empty] | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [23 empty] | All: | Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5 | [5 empty] | Mammography System Compliance Inspection | [10 empty] | Mammography System Compliance Inspection
6 | [23 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah |  | Print Area |  | Change | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [6 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  | Inspector |  | Change | Eugene Mah | kV Set | mAs Set | Add Filt | Target | Filter | kVp | ms | Dose (mGy) | Dose rate (mGy/s) | HVL (mm Al) | Corr kV
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | 24 | 50 | 0
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | 24 | 50 | 0
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | 24 | 50 | 0.4
13 |  |  | Survey ID: |  |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | 24 | 50 | 0.4
14 | [13 empty] | Room Number: |  |  |  | Accession Number: | [5 empty] | Room Number: |  |  |  | 24 | 50 | 0.5
15 |  | X-Ray Generator | [21 empty] | Site Number: |  |  |  | 24 | 50 | 0.5
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [10 empty] | Authorized Use: |  |  |  | 24 | 50 | 0.4
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: | [5 empty] | Date of Installation: |  |  |  | 24 | 50 | 0.4
18 |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number: | [5 empty] | Number of X-ray Tubes: |  |  |  | 25 | 50 | 0
19 | [13 empty] | Max kVp: |  |  |  | Max mA: | [5 empty] | Accesssion Number: |  |  |  | 25 | 50 | 0
20 |  | X-Ray Tube 1 | [21 empty] | X-Ray Generator |  |  |  | 25 | 50 | 0.4
21 |  |  | Tube Designation/Use: |  |  |  |  | Target: |  |  |  |  | X-Ray Tube 1 |  |  |  |  | Target: | [5 empty] | Manufacturer: |  |  |  | 25 | 50 | 0.4
22 |  |  | Manufacture Date: | [10 empty] | Tube Designation/Use: | [9 empty] | Model: |  |  |  | 25 | 50 | 0.5
23 |  | Insert | [5 empty] | Filter: | [5 empty] | Manufacture Date: | [9 empty] | Max kVp: |  |  |  | 25 | 50 | 0.5
24 |  |  | Manufacturer: | [9 empty] | Insert |  |  |  |  | Filter: | [5 empty] | Manufacture Date: |  |  |  | 25 | 50 | 0.5
25 |  |  | Model: | [10 empty] | Manufacturer: | [9 empty] | Serial Number: |  |  |  | 25 | 50 | 0.5
26 |  |  | Serial Number: |  |  |  | Focal Spot Sizes (mm) | [6 empty] | Model: | [9 empty] | Max mA: |  |  |  | 26 | 50 | 0
27 |  | Housing | [5 empty] | Large: | [5 empty] | Serial Number: |  |  |  | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube 1 |  |  |  | 28 | 20 | 0
28 |  |  | Manufacturer: |  |  |  |  | Small: |  |  |  |  | Housing |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | 28 | 50 | 0
29 |  |  | Model: |  |  |  | Filtration | Inherent: | [5 empty] | Manufacturer: |  |  |  | Small: | [5 empty] | Manufacture Date: |  |  |  | 28 | 50 | 0
30 |  |  | Serial Number: |  |  |  |  | Added: | [5 empty] | Model: | [9 empty] | Manufacturer: |  |  |  | 28 | 50 | 0
31 | [13 empty] | Serial Number: |  |  |  | Filtration | [5 empty] | Model: |  |  |  | 28 | 50 | 0
32 | [17 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | 28 | 50 | 0.5
33 | [5 empty] | Technique Chart | [11 empty] | Added: | [5 empty] | Manufacturer: |  |  |  | 28 | 50 | 0.5
34 | [11 empty] | 2 | Is this a breast tomosynthesis unit | [10 empty] | Model: |  |  |  | 28 | 50 | 0.6
35 | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  | 1 | Is this unit used for screening only? | [10 empty] | Serial Number: |  |  |  | 28 | 50 | 0.6
36 | Breast | [22 empty] | Large: |  |  |  | 28 | 50 | 0.6
37 | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | [6 empty] | Inspection Results | [6 empty] | Small: |  |  |  | 28 | 50 | 0.6
38 | <3 cm | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [11 empty] | Target 1: |  |  |  | 28 | 100 | 0
39 | 3 to 5 cm | [11 empty] | DHEC Registration sticker is present, clearly visible and legible | [10 empty] | Target 2: |  |  |  | 28 | 300 | 0
40 | 5 to 7 cm | [11 empty] | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [10 empty] | Filter 1: |  |  |  | 30 | 50 | 0
41 | >7 cm | [11 empty] | Radiation warning label posted on the generator control panel | [10 empty] | Filter 2: |  |  |  | 32 | 50 | 0
42 | [12 empty] | Operator manuals are available. | [10 empty] | Filter 3: |  |  |  | 32 | 50 | 0
43 | [5 empty] | Inspection Results | [6 empty] | Monthly radiation monitoring reports are posted. | [10 empty] | Inherent: |  |  |  | 32 | 50 | 0.5
44 | Rule Number | [8 empty] | Compliance | [13 empty] | Added |  |  |  | 32 | 50 | 0.5
45 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Mammography Unit Assembly Evaluation | [6 empty] | Technique Chart |  |  |  | 32 | 50 | 0.6
46 | DHEC RHB 10.2.1 |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  |  | Free-standing unit is mechanically stable | [10 empty] | Fatty Breast |  |  |  | 32 | 50 | 0.6
47 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD |  |  | All moving parts move smoothly, without obstructions to motion | [10 empty] | Target/Filter: |  |  |  | 32 | 50 | 0.6
48 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | All locks and detents work properly | [10 empty] | kVp: |  |  |  | 32 | 50 | 0.6
49 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | Image receptor holder assembly is free from vibrations | [10 empty] | Density: |  |  |  | 34 | 50 | 0
50 | [5 empty] | Mammography Unit Assembly Evaluation | [5 empty] | 3 | Image receptor slides smoothly into holder assembly | [10 empty] | Target/Filter: |  |  |  | 34 | 50 | 0
51 |  |  | Free-standing unit is mechanically stable | [6 empty] | TBD |  |  | Image receptor is held securely by assembly in any orientation | [10 empty] | kVp: |  |  |  | 34 | 50 | 0.4
52 |  |  | All moving parts move smoothly, without obstructions to motion | [6 empty] | TBD |  |  | Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm | [10 empty] | Density: |  |  |  | 34 | 50 | 0.4
53 |  |  | All locks and detents work properly | [6 empty] | TBD |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [10 empty] | Target/Filter: |  |  |  | 34 | 50 | 0.5
54 |  |  | Image receptor holder assembly is free from vibrations | [6 empty] | TBD |  |  | Operator technique control charts are posted | [10 empty] | kVp: |  |  |  | 34 | 50 | 0.5
55 |  |  | Image receptor slides smoothly into holder assembly | [6 empty] | NA |  |  | Operator protected during exposure by adequate radiation shielding | [10 empty] | Density: |  |  |  | 34 | 50 | 0.6
56 |  |  | Image receptor is held securely by assembly in any orientation | [6 empty] | TBD |  |  | All indicator lights working properly | [10 empty] | Target/Filter: |  |  |  | 34 | 50 | 0.6
57 |  |  | Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm | [6 empty] | TBD |  |  | Auto decompression can be overridden to maintain compression (and status displayed) | [10 empty] | kVp: |  |  |  | 28 | 50 | 0
58 |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [6 empty] | TBD |  |  | Manual emergency compression release can be activated in the event of a power failure | [10 empty] | Density: |  |  |  | 28 | 50 | 0
59 |  |  | Operator technique control charts are posted | [6 empty] | TBD | [13 empty] | 50/50 Breast |  |  |  | 28 | 50 | 0.5
60 |  |  | Operator protected during exposure by adequate radiation shielding | [6 empty] | TBD | [6 empty] | MQSA Equipment Standards | [6 empty] | Target/Filter: |  |  |  | 28 | 50 | 0.5
61 |  |  | All indicator lights working properly | [6 empty] | TBD |  |  | Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position | [10 empty] | kVp: |  |  |  | 28 | 50 | 0.6
62 |  |  | Autodecompression can be overridden to maintain compression | [6 empty] | TBD |  |  | Tube-receptor assembly locking mechanism does not fail during power interruption | [10 empty] | Density: |  |  |  | 28 | 50 | 0.6
63 |  |  | Manual emergency compression release in the event of a power failure | [6 empty] | TBD |  |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [10 empty] | Target/Filter: |  |  |  | 28 | 50 | 0.6
64 | [12 empty] | Moving grids are present for all image receptor sizes | [10 empty] | kVp: |  |  |  | 28 | 50 | 0.6
65 | [12 empty] | Magnification can be done without grid | [10 empty] | Density: |  |  |  | 30 | 50 | 0
66 | [12 empty] | Beam limiting device allows useful beam to extend beyond chest wall edge of receptor | [10 empty] | Target/Filter: |  |  |  | 30 | 50 | 0
67 | Date: | [8 empty] | Inspector: | Eugene Mah | 1 | Average illumination of the light field is at least 160 lux | [10 empty] | kVp: |  |  |  | 30 | 50 | 0.5
68 | Room Number | [8 empty] | Survey ID: |  |  | Magnification mode is available for non-interventional problem solving procedures | [10 empty] | Density: |  |  |  | 30 | 50 | 0.5
1 | [10 empty] | Medical University of South Carolina |  | System provides magnification factor between 1.4 – 2.0 | [10 empty] | Target/Filter: |  |  |  | 30 | 50 | 0.6
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Prior indication of selected focal spot is provided | [10 empty] | kVp: |  |  |  | 30 | 50 | 0.6
3 | Rule Number |  |  |  |  | MQSA Equipment Standards | [6 empty] | Prior indication of selected target is provided | [10 empty] | Density: |  |  |  | 30 | 50 | 0.7
4 | 900.12(b)(3i) |  | Tube-receptor can be fixed into any operating position and does not move | [6 empty] | TBD |  |  | Post exposure display of target/focal spot is provided | [10 empty] | Dense Breast |  |  |  | 30 | 50 | 0.7
5 | 900.12(b)(3ii) |  | Tube-receptor locking mechanism does not fail during power interruption | [6 empty] | TBD |  |  | Initial power driven compression can be activated hands-free from both sides | [10 empty] | Target/Filter: |  |  |  | 32 | 50 | 0
6 | 900.12(b)(4i) |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [6 empty] | TBD |  |  | Fine adjustment of compression can be activated from both sides | [10 empty] | kVp: |  |  |  | 32 | 50 | 0
7 | 900.12(b)(4ii) |  | Moving grids are present for all image receptor sizessizes | [6 empty] | TBD |  |  | Compression paddles match the sizes of all image receptors | [10 empty] | Density: |  |  |  | 32 | 50 | 0.6
8 | 900.12(b)(4iii) |  | Magnification can be done without grid | [6 empty] | TBD |  |  | Compression paddles do not flex by more than 1 cm | [10 empty] | Target/Filter: |  |  |  | 32 | 50 | 0.6
9 | 900.12(b)(5i) |  | Beam limiting device allows useful beam to extend beyond chest wall edge | [6 empty] | TBD |  |  | Compression paddles are flat and parallel to the breast support table | [10 empty] | kVp: |  |  |  | 32 | 50 | 0.7
10 | 900.12(b)(5ii) |  | Average illumination of the light field is at least 160 lux | [6 empty] | YES |  |  | Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor | [10 empty] | Density: |  |  |  | 32 | 50 | 0.7
11 | 900.12(b)(6i) |  | Magnification mode available for non-interventional problem solving | [6 empty] | TBD |  |  | Chest wall edge may be bent upwards for patient comfort and does not appear in the image | [10 empty] | Target/Filter: |  |  |  | 32 | 50 | 0.7
12 | 900.12(b)(6ii) |  | System provides magnification factor between 1.4 – 2.0 | [6 empty] | TBD |  |  | X-ray field extends beyond the chest wall edge of the receptor | [10 empty] | kVp: |  |  |  | 32 | 50 | 0.7
13 | 900.12(b)(7i) |  | Prior indication of selected focal spot is provided | [6 empty] | TBD |  |  | Manual selection of mAs or mA and/or time is available | [10 empty] | Density: |  |  |  | 34 | 50 | 0
14 | 900.12(b)(7i) |  | Prior indication of selected target is provided | [6 empty] | TBD |  |  | Prior indication of selected technique factors is provided | [10 empty] | Target/Filter: |  |  |  | 34 | 50 | 0
15 | 900.12(b)(7iii) |  | Post exposure display of target/focal spot is provided | [6 empty] | TBD |  |  | Post exposure display of technique factors is provided for AEC exposures | [10 empty] | kVp: |  |  |  | 34 | 50 | 0.6
16 | 900.12(b)(8)(iB) |  | Initial power driven compression can be activated hands-free from both sides | [6 empty] | TBD |  |  | AEC mode is operable in all combinations of equipment configuration | [10 empty] | Density: |  |  |  | 34 | 50 | 0.6
17 |  |  | Fine adjustment of compression can be activated from both sides | [6 empty] | TBD |  |  | AEC sensor positioning is flexible | [14 empty] | 34 | 50 | 0.7
18 | 900.12(b)(8)(iiA) |  | Compression paddles match the sizes of all image receptors | [6 empty] | TBD |  |  | AEC sensor position and size is indicated on the compression paddle | [10 empty] | 2D MGD: |  |  |  | 34 | 50 | 0.7
19 | 900.12(b)(8)(iiB) |  | Compression paddles do not flex by more than 1 cm | [6 empty] | TBD |  |  | Selected position of the AEC sensor is clearly indicated | [10 empty] | mGy/mAs: |  |  |  | 34 | 50 | 0.7
20 | 900.12(b)(8)(iiC) |  | Compression paddles are flat and parallel to the breast support table | [6 empty] | TBD |  |  | AEC density control setting is provided | [10 empty] | mGy/s: |  |  |  | 34 | 50 | 0.7
21 | 900.12(b)(8)(iiD) |  | Chest wall edge of the compression paddle is straight and parallel to receptor | [6 empty] | TBD |  | 3 | X-ray film used is appropriate for mammography | [10 empty] | 3D MGD: |  |  | NA | 36 | 50 | 0
22 | 900.12(b)(8)(iiE) |  | Chest wall edge may be bent upwards and does not appear in the image | [6 empty] | TBD |  | 3 | Intensifying screens used are appropriate for mammography and matched to the film used | [10 empty] | Combo 2D MGD: |  |  | NA | 38 | 50 | 0
23 |  |  | X-ray field extends beyond the chest wall edge of the receptor | [6 empty] | TBD |  | 3 | Film processing chemicals are appropriate for the film being used | [10 empty] | Combo 3D MGD: |  |  | NA | 28 | 50 | 0
24 | 900.12(b)(9i) |  | Manual selection of mAs or mA and/or time is available | [6 empty] | TBD |  |  | Lights for film illumination i.e. hot lights are available | [10 empty] | Combo 2D+3D MGD: |  |  | NA | 28 | 50 | 0
25 | 900.12(b)(9ii) |  | Prior indication of selected technique factors is provided | [6 empty] | TBD |  |  | Film masking devices are available to all interpreting physicians | [10 empty] | System Resolution |  |  |  | 28 | 50 | 0.4
26 | 900.12(b)(9iii) |  | Post exposure display of technique factors is provided for AEC exposures | [6 empty] | TBD | [13 empty] | 0 |  |  |  | 28 | 50 | 0.4
27 | 900.12(b)(10i) |  | AEC mode is operable in all combinations of equipment configuration | [6 empty] | TBD | [6 empty] | Mammography Phototimer Technique Chart | [6 empty] | 0 |  |  |  | 28 | 50 | 0.5
28 | 900.12(b)(10ii) |  | AEC sensor positioning is flexible | [6 empty] | TBD | [13 empty] | 0 |  |  |  | 28 | 50 | 0.5
29 |  |  | AEC sensor position and size is indicated on the compression paddle | [6 empty] | TBD |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast | [8 empty] | 28 | 50 | 0.6
30 |  |  | Selected position of the AEC sensor is clearly indicated | [6 empty] | TBD |  | Breast | [11 empty] | Image Quality (Printer) |  |  |  | 28 | 50 | 0.6
31 | 900.12(b)(10iii) |  | AEC density control setting is provided | [6 empty] | TBD |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  |  | kV: |  |  | 0 | 30 | 50 | 0
32 | 900.12(b)(11) |  | X-ray film used is appropriate for mammography | [6 empty] | NA |  | <3 cm | [11 empty] | mA: |  |  | 0 | 30 | 50 | 0
33 | 900.12(b)(12) |  | Intensifying screens used are appropriate for mammography and matched to film | [6 empty] | NA |  | 3 to 5 cm | [11 empty] | Bkg Density: |  |  |  | 30 | 50 | 0.4
34 | 900.12(b)(13) |  | Film processing chemicals are appropriate for the film being used | [6 empty] | NA |  | 5 to 7 cm | [11 empty] | Inside disk: |  |  |  | 30 | 50 | 0.4
35 | 900.12(b)(14) |  | Lights for film illumination i.e. hot lights are available | [6 empty] | TBD |  | >7 cm | [11 empty] | Outside disk: |  |  |  | 30 | 50 | 0.5
36 | 900.12(b)(15) |  | Film masking devices are available to all interpreting physicians | [6 empty] | TBD | [13 empty] | Difference: |  |  |  | 30 | 50 | 0.5
37 | [5 empty] | Comments | [5 empty] | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | Fibers: |  |  |  | 30 | 50 | 0.6
38 | [11 empty] | Breast | [11 empty] | Specks: |  |  |  | 30 | 50 | 0.6
39 | [11 empty] | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  |  | Masses: |  |  |  | 32 | 50 | 0
40 | [11 empty] | <3 cm | [11 empty] | Image Quality (Acq Unit) |  |  |  | 32 | 50 | 0
41 | [11 empty] | 3 to 5 cm | [11 empty] | 2D kV: |  |  | 0 | 32 | 50 | 0.5
42 | [11 empty] | 5 to 7 cm | [11 empty] | mA: |  |  | 0 | 32 | 50 | 0.5
43 | [11 empty] | >7 cm | [11 empty] | Exp Ind: |  |  | 0 | 32 | 50 | 0.6
44 | [23 empty] | Fibers: |  |  |  | 32 | 50 | 0.6
45 | [16 empty] | Inspection Results | [6 empty] | Specks: |  |  |  | 32 | 50 | 0.7
46 | [11 empty] | Compression Force Indicator | [11 empty] | Masses: |  |  |  | 32 | 50 | 0.7
47 | [13 empty] | Indicated | Measured | After 20s | Slip | [6 empty] | Mag kV: |  |  |  | 34 | 50 | 0
48 | [12 empty] | Power | [10 empty] | mA: |  |  |  | 34 | 50 | 0
49 | [12 empty] | Manual | [10 empty] | Exp Ind: |  |  |  | 34 | 50 | 0.5
50 | [12 empty] | Criteria: | Minimum compression force – 25 lbs | [9 empty] | Fibers: |  |  |  | 34 | 50 | 0.5
51 | [13 empty] | Maximum compression force – 45 lbs | [9 empty] | Specks: |  |  |  | 34 | 50 | 0.6
52 | [23 empty] | Masses: |  |  |  | 34 | 50 | 0.6
53 | [11 empty] | Compression Thickness Indicator | [11 empty] | 3D kV: |  |  |  | 34 | 50 | 0.7
54 | [12 empty] | BR-12 (cm) | Indicated | Difference | [8 empty] | mA: |  |  |  | 34 | 50 | 0.7
55 | [12 empty] | 1 | [10 empty] | Exp Ind: |  |  |  | 38 | 50 | 0
56 | [12 empty] | 2 | [10 empty] | Fibers: |  |  |  | 38 | 50 | 0
57 | [12 empty] | 4 |  |  | Mean: | [7 empty] | Specks: |  |  |  | 38 | 50 | 0.6
58 | [12 empty] | 4 |  |  | StDev: | [7 empty] | Masses: |  |  |  | 38 | 50 | 0.6
59 | [12 empty] | 4 | [14 empty] | 38 | 50 | 0.7
60 | [12 empty] | 4 |  |  |  | Criteria: | Difference between indicated and actual thickness is ± 5mm | [5 empty] | SNR: |  |  |  | 38 | 50 | 0.7
61 | [12 empty] | 6 | [10 empty] | CNR: |  |  |  | 38 | 50 | 0.8
62 | [12 empty] | 8 | [14 empty] | 38 | 50 | 0.8
63 | [23 empty] | Comment Page 1
64 | [12 empty] | Paddle flex is less than 1 cm |  |  |  | Compression thickness indicator:
65 | [12 empty] | Overall unit assembly
66
67 | Date: | [8 empty] | Inspector: | Eugene Mah | Artifact Evaluation
68 | Room Number | [8 empty] | Survey ID: |  |  | Target:
1 | [10 empty] | Medical University of South Carolina |  | Filter:
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Focal spot:
3 | Compression Force Indicator | [11 empty] | kV:
4 |  |  | Indicated | Measured | After 20s | Slip | [6 empty] | mAs:
5 |  | Power | [10 empty] | EI:
6 |  | Manual | [5 empty] | Acceptable: |  |  |  |  | Acceptable:
7 |  | Criteria: | Minimum compression force – 25 lbs
8 |  |  | Maximum compression force – 45 lbs
9 | [12 empty] | Comments:
10 | Compression Thickness Indicator | [11 empty] | Do not exceed 390 chars
11 |  | BR-12 (cm) | Indicated | Difference |  |  |  | Paddle flex is less than 1 cm: |  |  |  |  | New:
12 |  | 1 | [5 empty] | Overall unit assembly:
13 |  | 2 | [5 empty] | Compression thickness indicator: |  |  |  | Light Field Brightness |  |  |  | Enter 1 for YES, 2 for NO, 3 for NA
14 |  | 4 |  |  | Mean: | [7 empty] | Measurements (lux) |  |  |  |  | Average Illumination
15 |  | 4 |  |  | StDev: | [7 empty] | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux)
16 |  | 4 |  |  | Criteria: | Difference between indicated and actual thickness is ± 5mm | [14 empty] | Acceptable:
17 |  | 4
18 |  | 6 | [9 empty] | Light Field/X-ray Field Alignment
19 |  | 8 | [9 empty] | SID (cm): | 65
20 | [12 empty] | 24x30 | 18x24 |  |  | 18x24 | Small breast
21 | Artifact Evaluation | [11 empty] | 24x30 | Left | Center | Right | Tomo | 24x30
22 |  | Receptor: | [9 empty] | Left:
23 |  | Target: | [9 empty] | Right:
24 |  | Filter: | [9 empty] | Anterior:
25 |  | Focal spot: | [9 empty] | Chest Wall:
26 |  | kV: | [11 empty] | Enter deviation in cm
27 |  | mAs:
28 |  | EI: | [10 empty] | % Deviation
29 |  | Acceptable: | [10 empty] | 24x30 | 18x24 |  |  | 18x24 | Small breast
30 |  | Criteria: | No visible artifacts in the images | [9 empty] | 24x30 | Left | Center | Right | Tomo | 24x30
31 |  | Comments: | [9 empty] | Left:
32 | [11 empty] | Right:
33 | [11 empty] | Anterior:
34 | [11 empty] | Chest Wall:
35 | Light Field Brightness | [10 empty] | Left-Right
36 |  | Measurements (lux) |  |  |  |  | Average Illumination |  |  |  |  | Ant-Chest
37 |  | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux) | [5 empty] | Criteria: | Sum of left and right edge deviation is less than 2% SID
38 | [9 empty] | Acceptable: |  |  |  | Sum of anterior and chest wall edge deviation is less than 2% SID
39 |  | Criteria: | Average light field illumination >= 160 lux. | [10 empty] | Enter deviation in cm. Positive deviation indicates x-ray field extends beyond light field
40 | [11 empty] | X-ray Field/Image Receptor Alignment
41 | Light fField/X-ray Field Alignment |  |  |  | X-ray Field/Image Receptor Alignment
42 | SID (cm): | [11 empty] | 24x30 | 18x24 |  |  | 24x30 | Small breast
43 | 65 | Light Field/X-ray Field % Deviation | [10 empty] | 24x30 | Left | Center | Right | Tomo | 24x30
44 |  | 24x30 | 18x24 |  |  | 24x30 | Small breast |  |  |  |  | Left:
45 |  | 24x30 | Left | Center | Right | Tomo | Tomo |  |  |  |  | Right:
46 | Left: | [10 empty] | Anterior:
47 | Right: | [10 empty] | Chest Wall:
48 | Anterior: | [10 empty] | Paddle:
49 | Chest Wall: | [12 empty] | Enter deviation in cm. Positive deviation indicates x-ray field extends beyond image receptor
50 | Left-Right:
51 | Ant-Chest: | [11 empty] | % Deviation
52 | Acceptable: | [11 empty] | 24x30 | 18x24 |  |  | 24x30 | Small breast
53 |  | X-ray Field/Image Receptor % Deviation | [10 empty] | 24x30 | Left | Center | Right | Tomo
54 | Left: | [10 empty] | Left:
55 | Right: | [10 empty] | Right:
56 | Anterior: | [10 empty] | Anterior:
57 | Chest Wall: | [10 empty] | Chest Wall:
58 | Paddle: | [10 empty] | Paddle:
59 | Acceptable: | [11 empty] | X-ray field extends beyond the chest wall edge of the receptor
60 | Criteria: | Sum of left/right and chest wall/anterior deviation is less than 2% SID | [10 empty] | Chest wall edge of the compression paddle is not visible in the image
61 |  | X-ray field does not extend beyond image receptor by more than 2% SID
62 |  | X-ray field extends beyond the chest wall edge of image receptor. | [10 empty] | Criteria: | X-ray field does not extend beyond image receptor by more than 2% SID
63 |  | X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side | [11 empty] | X-ray field extends beyond the chest wall edge of image receptor.
64 |  | X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side | [11 empty] | X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side
65 |  | Difference between chest wall edge of the receptor and compression paddle is less than 1% SID | [11 empty] | X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side
66 | [13 empty] | Difference between chest wall edge of the receptor and compression paddle is less than 1% SID
67 | Date: | [8 empty] | Inspector: | Eugene Mah
68 | Room Number | [8 empty] | Survey ID: |  | System Resolution |  | Large | Small | Large | Small
1 | [10 empty] | Medical University of South Carolina |  | Mode (2D/3D):
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Target:
3 | AEC Thickness Tracking – 2D | [11 empty] | kV:
4 | AEC Mode: |  |  | AEC Position: | [8 empty] | mAs:
5 |  | Thickness |  |  |  |  | Reported | [5 empty] | lp/mm:
6 |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | Acceptable:
7 |  | 2 | [10 empty] | Previous:
8 |  | 4
9 |  | 4 | 0 | 0 | [8 empty] | Criteria: | Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode
10 |  | 4 | 0 | 0 | [9 empty] | Limiting system resolution must be 3 lp/mm or higher in 3D mode
11 |  | 4 | 0 | 0 | [7 empty] | AEC Thickness Tracking – 2D
12 |  | 6 | [9 empty] | AEC Mode: |  |  | AEC Position: |  |  | Dose Table: |  | Enter -1 if not applicable
13 |  | 8 | [10 empty] | Thickness |  |  |  |  | Corrected | Reported
14 |  | 4 | [10 empty] | (cm) | Filter | kV | mAs | EI | EI | MGD (mGy) | Deviation
15 |  |  |  |  | Average: | [7 empty] | 2
16 |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [9 empty] | 4
17 | [12 empty] | 4 | 0 | 0
18 | AEC Thickness Tracking – 3D | [11 empty] | 4 | 0 | 0
19 | AEC Mode: |  |  | AEC Position: | [8 empty] | 4 | 0 | 0
20 |  | Thickness |  |  |  |  | Reported | [5 empty] | 6
21 |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | 8
22 |  | 2 | [9 empty] | Mag: | 4
23 |  | 4 | [14 empty] | Average:
24 |  | 4 | 0 | 0 | [8 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
25 |  | 4 | 0 | 0
26 |  | 4 | 0 | 0 | [7 empty] | AEC Thickness Tracking – 3D
27 |  | 6 | [9 empty] | AEC Mode: |  |  | AEC Position: |  |  | Dose Table:
28 |  | 8 | [10 empty] | Thickness |  |  |  |  | Corrected | Reported
29 |  |  |  |  | Average: | [7 empty] | (cm) | Filter | kV | mAs | EI | EI | MGD (mGy) | Deviation
30 |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [9 empty] | 2
31 | [12 empty] | 4
32 | AEC Density Control | [11 empty] | 4 | 0 | 0
33 | AEC Mode: | 0 |  | AEC Position: | 0 |  | kV: | [5 empty] | 4 | 0 | 0
34 |  |  |  |  | Reported | [7 empty] | 4 | 0 | 0
35 |  | Density | mAs | EI | MGD (mGy) | Change | Acceptable | [5 empty] | 6
36 |  | -3 | [10 empty] | 8
37 |  | -2 | [14 empty] | Average:
38 |  | -1 | [10 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
39 |  | 0
40 |  | 1
41 |  | 2 | [9 empty] | AEC Density Control
42 |  | 3 | [9 empty] | AEC Mode: | 0 |  | AEC Position: | 0 | Reported |  |  |  | Selenia |  | Selenia |  | Selenia Dimensions
43 |  | 4 | [10 empty] | Density | kV | mAs | EI | MGD (mGy) | Change |  |  | Low | High | Low | High | Low | High
44 |  | Criteria: | Change in pixel value at each density step is within specified ranges | [9 empty] | -3 | [7 empty] | 0.5 | 0.61 | 0.5 | 0.61 | 0.56 | 0.66
45 | [12 empty] | -2 | [7 empty] | 0.63 | 0.77 | 0.63 | 0.77 | 0.66 | 0.78
46 | System Resolution |  | Large | Small | Large | Small | [6 empty] | -1 | [7 empty] | 0.77 | 0.94 | 0.77 | 0.94 | 0.78 | 0.92
47 |  | Mode (2D/3D): | [10 empty] | 0
48 |  | Target: | [10 empty] | 1 | [7 empty] | 1.04 | 1.27 | 1.04 | 1.27 | 1.06 | 1.24
49 |  | kV: | [5 empty] | Acceptable: |  |  |  |  | 2 | [7 empty] | 1.17 | 1.43 | 1.17 | 1.43 | 1.22 | 1.43
50 |  | mAs: | [10 empty] | 3 | [7 empty] | 1.31 | 1.6 | 1.31 | 1.6 | 1.4 | 1.64
51 |  | lp/mm: | [10 empty] | 4 | [7 empty] | 1.44 | 1.76 | 1.44 | 1.76 | 1.61 | 1.89
52 |  | Criteria: | Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode
53 | [12 empty] | Criteria: | Change in pixel value at each density step is within specified ranges
54 | Image Quality Evaluation
55 |  | Printer |  |  | Acq Workstation |  |  |  | Review Workstation |  |  | Mean Glandular Dose – 2D
56 |  | kV: | 0 |  | Normal | Mag | Tomo |  | Left | Right |  |  | Phantom: |  |  | kV:
57 |  | mA: | 0 | kV: | 0 |  |  |  | Workstation: |  |  |  | Target/Filter: |  |  | mA:
58 |  | Background Density: |  | mA: | 0 |  |  | Fibers: | [7 empty] | Reported |  | Air KERMA
59 |  | Inside disk: |  | EI: | 0 |  |  | Specks: | [5 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
60 |  | Outside disk: |  | Fibers: |  |  |  | Masses: | [11 empty] | ESE (mGy):
61 |  | Difference: |  | Specks: |  |  |  |  | Workstation: | [10 empty] | DGN (mrad/R):
62 |  | Fibers: |  | Masses: |  |  |  | Fibers: | [11 empty] | MGD (mGy):
63 |  | Specks: |  | Criteria: |  |  |  | Specks: | [11 empty] | Prev MGD (mGy):
64 |  | Masses: |  | 2D – Must be able to see 5 fibers, 4 speck groups, 4 masses |  |  |  | Masses: |  |  |  |  | Average: | [6 empty] | % Change:
65 |  |  |  | 3D – Must be able to see 4 fibers, 3 speck groups, 3 masses | [8 empty] | Coeff of Var:
66
67 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
68 | Room Number | [8 empty] | Survey ID: |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
1 | [10 empty] | Medical University of South Carolina
2 | [10 empty] | Mammography System Compliance Inspection | Mean Glandular Dose – 3D
3 | Mean Glandular Dose – 2D | [11 empty] | Phantom: | 0 |  | kV:
4 |  | Phantom: |  |  | kV: | [7 empty] | Target/Filter: |  |  | mA:
5 |  | Target/Filter: |  |  | mA: | [10 empty] | Reported |  | Air KERMA
6 |  |  |  |  | Reported |  | Air KERMA | [6 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
7 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | [10 empty] | ESE (mGy):
8 | [8 empty] | ESE (mGy): | [10 empty] | DGN (mrad/R): | #N/A
9 | [8 empty] | DGN (mrad/R): | [10 empty] | MGD (mGy): | NA
10 | [8 empty] | MGD (mGy): | [10 empty] | Prev MGD (mGy):
11 | [8 empty] | Prev MGD (mGy): |  |  |  | Average: | [6 empty] | % Change:
12 |  | Average: | [6 empty] | % Change: |  |  |  | Coeff of Var:
13 |  | Coeff of Var:
14 |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
15 |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: |  |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
16 | [13 empty] | MGD for 2D + 3D acquisitions is less than 3 mGy
17 | Mean Glandular Dose – 3D
18 |  | Phantom: | 0 |  | kV:
19 |  | Target/Filter: |  |  | mA: | [6 empty] | Mean Glandular Dose – Combo
20 |  |  |  |  | Reported |  | Air KERMA | [5 empty] | Phantom: | 0
21 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
22 | [8 empty] | ESE (mGy): |  |  | Combo Mode 2D
23 | [8 empty] | DGN (mrad/R): | #N/A |  |  | Target/Filter: |  |  | kV:
24 | [8 empty] | MGD (mGy): | NA | [5 empty] | mA:
25 | [8 empty] | Prev MGD (mGy): | [6 empty] | Reported |  | Air KERMA
26 |  | Average: | [6 empty] | % Change: |  |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
27 |  | Coeff of Var: | [17 empty] | ESE (mGy):
28 |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | [10 empty] | DGN (mrad/R):
29 |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | [10 empty] | MGD (mGy): | NA
30 | [19 empty] | Prev MGD (mGy):
31 | Mean Glandular Dose – Combo | [11 empty] | Average: | [6 empty] | % Change:
32 |  | Phantom: | 0 | [9 empty] | Coeff of Var:
33 |  |  | Combo Mode 2D Target/Filter: | [16 empty] | Deviation from Reported MGD: | #VALUE!
34 |  |  |  |  | Reported |  | Air KERMA |  | kV: |  |  | Combo Mode 3D
35 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  | Target/Filter: |  |  | kV:
36 | [8 empty] | ESE (mGy): | [6 empty] | mA:
37 | [8 empty] | DGN (mrad/R): | [6 empty] | Reported |  | Air KERMA
38 | [8 empty] | MGD (mGy): | NA |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
39 | [8 empty] | Prev MGD (mGy): | [10 empty] | ESE (mGy):
40 |  | Average: | [6 empty] | % Change: | [10 empty] | DGN (mrad/R):
41 |  | Coeff of Var: | [17 empty] | MGD (mGy): | NA
42 | [8 empty] | Deviation from Reported MGD: | #VALUE! | [9 empty] | Prev MGD (mGy):
43 | [12 empty] | Average: | [6 empty] | % Change:
44 |  |  | Combo Mode 3D Target/Filter: | [9 empty] | Coeff of Var:
45 |  |  |  |  | Reported |  | Air KERMA |  | kV: | [10 empty] | Deviation from Reported MGD: | #VALUE!
46 |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | [10 empty] | Combo MGD (2D + 3D): | NA
47 | [8 empty] | ESE (mGy): | [10 empty] | Prev Combo MGD:
48 | [8 empty] | DGN (mrad/R): |  |  | kVp Accuracy and Output
49 | [8 empty] | MGD (mGy): | NA |  | Meter: | Piranha |  |  | Calibration Date: | 2017-10-06
50 | [8 empty] | Prev MGD (mGy): |  |  | Serial #: | CB2-17090320 |  |  | Calibration Due: | 2019-10-06
51 |  | Average: | [6 empty] | % Change:
52 |  | Coeff of Var: | [14 empty] | Measured Values
53 |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | #VALUE! |  |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
54 |  |  | Deviation from reported MGD is less than15% | [11 empty] | 24 | 50
55 | [8 empty] | Combo MGD: | NA |  |  |  |  | 25 | 50
56 | [8 empty] | Prev Combo MGD: | [5 empty] | 26 | 50
57 | [14 empty] | 28 | 50
58 | [14 empty] | 30 | 50
59 | Signal to Noise Ratio/Contrast to Noise Ratio | [13 empty] | 32 | 50
60 |  | kV: | 0 |  |  | Current | Previous | Change | Acceptable | [5 empty] | 34 | 50
61 |  | mAs: | 0 |  | SNR: |  |  |  | Pass |  |  |  | Criteria: | kV accuracy is <= 5%
62 |  | Target: |  |  | CNR: |  |  |  | NA
63 |  | Filter: | [14 empty] | Measured Values
64 |  | Criteria: | SNR must be >= 40 | [9 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
65 |  |  | CNR does not change by more than 15% | [11 empty] | 28 | 50
66 | [14 empty] | 30 | 50
67 | Date: | [8 empty] | Inspector: | Eugene Mah |  |  |  | 32 | 50
68 | Room Number | [8 empty] | Survey ID: |  |  |  |  | 34 | 50
1 | [10 empty] | Medical University of South Carolina |  |  |  | 36 | 50
2 | [5 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  |  |  | 38 | 50
3 | Meter: | Piranha |  |  |  | Calibration Date: | 2017-10-06 | [5 empty] | Criteria: | kV accuracy is <= 5%
4 | SN: | CB2-17090320 |  |  |  | Calibration Due: | 2019-10-06
5 | kVp Accuracy | [15 empty] | Measured Values
6 | Target: |  |  |  | Target: |  |  |  | Target: |  |  |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
7 | Filter: |  |  |  | Filter: |  |  |  | Filter: | [5 empty] | 28 | 50
8 | mAs: | 50 |  |  | mAs: | 50 |  |  | mAs: | 50 |  |  |  |  | 30 | 50
9 | Indicated | Measured |  |  | Indicated | Measured |  |  | Indicated | Measured |  |  |  |  | 32 | 50
10 | kVp | kVp | Error (%) |  | kVp | kVp | Error (%) |  | kVp | kVp | Error (%) |  |  |  | 34 | 50
11 | 24 |  |  |  | 28 |  |  |  | 28 | [5 empty] | 38 | 50
12 | 25 |  |  |  | 30 |  |  |  | 30 |  |  |  | Criteria: | kV accuracy is <= 5%
13 | 26 |  |  |  | 32 |  |  |  | 32
14 | 28 |  |  |  | 34 |  |  |  | 34 |  |  | kV, Exposure, and Output Reproducibility
15 | 30 |  |  |  | 36 |  |  |  | 38
16 | 32 |  |  |  | 38 | [11 empty] | Measured Values
17 | 34 | [11 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
18 |  | Acceptable: |  |  |  | Acceptable: |  |  |  | Acceptable: |  |  |  |  | 28 | 50
19 | [14 empty] | 28 | 50
20 |  | Criteria: | kVp accurate to within 5% of indicated | [11 empty] | 28 | 50
21 | [14 empty] | 28 | 50
22 | kVp and Output Reproducibility | [14 empty] | Mean:
23 | Target: |  | Filter: | [12 empty] | Std Dev:
24 | mAs: | 50 | [13 empty] | Coeff of Var:
25 | Indicated | Measured | [13 empty] | Previous:
26 | kVp | kVp | mGy | mGy/mAs | mGy/s | [7 empty] | Criteria: | Coefficient of variation <= 2%
27 | 28 | [12 empty] | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
28 | [13 empty] | Minimum acceptable output is 2 mGy/s at 28 kVp for W target
29 | [11 empty] | Exposure Linearity
30 | [16 empty] | Measured Values
31 | Mean: | [11 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
32 | Std Dev: | [13 empty] | 28 | 20
33 | Coeff of Var: | [13 empty] | 28 | 50
34 | Acceptable: | [13 empty] | 28 | 100
35 |  | Criteria: | Coefficient of variation <= 2% | [11 empty] | 28 | 300
36 |  |  | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target | [9 empty] | Criteria: | Coefficient of linearity is <= 10% |  |  |  |  | Coeff of Linearity:
37 |  |  | Minimum acceptable output is 2 mGy/s at 28 kVp for W target
38 | [12 empty] | Unit Can Maintain Output for 3 sec at 28 kVp
39 | Exposure Linearity
40 | Target: |  | Filter: | [8 empty] | Half Value Layer – 2D |  | 1 | Compression paddle used?
41 | [12 empty] | Target/Filter | / | / | / | / | / | / | / | /
42 | Set mAs | kVp | mGy | mGy/mAs | mGy/s | [7 empty] | kV | 24 | 25 | 28 | 32 | 28 | 30 | 32 | 34
43 | 20 | [11 empty] | Al (mm) | Air Kerma (mGy)
44 | 50 | [11 empty] | 0
45 | 100 | [11 empty] | 0
46 | 300 | [11 empty] | HVL
47 |  |  | Coeff of Linearity: | [9 empty] | 0 mm
48 |  | Criteria: | Coefficient of linearity is <= 10% | [9 empty] | Min | 0.27 | 0.28 | 0.31 | 0.35 | 0.31 | 0.33 | 0.35 | 0.37
49 | [12 empty] | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.47 | 0.49 | 0.51 | 0.53
50 | Half Value Layer
51 | Target/Filter | / | / | / | / | / | / | / | /
52 | kV | 24 | 25 | 28 | 32 | 28 | 30 | 32 | 34 |  |  |  | Criteria: | HVL is between the minimum and maximum limits
53 | HVL | [10 empty] | Half Value Layer – 3D
54 | 0 mm | [11 empty] | Target/Filter | / | / | / | / | /
55 | Min | 0.27 | 0.28 | 0.31 | 0.35 | 0.31 | 0.33 | 0.35 | 0.37 |  |  |  | kV | 28 | 30 | 32 | 34 | 38
56 | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.47 | 0.49 | 0.51 | 0.53 |  |  |  | Al (mm) | Air Kerma (mGy)
57 | Acceptable: | [11 empty] | 0
58 | [12 empty] | 0
59 | Target/Filter | / | / | / | / | / | [6 empty] | HVL
60 | kV | 28 | 30 | 32 | 34 | 38 | [6 empty] | 0 mm
61 | HVL | [11 empty] | Min | 0.31 | 0.33 | 0.35 | 0.37 | 0.41
62 | 0 mm | [11 empty] | Max | 0.4 | 0.42 | 0.44 | 0.46 | 0.57
63 | Min | 0.31 | 0.33 | 0.35 | 0.37 | 0.41
64 | Max | 0.4 | 0.42 | 0.44 | 0.46 | 0.57 | [6 empty] | Criteria: | Mo/Rh targets – HVL is between the minimum and maximum limits
65 | Acceptable: | NA | NA | NA | NA | NA | [7 empty] | W targets – HVL is above the minimum limit
66 |  | Criteria: | Mo/Rh targets – HVL is between the minimum and maximum limits | [8 empty] | Image Quality Evaluation
67 | Date: | [8 empty] | Inspector: | Eugene Mah | Printer | Model: |  | SN: |  |  | QC Manual Version:
68 | Room Number | [8 empty] | Survey ID:
[15 empty] | Previous |  |  |  | Previous
[13 empty] | kV: | 0 |  |  | Fibers:
[13 empty] | mAs: | 0 |  |  | Specks:
[13 empty] | Background Density: |  |  |  | Masses:
[13 empty] | Inside disk: |  |  |  | Acceptable:
Outside disk:
Difference:
Workstation
[14 empty] | Workstation: |  |  |  | Workstation:
[14 empty] | Left | Right |  |  | Left | Right
[13 empty] | Fibers: |  |  |  | Fibers:
[13 empty] | Specks: |  |  |  | Specks:
[13 empty] | Masses: |  |  |  | Masses:
Acquisition Unit
[13 empty] | 2D | Previous | Mag | Previous | 3D | Previous
[12 empty] | kV: | 0
[12 empty] | mAs: | 0
[12 empty] | EI: | 0
Fibers:
Specks:
Masses:
[13 empty] | Criteria: | 2D – Must be able to see 5 fibers, 4 speck groups, 4 masses
3D – Must be able to see 4 fibers, 3 speck groups, 3 masses
Signal to Noise
[17 empty] | Mean | SD
[13 empty] | kV: | 0 |  | In acrylic:
[13 empty] | mAs: | 0 |  | Background: |  |  | 8.03
Target:
[13 empty] | Filter: |  |  |  | Current | Previous | Change | Acceptable
[16 empty] | SNR: |  |  |  | Pass
[16 empty] | CNR: |  |  |  | NA
[13 empty] | Criteria: | SNR must be >= 40
CNR does not change by more than 15%
Previous Year Comments
Comments:
[13 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[13 empty] | Do not exceed 390 chars |  | 0

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
| kVp | [9 empty] | kVp | [9 empty] | kVp
HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
0.23 | 116 | [8 empty] | 0.23 | [9 empty] | 0.23
0.24 | 121 | 124 | [7 empty] | 0.24 | [9 empty] | 0.24
0.25 | 126 | 129 | 131 | [6 empty] | 0.25 | [9 empty] | 0.25
0.26 | 130 | 133 | 135 | 138 | [5 empty] | 0.26 | [9 empty] | 0.26
0.27 | 135 | 138 | 140 | 142 | 143 |  |  |  |  | 0.27 | [9 empty] | 0.27
0.28 | 140 | 142 | 144 | 146 | 147 | 149 |  |  |  | 0.28 |  |  | 149 | 151 | 154 |  |  |  |  | 0.28 |  |  | 150 | 155 | 159
0.29 | 144 | 146 | 148 | 150 | 151 | 153 | 154 |  |  | 0.29 |  |  | 154 | 156 | 158 | 159 |  |  |  | 0.29 |  |  | 155 | 160 | 164 | 168
0.3 | 149 | 151 | 153 | 155 | 156 | 157 | 158 | 159 |  | 0.3 |  |  | 158 | 160 | 162 | 162 | 163 |  |  | 0.3 |  |  | 160 | 164 | 168 | 172 | 176
0.31 | 154 | 156 | 157 | 159 | 160 | 161 | 162 | 163 | 164 | 0.31 |  |  | 163 | 164 | 166 | 166 | 167 | 167 |  | 0.31 |  |  | 165 | 168 | 172 | 174 | 180 | 182
0.32 | 158 | 160 | 162 | 163 | 164 | 166 | 167 | 168 | 168 | 0.32 |  |  | 167 | 169 | 171 | 171 | 171 | 172 | 172 | 0.32 |  |  | 169 | 173 | 177 | 181 | 184 | 186 | 188
0.33 | 163 | 165 | 166 | 168 | 169 | 170 | 171 | 173 | 173 | 0.33 |  |  | 171 | 173 | 175 | 176 | 176 | 176 | 176 | 0.33 |  |  | 174 | 178 | 181 | 185 | 188 | 190 | 192
0.34 | 168 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 0.34 |  |  | 176 | 178 | 179 | 179 | 180 | 180 | 180 | 0.34 |  |  | 179 | 183 | 186 | 190 | 193 | 195 | 196
0.35 |  | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 0.35 |  |  | 180 | 181 | 183 | 183 | 184 | 185 | 185 | 0.35 |  |  | 184 | 187 | 190 | 194 | 197 | 199 | 201
0.36 |  |  | 179 | 181 | 182 | 183 | 184 | 185 | 185 | 0.36 |  |  | 185 | 186 | 187 | 187 | 188 | 188 | 189 | 0.36 |  |  | 189 | 192 | 195 | 198 | 201 | 204 | 205
0.37 |  |  |  | 185 | 186 | 187 | 188 | 189 | 190 | 0.37 |  |  | 189 | 190 | 191 | 191 | 192 | 193 | 193 | 0.37 |  |  | 193 | 196 | 199 | 202 | 205 | 207 | 209
0.38 |  |  |  |  | 190 | 191 | 192 | 193 | 194 | 0.38 |  |  | 193 | 194 | 196 | 196 | 197 | 197 | 197 | 0.38 |  |  | 198 | 201 | 204 | 207 | 209 | 211 | 213
0.39 |  | 174 178: 196; 175 179: 197; 176 180: 198; 177 181: 198; 0.34 0.35: 0.39 |  |  | 198 | 199 | 200 | 200 | 201 | 201 | 202 | 0.39 |  |  | 203 | 206 | 208 | 211 | 214 | 216 | 217
0.4 |  | 175 179: 201; 176 180: 202; 177 181: 203; 0.34 0.35: 0.4 |  |  | 202 | 203 | 204 | 204 | 205 | 205 | 206 | 0.4 |  |  | 208 | 211 | 213 | 216 | 218 | 220 | 221
0.41 |  | 176 180: 206; 177 181: 207; 0.34 0.35: 0.41 |  |  | 206 | 207 | 208 | 208 | 209 | 209 | 210 | 0.41 |  |  | 213 | 215 | 217 | 220 | 222 | 224 | 225
0.42 |  | 177 181: 211; 0.34 0.35: 0.42 |  |  | 211 | 211 | 212 | 212 | 213 | 213 | 214 | 0.42 |  |  | 218 | 220 | 222 | 224 | 226 | 228 | 229
Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast | [12 empty] | Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
| kVp | [12 empty] | kVp
HVL | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | HVL | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39
0.3 | 152 | 157 | 163 | 166 | 170 | 173 | 175 | 177 | 179 | 182 | 184 | 187 | 0.4 | 222 | 226 | 229 | 231 | 234 | 236 | 239 | 241 | 244 | 246 | 248 | 250 | 252
0.325 | 163 | 169 | 174 | 177 | 181 | 183 | 186 | 188 | 190 | 192 | 195 | 197 | 0.425 | 233 | 236 | 239 | 242 | 244 | 246 | 248 | 251 | 253 | 256 | 258 | 260 | 262
0.35 | 175 | 180 | 185 | 188 | 191 | 194 | 196 | 198 | 200 | 202 | 205 | 207 | 0.45 | 244 | 247 | 249 | 252 | 254 | 256 | 258 | 260 | 263 | 265 | 267 | 269 | 271
0.375 | 186 | 191 | 196 | 199 | 202 | 205 | 207 | 209 | 211 | 213 | 215 | 218 | 0.475 | 254 | 257 | 260 | 262 | 264 | 266 | 268 | 270 | 273 | 275 | 277 | 279 | 281
0.4 | 198 | 203 | 207 | 210 | 213 | 215 | 217 | 219 | 221 | 223 | 226 | 228 | 0.5 | 265 | 267 | 270 | 272 | 274 | 276 | 278 | 280 | 282 | 284 | 286 | 288 | 290
0.425 | 209 | 214 | 218 | 221 | 224 | 226 | 228 | 230 | 232 | 234 | 236 | 238 | 0.525 | 275 | 278 | 280 | 282 | 284 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 300
0.45 | 221 | 226 | 230 | 232 | 235 | 237 | 238 | 240 | 242 | 244 | 246 | 248 | 0.55 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 299 | 301 | 303 | 305 | 307 | 309
0.475 | 233 | 237 | 241 | 243 | 245 | 247 | 249 | 251 | 253 | 254 | 256 | 258 | 0.575 | 296 | 298 | 3090 | 302 | 304 | 305 | 307 | 309 | 311 | 313 | 315 | 317 | 318
0.5 | 244 | 248 | 252 | 254 | 256 | 258 | 260 | 261 | 263 | 265 | 267 | 269 | 0.6 | 306 | 308 | 310 | 312 | 313 | 315 | 317 | 319 | 320 | 322 | 324 | 326 | 328
0.525 | 256 | 260 | 263 | 265 | 267 | 269 | 270 | 272 | 273 | 275 | 277 | 279 | 0.625 | 316 | 318 | 320 | 322 | 323 | 325 | 326 | 328 | 330 | 332 | 333 | 335 | 337
0.55 | 267 | 271 | 274 | 276 | 278 | 279 | 281 | 282 | 284 | 285 | 287 | 289 | 0.65 | 326 | 328 | 330 | 331 | 333 | 334 | 336 | 338 | 339 | 341 | 343 | 344 | 346
0.575 | 279 | 282 | 285 | 287 | 288 | 290 | 291 | 292 | 294 | 296 | 297 | 299 | 0.675 | 336 | 338 | 339 | 341 | 342 | 344 | 345 | 347 | 349 | 350 | 352 | 354 | 355
0.6 | 290 | 293 | 296 | 297 | 299 | 300 | 301 | 303 | 304 | 306 | 308 | 310 | 0.7 | 346 | 348 | 349 | 350 | 352 | 353 | 355 | 356 | 358 | 359 | 361 | 363 | 364
0.625 | 301 | 304 | 306 | 308 | 309 | 310 | 312 | 313 | 315 | 316 | 318 | 320 | 0.725 | 356 | 357 | 358 | 360 | 361 | 362 | 364 | 365 | 367 | 368 | 370 | 372 | 373
0.65 | 312 | 314 | 317 | 318 | 320 | 321 | 322 | 323 | 325 | 326 | 328 | 330 | 0.75 | 365 | 367 | 368 | 369 | 370 | 372 | 373 | 375 | 376 | 378 | 379 | 381 | 382
0.675 | 322 | 325 | 327 | 328 | 330 | 331 | 333 | 333 | 335 | 336 | 338 | 340 | 0.775 | 374 | 376 | 377 | 378 | 379 | 381 | 382 | 383 | 385 | 386 | 388 | 390 | 391
0.7 | 333 | 335 | 337 | 339 | 340 | 341 | 342 | 343 | 345 | 346 | 348 | 350 | 0.8 | 384 | 385 | 386 | 387 | 388 | 390 | 391 | 392 | 394 | 395 | 397 | 398 | 400
0.725 | 342 | 345 | 347 | 348 | 349 | 351 | 352 | 353 | 354 | 356 | 358 | 360 | 0.825 | 393 | 394 | 395 | 396 | 397 | 399 | 400 | 401 | 403 | 404 | 406 | 407 | 408
0.75 | 352 | 355 | 357 | 358 | 359 | 360 | 361 | 363 | 364 | 366 | 368 | 369 | 0.85 | 402 | 403 | 404 | 405 | 406 | 407 | 409 | 410 | 411 | 413 | 414 | 416 | 417
0.775 | 361 | 365 | 367 | 368 | 369 | 370 | 371 | 372 | 374 | 375 | 377 | 379 | 0.875 | 410 | 411 | 412 | 413 | 415 | 416 | 417 | 418 | 420 | 421 | 423 | 424 | 425
0.8 | 369 | 374 | 376 | 377 | 378 | 379 | 380 | 382 | 383 | 385 | 387 | 389 | 0.9 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 427 | 428 | 429 | 431 | 432 | 434
Exposure to Dose Conversion Factors W/Al for 4.2cm 50/50 Breast
kVp
HVL | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49
0.2 | 120 | 125 | 133 | 138 | 143 | 148 | 156 | 160 | 165 | 169 | 177 | 181 | 185 | 188 | 195 | 198 | 200 | 203 | 208 | 216 | 219 | 222 | 225 | 231 | 234
0.25 | 143 | 148 | 155 | 160 | 165 | 169 | 176 | 181 | 185 | 189 | 196 | 200 | 204 | 207 | 213 | 216 | 219 | 221 | 226 | 233 | 236 | 239 | 241 | 247 | 250
0.3 | 166 | 171 | 177 | 182 | 187 | 191 | 197 | 202 | 206 | 210 | 216 | 220 | 223 | 226 | 232 | 235 | 237 | 239 | 244 | 250 | 253 | 256 | 258 | 264 | 266
0.35 | 189 | 195 | 200 | 204 | 209 | 213 | 218 | 222 | 226 | 230 | 236 | 239 | 243 | 246 | 250 | 253 | 255 | 257 | 262 | 267 | 270 | 273 | 275 | 280 | 282
0.4 | 212 | 217 | 222 | 226 | 231 | 235 | 239 | 243 | 247 | 251 | 255 | 258 | 262 | 265 | 269 | 271 | 273 | 275 | 280 | 285 | 287 | 290 | 292 | 296 | 298
0.45 | 234 | 239 | 244 | 248 | 252 | 256 | 260 | 264 | 268 | 271 | 275 | 278 | 281 | 284 | 287 | 290 | 292 | 294 | 298 | 302 | 304 | 307 | 309 | 313 | 315
0.5 | 256 | 261 | 265 | 269 | 273 | 277 | 280 | 284 | 288 | 291 | 294 | 297 | 300 | 303 | 306 | 308 | 310 | 312 | 316 | 319 | 321 | 324 | 326 | 329 | 331
0.55 | 278 | 282 | 286 | 290 | 293 | 297 | 300 | 304 | 307 | 310 | 313 | 316 | 319 | 322 | 324 | 326 | 328 | 330 | 334 | 336 | 339 | 341 | 343 | 345 | 347
0.6 | 300 | 303 | 307 | 310 | 313 | 317 | 320 | 323 | 326 | 329 | 332 | 335 | 337 | 340 | 342 | 344 | 346 | 348 | 351 | 354 | 356 | 358 | 360 | 362 | 363
0.65 | 321 | 324 | 327 | 330 | 333 | 336 | 339 | 342 | 345 | 347 | 350 | 353 | 355 | 358 | 360 | 362 | 363 | 365 | 368 | 370 | 372 | 374 | 376 | 378 | 380
0.7 | 341 | 344 | 347 | 350 | 352 | 355 | 358 | 361 | 363 | 366 | 368 | 370 | 373 | 375 | 377 | 379 | 380 | 382 | 385 | 387 | 389 | 390 | 392 | 394 | 395
0.75 | 360 | 363 | 366 | 369 | 370 | 374 | 376 | 379 | 381 | 383 | 386 | 388 | 390 | 392 | 394 | 395 | 397 | 398 | 401 | 403 | 404 | 406 | 408 | 409 | 411
0.8 | 379 | 382 | 385 | 387 | 389 | 392 | 394 | 396 | 398 | 401 | 403 | 405 | 407 | 408 | 410 | 412 | 413 | 414 | 417 | 418 | 420 | 421 | 423 | 424 | 426
0.85 | 398 | 400 | 403 | 405 | 407 | 409 | 411 | 413 | 415 | 417 | 419 | 421 | 423 | 425 | 426 | 427 | 429 | 430 | 432 | 434 | 435 | 436 | 438 | 439 | 440
0.9 | 415 | 418 | 420 | 422 | 424 | 46 | 428 | 430 | 432 | 434 | 435 | 437 | 439 | 440 | 442 | 443 | 444 | 445 | 447 | 449 | 450 | 451 | 452 | 454 | 455
0.95 | 432 | 435 | 437 | 438 | 440 | 442 | 444 | 446 | 448 | 449 | 451 | 453 | 454 | 455 | 457 | 458 | 459 | 460 | 462 | 463 | 464 | 465 | 467 | 468 | 469
1 | 448 | 451 | 453 | 454 | 456 | 458 | 460 | 461 | 463 | 465 | 466 | 467 | 469 | 470 | 471 | 472 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482
kV Accuracy/Output |  | 170 174: Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV^2 | [8 empty] | Radcal kV correction for Hologic Selenia/Dimension
Target | kV | kV^2 | mGy/mAs | mGy/s |  | Target | kV | kV^2 | mGy/mAs | mGy/s |  | Target | kV | kV^2 | mGy/mAs | mGy/s |  |  |  | Selenia |  | Dimension
| 24 | 170 174: 576; 175 179: 28; 176 180: 784 |  |  | 28 | 784 |  |  |  |  | Measured | W/Rh | W/Ag | W/Rh | W/Ag | W/Al
Filter | 25 | 170 174: 625; 174 178: Filter; 175 179: 30; 176 180: 900 |  | Filter | 30 | 900
| 26 | 170 174: 676; 175 179: 32; 176 180: 1024 |  |  | 32 | 1024
| 28 | 170 174: 784; 175 179: 34; 176 180: 1156 |  |  | 34 | 1156
| 30 | 170 174: 900; 175 179: 36; 176 180: 1296 |  |  | 38 | 1444
| 32 | 170 174: 1024; 175 179: 38; 176 180: 1444 |  |  |  | Slope: | #DIV/0! | #DIV/0!
| 34 | 170 174: 1156; 176 180: Slope:; 177 181: #DIV/0!; 0.34 0.35: #DIV/0! |  |  |  | Intercept: | #DIV/0! | #DIV/0!
|  | Slope: | #DIV/0! | #DIV/0! |  |  |  | Intercept: | #DIV/0! | #DIV/0! |  |  |  | R^2: | #DIV/0! | #DIV/0!
|  | Intercept: | #DIV/0! | #DIV/0! |  |  |  | R^2: | #DIV/0! | #DIV/0!
|  | R^2: | #DIV/0! | #DIV/0!
Half Value Layer
Target | kV | HVL | Target | kV | HVL | Target | kV | HVL
| 24 | 170 174: #DIV/0!; 172 176: 28; 173 177: #DIV/0!; 175 179: 28
Filter | 25 | #DIV/0! | Filter | 30 | #DIV/0! | Filter | 30
| 26 | 170 174: 0; 172 176: 32; 173 177: #DIV/0!; 175 179: 32
| 28 | 170 174: #DIV/0!; 172 176: 34; 173 177: #DIV/0!; 175 179: 34
| 30 | 170 174: 0; 172 176: 36; 173 177: 0; 175 179: 38
| 32 | 170 174: #DIV/0!; 172 176: 38; 173 177: 0
| 34 | 170 174: #DIV/0!
| Slope: | 170 174: #DIV/0!; 172 176: Slope:; 173 177: #DIV/0!; 175 179: Slope:; 176 180: #DIV/0!
| Intercept: | 170 174: #DIV/0!; 172 176: Intercept:; 173 177: #DIV/0!; 175 179: Intercept:; 176 180: #DIV/0!
| R^2: | 170 174: #DIV/0!; 172 176: R^2:; 173 177: #DIV/0!; 175 179: R^2:; 176 180: #DIV/0!
Target | Filter | 170 174: C; 172 176: Critiera:; 173 177: Min HVL > kV/100 (no paddle)
Mo | Mo | 170 174: 0.12; 173 177: Min HVL > kV/100 + 0.03 (with paddle)
Mo | Rh | 170 174: 0.19; 173 177: Max HVL < kV/100 + C
Rh | Rh | 170 174: 0.22
DGN | #VALUE! | 174 178: Combo
Mo/Mo | Mo/Rh | Rh/Rh | W/Rh | W/Ag | W/Al | W/Rh | W/Ag | W/Al
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Radcal kV Correction for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Rh | < | 27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| >= | 27.586 | -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Ag | < | 30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| >= | 30.1 | -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
DGN Mo/Mo: Radcal kV Correction for Hologic Dimension |  | Barracuda/MPD kVp correction table for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | W/Rh (RQ M6) |  |  | W/Ag (RQ M10)
W/Rh | < | 26.9 | 138.8867 | -10.7264 | 0.26216 | -0.00082 |  |  |  | kV | LFS | SFS | kV | LFS | SFS
| >= | 26.9 | -5009.7752 | 605.732 | -27.3019 | 0.54671 | -0.0040986 |  |  | 22 | 0.2 | 0.1 | 22 | 0.2 | -0.2
[10 empty] | 23 | 0.2 | -0.1 | 23 | 0.4 | -0.1
W/Ag | kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | 24 | 0.1 | -0.4 | 24 | 0.4 | 0
| < | 28.7 | 296.3419 | -31.6292 | 1.1803 | -0.01417 |  |  |  | 25 | 0.1 | -0.3 | 25 | 0.5 | -0.1
#VALUE! Mo/Rh: >=; #VALUE! Rh/Rh: 28.7; #VALUE! W/Rh: 4.8345; #VALUE! W/Ag: 0.91924 |  | 26 | 0 | -0.2 | 26 | 0.5 | -0.2
[10 empty] | 27 | 0.1 | -0.3 | 27 | 0.7 | -0.2
W/Al |  |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | 28 | 0.2 | -0.5 | 28 | 0.9 | -0.1
#VALUE! Mo/Rh: <; #VALUE! Rh/Rh: 28.7; #VALUE! W/Rh: 49.3111; #VALUE! W/Ag: -2.9302; #VALUE! W/Al: 0.07379 |  | 29 | 0.4 | -0.2 | 29 | 0.8 | -0.3
#VALUE! Rh/Rh: 28.7 < kV < 30.1; #VALUE! W/Rh: -24.875; #VALUE! W/Ag: 1.8032 |  | 30 | 0.6 | 0 | 30 | 0.8 | -0.4
#VALUE! Mo/Rh: >=; #VALUE! Rh/Rh: 30.1; #VALUE! W/Rh: -4.8346; #VALUE! W/Ag: 1.1571 |  | 31 | 0.8 | 0.1 | 31 | 0.8 | -0.3
[10 empty] | 32 | 1 | 0.2 | 32 | 0.8 | -0.3
DGN Mo/Mo: CNR Correction (XX8xxxxx series detectors) |  | 33 | 1 | 0.4 | 33 | 0.9 | -0.3
DGN Mo/Mo: Conventional |  | 34 | 1.1 | 0.6 | 34 | 1.1 | -0.3
Thickness (cm) | 0 | 1 | 2 | 3 | 4 | 5 | 6 |  |  | 35 | 1.2 | 0.6 | 35 | 1.1 | -0.2
2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 36 | 1.4 | 0.6 | 36 | 1.2 | -0.2
4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 37 | 1.5 | 0.7 | 37 | 1.1 | -0.1
6 | 1.04 | 1.02 | 1.01 | 1.08 | 1.35 | 1.35 | 1.31 |  |  | 38 | 1.6 | 0.9 | 38 | 1.1 | 0
8 | 1.23 | 1.2 | 1.18 | 1.1 | 1.82 | 1.85 | 1.82 |  |  | 39 | 1.7 | 0.8 | 39 | 1.3 | 0.3
Magnification
Thickness (cm) | 0 | 1 | 2 | 3 | 4
2 | 1.15 | 1.15 | 1.15 | 1.15 | 1.5
4 | 1.15 | 1.15 | 1.15 | 1.15 | 1.5
6 | 1.19 | 1.18 | 1.18 | 1.18 | 1.55
8 | 1.28 | 1.24 | 1.22 | 1.29 | 1.67
Stereotactic biopsy |  |  | CE2D
Thickness (cm) | 0 |  | Thickness | 0 | 1
2 | 0.91 |  | 2 | 1 | 0.72
4 | 1 |  | 4 | 1 | 0.9
6 | 1.32 |  | 6 | 1.91 | 1.71
8 | 1.88 |  | 8 | 1.81 | 2.22
Tomo 15 Proj STD |  |  |  | Tomo 15 Proj ENH
Thickness (cm) | 0 | 1 |  | Thickness | 0
2 | 0.7 | 0.7 |  | 2 | 0.57
4 | 0.91 | 0.91 |  | 4 | 0.91
6 | 1.55 | 1.55 |  | 6 | 1.68
8 | 2.78 | 2.78 |  | 8 | 1.93
CNR Correction (Selenia Dimension W/Rh)
Thickness (cm) | 0 |  | 1 |  | 2 |  | 2
| LFS | SFS | LFS | SFS | LFS | SFS | LFS | SFS
0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
3 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
5 | 1 | 1 | 1 | 1 | 1.15 | 1 | 1 | 1
6 | 1 | 1 | 1 | 1.15 | 1.15 | 1 | 1 | 1
7 | 1.1 | 1 | 1.1 | 1.15 | 1.15 | 1 | 1 | 1
8 | 1.15 | 1 | 1.15 | 1.15 | 1.15 | 1 | 1 | 1
9 | 1.1 | 1 | 1.1 | 1.15 | 1.15 | 1 | 1 | 1
10 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
11 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
12 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
13 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
14 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## DataPage
Sheet | MAMMO_HOL
SurveyID
Light field
Resolution | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0
2D MGD
3D MGD | NA
Combo MGD | NA
kV Accuracy | / | kV | Output | / | kV | Output | / | kV | Output
Output | 24 |  |  | 28 |  |  | 28
| 25 |  |  | 30 |  |  | 30
| 26 |  |  | 32 |  |  | 32
| 28 |  |  | 34 |  |  | 34
| 30 |  |  | 36 |  |  | 38
| 32 |  |  | 38
34
Linearity | /
20
50
100
300
HVL | / | 24
| / | 25
| / | 28
| / | 32
| / | 28
| / | 30
| / | 32
| / | 34
| / | 28
| / | 30
| / | 32
| / | 34
| / | 38
Acq | 2D | Mag | 3D | Film
Fibers | 0 | 0 | 0 | 0
Specks | 0 | 0 | 0 | 0
Masses | 0 | 0 | 0 | 0
SNR target/filter | /
SNR
CNR
Note on B4: Eugene Mah: Target/Filter
Note on B5: Eugene Mah: kV
Note on B6: Eugene Mah: Line pairs/mm

## HVLProcessing
24 | kVp |  |  |  |  | 28 | kVp |  |  |  |  | 28 | kVp
mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE!
0 | [5 empty] | 0 | [5 empty] | 0
0.4 | [5 empty] | 0.5 | [5 empty] | 0.4
|  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP)
25 | kVp |  |  |  |  | 30 | kVp |  |  |  |  | 30 | kVp
mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE!
0 | [5 empty] | 0 | [5 empty] | 0
0.4 | [5 empty] | 0.5 | [5 empty] | 0.4
|  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP)
28 | kVp |  |  |  |  | 32 | kVp |  |  |  |  | 32 | kVp
mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE!
0 | [5 empty] | 0 | [5 empty] | 0
0.5 | [5 empty] | 0.6 | [5 empty] | 0.5
|  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP)
32 | kVp |  |  |  |  | 34 | kVp |  |  |  |  | 34 | kVp
mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE!
0 | [5 empty] | 0 | [5 empty] | 0
0.5 | [5 empty] | 0.6 | [5 empty] | 0.5
|  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP) |  |  |  |  | Data Exp | LN(EXP)
34 | kVp | [10 empty] | 38 | kVp
mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE! | [6 empty] | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp | #VALUE!
0 | [11 empty] | 0
0.4 | [11 empty] | 0.6
|  |  | Data Exp | LN(EXP) | [10 empty] | Data Exp | LN(EXP)

## Corrected kV
GENERAL
Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
Read the true corrected kV in the adjacent cell to the right.
If that cell is light green, the corrected kV is accurate.
If the cell is red, the entered kV is out of range and the correction should not be used.
If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.
Tungsten-Rhodium (W / Rh)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| Radcal kV | True kV
| 29 | 30.2821
Tungsten-Silver (W / Ag)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
| Radcal kV | True kV
| 37.1 | 38.9675
27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
0