## Gen_form
1 | [10 empty] | Revision 0.2-20260319 (Has 39 data records) | [10 empty] | PRINT AREAS
2 |  |  |  |  | Medical University of South Carolina | [9 empty] | Medical University of South Carolina | [6 empty] | FIRST,EXP_BLD,FS_KV,REPRO,LIN,HVL_IMGQUAL,OUTPUT,CTPG1
3 |  |  |  |  | Charleston, South Carolina | [9 empty] | Charleston, South Carolina | [6 empty] | ALL: | FIRST,COMPG1,COMPG2,EXP_BLD,FS_KV,TIM,REPRO,LIN,HVL_IMGQUAL,OUTPUT,CTPG1
4 | [23 empty] | REM: | (1) Go to last year's spreadsheet (see "**Note", below) and copy from ITS "THIS YEAR" range
5 |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | [8 empty] | Radiation Oncology IGRT Compliance Inspection | [10 empty] | (2) Come back to this sheet and paste (use paste special, "values") into "LAST YEAR" range
6 | [24 empty] | LAST | Cell AD6 | THIS
7 | [23 empty] | Measurement Parameter | YEAR |  | YEAR | (Locked) | RANGES: | LASTYEAR (Data, starting with cell AB8)
8 |  |  |  |  | System Information | [6 empty] | Date: |  |  |  | System Information |  |  | Inspector: | Eugene Mah |  |  |  | Print Area: |  | Change | FIRST,EXP_BLD,FS_KV,REPRO,LIN,HVL_IMGQUAL,OUTPUT,CTPG1 | Y2 |  | THISYEAR (Data, starting with cell AD8)
9 | [11 empty] | Last Year: | [6 empty] | Last Year: |  |  |  |  | Date: |  |  |  | O8
10 |  | Location | [9 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  |  | Inspector: |  | Change | Eugene Mah | V8
11 |  |  | Facility: |  |  |  | Site: | [5 empty] | Facility: |  |  |  | Site Number: | [6 empty] | Facility: |  |  |  | P11
12 |  |  | Department: |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [6 empty] | Department: |  |  |  | P12
13 |  |  | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [6 empty] | Area/Division: |  |  |  | P13
14 |  |  | Survey ID: |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: | [6 empty] | Survey ID: |  |  |  | P14
15 | [12 empty] | Room Number: |  |  |  | Accession Number: | [6 empty] | Room Number: |  |  |  | P15
16
17 |  | X-Ray Generator | [9 empty] | X-Ray Generator | [11 empty] | X-Ray Generator
18 |  |  | Manufacturer: |  |  |  | Date of Manufacture: | [5 empty] | Manufacturer: |  |  |  | Date of Manufacture: | [6 empty] | Manufacturer: |  |  |  | P18
19 |  |  | Model: |  |  |  | Serial Number: | [5 empty] | Model: |  |  |  | Serial Number: | [6 empty] | Model: |  |  |  | P19
20 |  |  | Max. kVp: |  |  |  | Max. mA: | [5 empty] | Max. kVp: |  |  |  | Max. mA: | [6 empty] | Max. kVp: |  |  |  | P20
21 |  |  | Software Version: | [9 empty] | Software Version: | [10 empty] | Software Version: |  |  |  | P21
22 | [23 empty] | Date of Manufacture: |  |  |  | T18
23 |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes |  |  |  | X-Ray Tube 1 | [5 empty] | Focal Spot Sizes (mm) | [5 empty] | Serial Number: |  |  |  | T19
24 |  |  | Tube Designation/Use: |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Large: | [6 empty] | Max. mA: |  |  |  | T20
25 |  |  | Date of Manufacture: |  |  |  | Small: | [5 empty] | Date of Manufacture: |  |  |  | Small:
26 | [16 empty] | Micro: | [6 empty] | X-Ray Tube 1
27 |  |  | Insert Manufacturer: |  |  |  |  | Filtration |  |  |  |  | Insert Manufacturer: |  |  |  |  | Filtration | [5 empty] | Tube Designation/Use: |  |  |  | P24
28 |  |  | Insert Model: |  |  |  | Inherent: | [5 empty] | Insert Model: |  |  |  | Inherent: | [6 empty] | Date of Manufacture: |  |  |  | P25
29 |  |  | Insert Serial Number: |  |  |  | Added: | [5 empty] | Insert Serial Number: |  |  |  | Added:
30 | [23 empty] | Insert Manufacturer: |  |  |  | P27
31 |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor |  |  |  |  | Housing Manufacturer: |  |  |  |  | Image Receptor | [5 empty] | Insert Model: |  |  |  | P28
32 |  |  | Housing Model: |  |  |  | Film: | [5 empty] | Housing Model: |  |  |  | Film: | [6 empty] | Insert Serial Number: |  |  |  | P29
33 |  |  | Housing Serial Number: |  |  |  | Screen: | [5 empty] | Housing Serial Number: |  |  |  | Screen:
34 | [23 empty] | Housing Manufacturer: |  |  |  | P31
35 | [23 empty] | Housing Model: |  |  |  | P32
36 | [23 empty] | Housing Serial Number: |  |  |  | P33
37
38 | [23 empty] | Site Number: |  |  |  | T11
39 | [23 empty] | Authorized Use: |  |  |  | T12
40 | [23 empty] | Date of Installation: |  |  |  | T13
41 | [23 empty] | Number of X-ray Tubes: |  |  |  | T14
42 | [23 empty] | Accession Number: |  |  |  | T15
43
44 | [23 empty] | Focal Spot Sizes
45 | [23 empty] | Large: |  |  |  | T24
46 | [23 empty] | Small: |  |  |  | T25
47 | [23 empty] | Micro: |  |  |  | T26
48 | [23 empty] | Filtration
49 |  |  |  |  | Inspection Results | [10 empty] | Inspection Results | [7 empty] | Inherent: |  |  |  | T28
50 | [23 empty] | Added: |  |  |  | T29
51 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA
52 | Rule Number |  |  |  | Labels, Notices, Postings |  |  |  | Compliance |  | 2 | 2 | Is this a new installation? | [10 empty] | Image Receptor
53 | [23 empty] | Film: |  |  |  | T32
54 | DES RHB 2.5.1.1 |  | DES registration sticker is present, clearly visible and legible. | [5 empty] | TBD | [5 empty] | Labels, Notices, Postings |  |  | 1 | 1 | Tube Number** |  |  |  | Screen: |  |  |  | T33
55 | DES RHB 10.2.1 |  | DES form SC-RHA-20 "Notice to Employees" posted or referenced. | [5 empty] | TBD | [10 empty] | for this room or
56 | DES RHB 4.2.7 |  | Pregnancy Warning sign is posted | [5 empty] | NA |  |  |  | DES registration sticker is present, clearly visible and legible. | [6 empty] | generator |  |  |  | New install? |  | Change | 2 | M52
57 | DES RHB 4.3.1 |  | Radiation warning label present on the generator control panel. | [5 empty] | TBD |  |  |  | DES form SC-RHA-20 "Notice to Employees" posted or referenced. | [5 empty] | **Note: Input only one of the |  |  |  |  | Tube Designation |  | Change | 1 | U54
58 | DES RHB 4.2.6 |  | Technique chart is present and accurate. |  |  |  | DATE: |  | NA |  | 3 |  | Pregnancy Warning sign is posted | [5 empty] | following numerals or names:
59 | JCAHO |  | Patient exposure/dose measurements are posted. |  |  |  | DATE: |  | NA |  |  |  | Radiation warning label present on the generator control panel. | [5 empty] | 1 | 2 |  |  |  | Table Bucky
60 |  |  | Tube loading charts for each focal spot are posted. | [5 empty] | NA |  | 3 | DATE: |  | Technique chart present. |  |  |  |  | AP | 3 |  |  |  | TB used? |  | Change | 1 | M409
61 |  |  | Tube loading charts are correct for installed tube(s). | [5 empty] | NA |  | 3 | DATE: |  | Patient exposure/dose measurements posted. |  |  |  |  | Front | Lat |  |  |  | Beam alignment on table top? |  | Change | 2 | M410
62 |  |  | Operators manuals are available. | [5 empty] | TBD |  | 3 |  | Tube loading charts for each focal spot are posted. | [5 empty] | Frontal | Lateral
63 |  |  | Monthly radiation monitoring reports are posted. | [5 empty] | TBD |  | 3 |  | Tube loading charts are correct for installed tube(s). | [5 empty] | Ceph | Pan |  |  |  | Source to Table Top (cm) |  | Change | 100 | M420
64 | [12 empty] | Operators manuals are available. | [5 empty] | Mobile | BCM |  |  |  | Table top to cassette (cm) |  | Change | 50 | M421
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  |  | Monthly radiation monitoring reports are posted.
66 | Room Number: | [6 empty] | Survey ID: | [15 empty] | Table bucky cm/in |  | Change | cm | V423
1 | [9 empty] | Medical University of South Carolina | [8 empty] | NEW VALUES |  |  |  |  | Bucky SID |  | Change | 150 | U423
2 |  |  |  |  | Inspection Results |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | [23 empty] | Indicated field size
4 | [23 empty] | Transverse 1 |  | Change | 5 | M192
5 |  |  |  |  | Generator |  |  |  | Compliance | [14 empty] | Longitudinal 1 |  | Change | 5 | N192
6 | DES RHB 4.3.9.1 |  | Means are provided to indicate technique factors prior to exposure. | [5 empty] | TBD | [14 empty] | Transverse 2 |  | Change | 10 | M193
7 |  |  | All technique indicators functional and are free of excessive wear. | [5 empty] | TBD | [14 empty] | Longitudinal 2 |  | Change | 10 | N193
8 |  |  | Technique overload circuits function to block excessive technique. | [5 empty] | TBD
9 |  |  | Anode heat calculator/indicator is functioning properly. | [5 empty] | NA | [14 empty] | X-RAY TUBE
10 | DES RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational. | [5 empty] | TBD | [6 empty] | COMMENTS | [6 empty] | Large Focus: | mA/mAs |  | Change | mAs | P231
11 | DES RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational. | [5 empty] | TBD |  |  | Note: Type all comments in the shaded cells. | [11 empty] | kVp |  |  |  | O233
12 | DES RHB 4.7.4.2.4 |  | Audible exposure end indicator present and operational. | [5 empty] | TBD | [5 empty] | Previous Year Comments: | [8 empty] | mA |  |  |  | P233
13 |  |  | Exposure delay time is not excessive. | [5 empty] | TBD |  |  | Comments: | [11 empty] | Time |  |  |  | Q233
14 | DES RHB 4.7.7 |  | kVp accurate to within 5% of indicated. | [5 empty] | TBD |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **The DES registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag. | [6 empty] | SID |  | Change | 150 | R233
15 | DES RHB 4.7.5 |  | kVp and exposure reproducibility coeff. of var. < 5%. | [5 empty] | NA |  |  | Additional comments: | [11 empty] | OFD |  | Change | 100 | S233
16 | DES RHB 4.7.7 |  | Timer accuracy within 10%. | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **A technique chart was not posted in the room. A technique chart reflecting techniques currently used with the CR system must be provided (DES RHB 4.2.8).
17 | DES 4.7.4.2.7 |  | Timer reproducibility coeff. of var. < 5%. | [5 empty] | NA |  |  | Additional comments: | [11 empty] | Last year Lrg FS resolution |  |  |  | U233
18 | DES RHB 4.7.8 |  | Exposure coefficient of linearity < 10%. | [5 empty] | TBD |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console. | [6 empty] | Previous 1 yr Lrg FS resolution |  |  |  | V233
19 | [11 empty] | Additional comments:
20 |  |  |  |  | X-ray Tube | [6 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Tube loading charts for each focal spot were not available at the time of inspection. These charts should be obtained and posted near the operator's console. | [5 empty] | Small Focus: | mA/mAs |  | Change | mAs | P236
21 |  |  | Focal spot resolution is within acceptable parameters | [5 empty] | TBD |  |  | Additional comments: | [11 empty] | kVp |  |  |  | O238
22 |  |  | The x-ray output @40" within +/-25% from previous measurements. | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **Operator's manuals for the imaging equipment were not available at the time of inspection. These should be obtained and made available to the operator of the equipment. | [6 empty] | mA |  |  |  | P238
23 | [11 empty] | Additional comments: | [11 empty] | Time |  |  |  | Q238
24 |  |  |  |  | Radiation Safety | [6 empty] | Do NOT exceed 390 |  | 0 |  | 0 | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office. | [6 empty] | SID |  | Change | 150 | R238
25 | DES RHB 4.7.4.2.1 |  | Continuous pressure on exposure switch required to make an exposure. | [5 empty] | TBD |  |  | Additional comments: | [11 empty] | OFD |  | Change | 100 | S238
26 |  |  | Spot film exposure blocked when there is no cassette | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | **The light field illumination failed to meet the minimum specification of 160 lux. The light source should be adjusted or replaced so as to meet this requirement.
27 |  |  | Spot film exposure blocked when serial exposures are completed | [5 empty] | NA |  |  | Additional comments: | [11 empty] | Last year Sm FS resolution |  |  | 0 | U238
28 | 21CFR 1020.31(g)(4) |  | Spot film device provides acceptable centering to the image receptor | [5 empty] | NA |  |  | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Previous 1 yr Sm FS resolution |  |  |  | V238
29 | DES RHB 4.7.4.2.2 |  | Exposure switch for selected tube is secured at a safe location | [5 empty] | TBD |  |  | Additional comments:
30 | DES RHB 4.7.4.2.5 |  | Operator has full visibility of the exposure factors and patient. | [5 empty] | TBD |  |  | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Test Equipment Utilized: |  | Change | Piranha CB2-19020491 | O251
31 | DES RHB 4.2.14.7 |  | Lead aprons available. |  | Half: 0 |  | Full: 0 |  | NA | [14 empty] | Piranha CB2-17090320
32 | DES RHB 4.2.14.7 |  | Lead gloves available. |  |  | Number (Left + Right): | 0 |  | NA
33 |  |  | A properly designed and installed apron rack is present. | [5 empty] | NA | [14 empty] | LFSDD |  | Change | cm | N257
34 | DES RHB 4.2.14.7 |  | Documentation of annual protective apparel integrity inspection is available. | [5 empty] | NA | [14 empty] | SFSDD |  | Change | cm | N275
35 |  |  | Patient restraint devices available. | [5 empty] | TBD | [5 empty] | DO NOT USE THIS SPACE
36 | DES RHB 4.2.12 |  | Gonadal shielding is present and functional. | [5 empty] | NA | [14 empty] | Timer
37 |  |  | Radiation warning light at room entrance is functional with selected tube | [5 empty] | NA | [14 empty] | Generator Type |  | Change | HF | O369
38 | [23 empty] | Timer Test Valid? |  | Change | 2 | M370
39 |  |  |  |  | Physical Inspection
40 |  |  | System clean and free of debris. | [5 empty] | TBD
41 |  |  | All cables free from excessive wear or restraints. | [5 empty] | TBD
42 |  |  | System free of missing or damaged components. | [5 empty] | TBD | [14 empty] | Large Focus Output
43 |  |  | Table movements and locks acceptable | [5 empty] | TBD | [14 empty] | SDD(in.) |  |  | 0 | AB188
44 |  |  | Tube crane movements and locks acceptable. | [5 empty] | TBD
45 |  |  | All other movements and locks functional. | [5 empty] | TBD | [6 empty] | COMMENTS PAGE 2 | [7 empty] | kVp |  |  |  | P1397
46 | [27 empty] | P1398
47 | [14 empty] | Previous Year Comments: | [12 empty] | P1399
48 | [11 empty] | Comments: | OPERATIONAL NOTES | [14 empty] | P1400
49 | [11 empty] | Do NOT exceed 390 |  | 17 |  | 0 | [11 empty] | P1401
50 | [11 empty] | Additional comments: | Isocenter is 100 cm from focal spot | [14 empty] | P1402
51 | [11 empty] | Do NOT exceed 390 |  | 35 |  | 0 | [11 empty] | P1403
52 | [11 empty] | Additional comments: | Collimator blade locations and field sizes are referenced to isocenter. | [14 empty] | P1404
53 | [11 empty] | Do NOT exceed 390 |  | 71 |  | 0
54 | [11 empty] | Additional comments: | Use 5x5cm field size for exposure/kV measurements. | [10 empty] | mGy/mAs |  |  |  | R1397
55 | [11 empty] | Do NOT exceed 390 |  | 50 |  | 0 | [7 empty] | @ 0 in. |  |  |  | R1398
56 | [11 empty] | Additional comments: | For generator check, align kV/exposure detectors at isocenter. | [14 empty] | R1399
57 | [11 empty] | Do NOT exceed 390 |  | 62 |  | 0 | [11 empty] | R1400
58 | [11 empty] | Additional comments: | [15 empty] | R1401
59 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1402
60 | [11 empty] | Additional comments: | [15 empty] | R1403 |  |  | 80 kV
61 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1404 |  |  | mm Al. | Avg. Exp.
62 | [11 empty] | Additional comments: | [18 empty] | #DIV/0! | #DIV/0!
63 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Small Focus Output | [6 empty] | #DIV/0! | #DIV/0!
64 | [11 empty] | Additional comments: | [11 empty] | SDD(in.) |  |  | 0 | N1411 |  |  | #DIV/0! | #DIV/0!
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Do NOT exceed 390 |  | 0 |  | 0 | [14 empty] | #DIV/0! | #DIV/0!
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | Additional comments: | [11 empty] | kVp |  |  |  | P1414 |  |  | #DIV/0! | #DIV/0!
1 | [9 empty] | Medical University of South Carolina |  | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1415 |  |  | #DIV/0! | #DIV/0!
2 |  |  |  |  | COMMENTS PAGE 1 |  |  |  |  | Radiation Oncology IGRT Compliance Inspection |  | Additional comments: | [15 empty] | P1416
3 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1417
4 | [11 empty] | Additional comments: | [15 empty] | P1418
5 | Comments: | [10 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | P1419
6 | [11 empty] | Additional comments:
7 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | mGy/mAs |  |  |  | R1414
8 | [11 empty] | Additional comments: | [11 empty] | @ 0 in. |  |  |  | R1415
9 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1416
10 | [11 empty] | Additional comments: | [15 empty] | R1417
11 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [11 empty] | R1418
12 | [11 empty] | Additional comments: | [15 empty] | R1419
13 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | mGy/mAs Output
14 | [11 empty] | Additional comments: | [11 empty] | Lrg Focus Max. |  | Change | NA | Q1406
15 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Lrg Focus Median |  | Change | NA | R1406
16 | [11 empty] | Additional comments: | [11 empty] | Lrg Focus Min |  | Change | NA | S1406
17 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Sml Focus Max. |  | Change | NA | Q1421
18 | [11 empty] | Additional comments: | [11 empty] | Sml Focus Median |  | Change | NA | R1421
19 | [11 empty] | Do NOT exceed 390 |  | 0 |  | 0 | [7 empty] | Sml Focus Min |  | Change | NA | S1421
20
21 | [23 empty] | COMMENT PAGE 1 |  |  | 0 | O79
22
23 | [26 empty] | 0 | O81
24 | [14 empty] | DO NOT USE THIS SPACE
25 | [26 empty] | 0 | O83
26
27 | [26 empty] | 0 | O85
28
29 | [26 empty] | 0 | O87
30
31 | [26 empty] | 0 | O89
32
33 | [26 empty] | 0 | O91
34
35 | [26 empty] | 0 | O93
36
37 | [26 empty] | 0 | O95
38
39 | [23 empty] | COMMENT PAGE 2 |  | Change | OPERATIONAL NOTES | O114
40
41 | [15 empty] | IMAGE RECEPTOR | [9 empty] | Change | Isocenter is 100 cm from focal spot | O116
42 | [10 empty] | Enter 1 for YES, 2 for NO
43 | [10 empty] | 1 | 1 | Is this tube used with a table bucky or other horizontal type film image receptor? | [12 empty] | Change | Collimator blade locations and field sizes are referenced to isocenter. | O118
44 | [10 empty] | 2 | 2 | Mag. Correction - Is beam alignment test being done with the film on the table top?
45 | [25 empty] | Change | Use 5x5cm field size for exposure/kV measurements. | O120
46 | [10 empty] | Beam Alignment |  |  |  |  | NOTE: | Use consistent units (cm/in.)
47 | [10 empty] | 100 | 100 | (cm) Source to Table |  |  |  | throughout |  | Last Year: | ?? | [5 empty] | Change | For generator check, align kV/exposure detectors at isocenter. | O122
48 | [10 empty] | 50 | 50 | (cm) Table top to receptor
49 | [18 empty] | Push button to select cm/in. | [7 empty] | 0 | O124
50 | [17 empty] | Bucky SID: | 150 | cm
51 | [17 empty] | Last Year | [8 empty] | 0 | O126
52 | [17 empty] | New
53 | [26 empty] | 0 | O128
54 | [12 empty] | Indicated Field Size @SID |  | Measured* Radiation |  | Field Size Indicators |  | Acceptable
55 | [10 empty] | New Values |  | (cm) |  | (cm) |  | Error (% SID) |  | Field Size Indicators | [7 empty] | 0 | O130
56 | [10 empty] | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal
57 | [10 empty] | 5 | 5 | 5 | 5 |  |  | NA | NA | NA | NA | [6 empty] | 0 | O132
58 | [10 empty] | 10 | 10 | 10 | 10 |  |  | NA | NA | NA | NA
59 | [10 empty] | *Dimension of highest optical density region | [15 empty] | 0 | O134
60 | [10 empty] | **Dimension between the image of the coins marking the edge of the light field
61 | [26 empty] | 0 | O136
62 | [11 empty] | Criteria: | The indicated field length/width are accurate to within 2% of the SID.
63 | [12 empty] | X-ray field and light field length/width are congruent within 2% SID. | [13 empty] | 0 | O138
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah | [17 empty] | 0 | O140
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina | [16 empty] | 0 | O142
2 |  |  |  |  | Inspection Results |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | [26 empty] | 0 | O144
4 |  |  |  |  | COMMENTS PAGE 2
5 | [26 empty] | 0 | O146
6
7 | Comments: | OPERATIONAL NOTES | [12 empty] | X-ray Tube |  |  |  |  | 1 Year | 2 Year | [5 empty] | 0 | O148
8 | [10 empty] | Focal Spot Resolution | [7 empty] | Measured | Previous | Previous
9 | [10 empty] | Large: |  | mm | mAs |  |  |  |  | Resolution | Resolution | Resolution | [5 empty] | 0 | O150
10 |  | Isocenter is 100 cm from focal spot | [10 empty] | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm)
11 | [15 empty] | 150 | 100 | 3 | [5 empty] | COMMENTS
12 | [11 empty] | New Values: |  |  |  | 150 | 100 | [6 empty] | Focal Spot Resolution |  |  |  | O223
13 |  | Collimator blade locations and field sizes are referenced to isocenter. | [16 empty] | Measured | Previous | Previous |  |  | Large Focus kV accuracy |  |  |  | O243
14 | [10 empty] | Small: |  | mm | mAs |  |  |  |  | Resolution | Resolution | Resolution |  |  | Small Focus kV accuracy |  |  |  | O259
15 | [12 empty] | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  |  | Large Focus Timer Reprod-1 |  |  |  | O291
16 |  | Use 5x5cm field size for exposure/kV measurements. | [13 empty] | 150 | 100 | 3 | [5 empty] | Large Focus Timer Reprod-2 |  |  |  | O293
17 | [11 empty] | New Values: |  |  |  | 150 | 100 | [6 empty] | Small Focus Timer Reprod-1 |  |  |  | O310
18 | [12 empty] | * SID = Source to Image (film) Distance. | [10 empty] | Small Focus Timer Reprod-2 |  |  |  | O312
19 |  | For generator check, align kV/exposure detectors at isocenter. | [10 empty] | ** OFD = Object (resolution pattern) to Film Distance. | [10 empty] | Exposure Timer |  |  |  | O350
20 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [12 empty] | Large Focus Linearity |  |  |  | O395
21 | [11 empty] | Focal spot resolution is within acceptable parameters | [11 empty] | Small Focus Linearity |  |  |  | O410
22 | [12 empty] | Criteria: | No significant change in resolution since previous measurement. | [9 empty] | Half Value Layer |  |  |  | O460
23
24 | [23 empty] | Tube Designation(column 1): |  | Change | 1 | U59
25 | [11 empty] | Comments: | [13 empty] | Change | AP | U60
26 | [11 empty] | Do NOT exceed 390 |  | 0 | [11 empty] | Change | Front | U61
27 | [11 empty] | New | [13 empty] | Change | Frontal | U62
28 | [15 empty] | Generator Performance | [9 empty] | Change | Ceph | U63
29 | [11 empty] | Test Equipment Used: | Piranha CB2-19020491 |  |  | Calibration Date: | [9 empty] | Change | Mobile | U64
30 | [11 empty] | New: | Piranha CB2-19020491 |  |  | Calibration Due: | [11 empty] | U65
31 | [10 empty] | Large Focus kV Accuracy and Output
32 | [23 empty] | Tube Designation(column 2): |  | Change | 2 | V59
33 | [14 empty] | Note: See focal spot data above for mA/mAs selector button | [10 empty] | Change | 3 | V60
34 | [10 empty] | Test | SDD* | Focal | Filter Pack/ |  |  | Set |  | Measured values | [6 empty] | Change | Lat | V61
35 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  |  |  |  | Change | Lateral | V62
36 | [11 empty] | 100 | Large |  | 60 | 20 | [9 empty] | Change | Pan | V63
37 | [11 empty] | 100 | Large |  | 80 | 20 | [9 empty] | Change | BCM | V64
38 | [11 empty] | 100 | Large |  | 100 | 20 | [11 empty] | V65
39 | [11 empty] | 100 | Large |  | 120 | 20
40 | [11 empty] | 100 | Large |  | 140 | 20 | [7 empty] | HVL Measured kVp: |  | Change | HVL @80 kVp (Avg kV) | S464
41 | [23 empty] | HVL Measured: |  | #DIV/0! | #DIV/0! | S465
42
43 | [23 empty] | No. Half Aprons: |  |  | 0 | R520
44 | [10 empty] | *SDD=Source to detector distance | [12 empty] | No. Full Aprons: |  |  | 0 | T520
45 | [11 empty] | Comments: | [11 empty] | No. Lead Gloves: |  |  | 0 | R521
46 | [11 empty] | Do NOT exceed 190 |  | 0 |  |  |  |  | Acceptable: | TBD
47 | [11 empty] | New | [11 empty] | Contact/access Info.:
48 | [10 empty] | Small Focus kV Accuracy and Output | [16 empty] | T1
49 | [27 empty] | T2
50 | [27 empty] | T3
51 | [14 empty] | Note: See focal spot data above for mA/mAs selector button | [12 empty] | T4
52 | [10 empty] | Test | SDD* | Focal | Filter Pack/ |  |  | Set |  | Measured values | [8 empty] | T5
53 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | [6 empty] | T6
54 | [11 empty] | 100 | Small |  | 50 | 10
55 | [11 empty] | 100 | Small |  | 70 | 10 | [7 empty] | CT Protocols
56 | [11 empty] | 100 | Small |  | 90 | 10 | [7 empty] | Adult Abdomen
57 | [11 empty] | 100 | Small |  | 110 | 10 | [7 empty] | Protocol name |  | Change | 125 kV Half Fan | P542
58 | [11 empty] | 100 | Small |  | 130 | 10 | [7 empty] | kVp |  | Change | 125 | P543
59 | [23 empty] | mA |  | Change | 60 | P544
60 | [10 empty] | *SDD=Source to detector distance | [12 empty] | mAs |  | Change | 1.2 | P545
61 | [11 empty] | Comments: | [11 empty] | Eff mAs |  |  |  | P546
62 | [11 empty] | New |  | 0 | [9 empty] | Time/rotation (s) |  |  |  | P547
63 | [13 empty] | (Combined comments) |  |  |  |  | Acceptable: | TBD |  |  |  | Scan FOV (cm) |  |  |  | P548
64 | [11 empty] | Do NOT exceed 190 |  | 0 | [9 empty] | Display FOV (cm) |  |  |  | P549
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Criteria: | kVp accurate to within 5% of indicated. | [10 empty] | Reconstruction Algorithm |  |  |  | P550
66 | Room Number: | [6 empty] | Survey ID: | [15 empty] | Axial/Helical |  | Change | A | P551
1 | [9 empty] | Medical University of South Carolina | [13 empty] | Z-Axis collimation |  | Change | 280 | P552
2 | [9 empty] | Radiation Oncology IGRT Compliance Inspection | [13 empty] | # of channels |  | Change | 1 | P553
3 |  |  |  |  | Exposure/Dose Measurement | [18 empty] | Table Increment/speed |  | Change | 280 | P554
4 | [23 empty] | Pitch |  | Change | 1 | P555
5 | [23 empty] | Recon scan width (mm) |  |  |  | P556
6 | [23 empty] | Recon scan interval (mm) |  |  |  | P557
7 | [23 empty] | Dose reduction techniques |  |  |  | P558
8 | [15 empty] | Generator Performance | [7 empty] | Indiated CTDIvol (mGy) |  |  |  | P559
9 | [23 empty] | Adult Chest
10 | [11 empty] | Test Equipment Used: | Piranha CB2-19020491 | [10 empty] | Protocol name |  | Change | 125 kV Full fan | Q542
11 | [23 empty] | kVp |  | Change | 125 | Q543
12 | [10 empty] | Large Focus kV, Exposure, Timer Reproducibility | [12 empty] | mA |  | Change | 60 | Q544
13 | [23 empty] | mAs |  | Change | 1.2 | Q545
14 | [10 empty] | Test | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values |  |  |  |  | Eff mAs |  |  |  | Q546
15 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  |  | Time/rotation (s) |  |  |  | Q547
16 | [11 empty] | 100 | Large |  | 80 | 20 | [7 empty] | Scan FOV (cm) |  |  |  | Q548
17 | [11 empty] | 100 | Large |  | 80 | 20 | [7 empty] | Display FOV (cm) |  |  |  | Q549
18 | [11 empty] | 100 | Large |  | 80 | 20 | [7 empty] | Reconstruction Algorithm |  |  |  | Q550
19 | [11 empty] | 100 | Large |  | 80 | 20 | [7 empty] | Axial/Helical |  | Change | A | Q551
20 | [11 empty] | 100 | Large |  | 80 | 20 | [7 empty] | Z-Axis collimation |  | Change | 280 | Q552
21 | [23 empty] | # of channels |  | Change | 1 | Q553
22 | [23 empty] | Table Increment/speed |  | Change | 280 | Q554
23 | [23 empty] | Pitch |  | Change | 1 | Q555
24 | [23 empty] | Recon scan width (mm) |  |  |  | Q556
25 | [23 empty] | Recon scan interval (mm) |  |  |  | Q557
26 | [10 empty] | Note: Type comments in the below shaded cell. | [12 empty] | Dose reduction techniques |  |  |  | Q558
27 | [11 empty] | Do NOT exceed 190 | [11 empty] | Indiated CTDIvol (mGy) |  |  |  | Q559
28 | [11 empty] | New |  | 0 | [9 empty] | Adult Head
29 | [11 empty] | Do NOT exceed 190 | [6 empty] | kVp | mGy | Time |  |  | Protocol name |  |  |  | R542
30 | [11 empty] | New |  | 0 |  |  |  | Acceptable: | NA | NA | NA |  |  | kVp |  |  |  | R543
31 | [10 empty] | Small Focus kV, Exposure, Timer Reproducibility | [12 empty] | mA |  |  |  | R544
32 | [23 empty] | mAs |  |  |  | R545
33 | [10 empty] | Test | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values |  |  |  |  | Eff mAs |  |  |  | R546
34 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy |  |  | Time/rotation (s) |  |  |  | R547
35 | [11 empty] | 100 | Small |  | 90 | 10 | [7 empty] | Scan FOV (cm) |  |  |  | R548
36 | [11 empty] | 100 | Small |  | 90 | 10 | [7 empty] | Display FOV (cm) |  |  |  | R549
37 | [11 empty] | 100 | Small |  | 90 | 10 | [7 empty] | Reconstruction Algorithm |  |  |  | R550
38 | [11 empty] | 100 | Small |  | 90 | 10 | [7 empty] | Axial/Helical |  |  |  | R551
39 | [11 empty] | 100 | Small |  | 90 | 10 | [7 empty] | Z-Axis collimation |  |  |  | R552
40 | [23 empty] | # of channels |  |  |  | R553
41 |  | LFS Output Slope: |  |  | SFS Output Slope: |  |  | Overall Slope: | [15 empty] | Table Increment/speed |  |  |  | R554
42 | [23 empty] | Pitch |  |  |  | R555
43 | [23 empty] | Recon scan width (mm) |  |  |  | R556
44 | [23 empty] | Recon scan interval (mm) |  |  |  | R557
45 |  |  |  |  | Image Receptor | [18 empty] | Dose reduction techniques |  |  |  | R558
46 | [11 empty] | Do NOT exceed 190 | [11 empty] | Indiated CTDIvol (mGy) |  |  |  | R559
47 | Indicated Field Size @SID |  | Measured* Radiation |  | Field Size Indicators |  | Acceptable |  |  |  |  | New |  | 0 | [9 empty] | CTDI Measurements
48 | (cm) |  | (cm) |  | Error (% SID) |  | Field Size Indicators |  |  |  |  | Do NOT exceed 190 | [11 empty] | Adult Abdomen |  |  |  | O577
49 | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal | Transverse | Longitudinal |  |  |  | New |  | 0 | [5 empty] | kVp | mGy |  |  | Adult Chest |  |  |  | P577
50 | 5 | 5 |  |  | NA | NA | NA | NA | [10 empty] | Acceptable: | NA | NA |  |  | Adult Head |  |  |  | Q577
51 | 10 | 10 |  |  | NA | NA | NA | NA |  |  |  | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%. | [10 empty] | CTDIvol/100 mAs
52 | *Dimension of highest optical density region | [22 empty] | Adult Abdomen |  |  |  | O583
53 | **Dimension between the image of the coins marking the edge of the light field | [22 empty] | Adult Chest |  |  |  | P583
54 | [23 empty] | Adult Head |  |  |  | Q583
55 | Criteria: | The indicated field length/width are accurate to within 2% of the SID. | [9 empty] | Combined Comments:
56 |  | X-ray field and light field length/width are congruent within 2% SID. | [9 empty] | Large Focus:
57
58 | [11 empty] | Small Focus:
59
60
61 | [11 empty] | Test Equipment Used: | Piranha CB2-19020491
62
63
64 | [10 empty] | Timer performance (Large Focal Spot)
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | NOTE: Enter "Dent" for Dental, "1*" for 1-phase, "3" for 3-phase, "HF" for High Frequency, "CP" for Constant Potential
1 | [9 empty] | Medical University of South Carolina |  | Generator Type: | HF | HF | (High Frequency)
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | 2 | 2 | Timer Test is Valid for this generator
3 | [10 empty] | Test | SDD | Focal | Filter Pack/ |  | Set Values* |  |  | Measured Values*
4 |  |  |  |  | X-RAY TUBE | [5 empty] | Setup | cm | Spot | Sensitivity | kV | mA/mAs | seconds | kV (max) | kV (avg) | seconds | mGy
5 | [11 empty] | 100 | Large |  | 80 | 2 | 0.01
6 | [11 empty] | 100 | Large |  | 80 | 4 | 0.02
7 | Focal Spot Resolution | [10 empty] | 100 | Large |  | 80 | 8 | 0.04
8 | [8 empty] | 1 Year | 2 Year |  | 100 | Large |  | 80 | 20 | 0.1
9 | Large: |  | mm |  |  |  |  | Measured | Previous | Previous |  | 100 | Large |  | 80 | 50 | 0.25
10 | [7 empty] | Resolution | Resolution | Resolution |  | 100 | Large |  | 80 | 80 | 0.4
11 |  | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm)
12 |  |  |  |  | 150 | 100 | 3
13
14 | Small: |  | mm |  |  |  |  | Measured | Previous | Previous
15 | [7 empty] | Resolution | Resolution | Resolution |  | *Record time in seconds for 3-phase, HF, and CP
16 |  | kVp | mAs | time | SID* | OFD** | Magnification | (lp/mm) | (lp/mm) | (lp/mm) |  | For 1-phase, record time in pulses up to 0.5 sec. (60 pulses), and seconds thereafter.
17 |  |  |  |  | 150 | 100 | 3
18 |  | * SID = Source to Image (film) Distance. | [9 empty] | Comments:
19 |  | ** OFD = Object (resolution pattern) to Film Distance. | [5 empty] | Acceptable: | TBD |  | Note: Type comments in the above shaded cell.
20 | [11 empty] | Additional comments:
21 |  | Criteria: | No significant change in resolution since previous measurement. | [8 empty] | New |  | 0
22 | [15 empty] | Time (sec) |  | Error | 3-Phase | 1-Phase | Dent
23 | Comments: | [14 empty] | Indicated | Measured | (Percent) | Acceptable | Acceptable | Acceptable
24 | [15 empty] | 0.01
25 | [15 empty] | 0.02
26 | [15 empty] | 0.04
27 | [15 empty] | 0.1
28 | [15 empty] | 0.25
29 |  |  |  |  | GENERATOR PERFORMANCE | [10 empty] | 0.4
30 | [13 empty] | Combined Comments
31
32 |  | Test Equipment Used: | Piranha CB2-19020491 |  |  | Calibration Date: |  |  | Calibration Due:
33
34 | kVp Accuracy | [10 empty] | Timer Specifications
35 | [11 empty] | Generator |  |  |  | Generator
36 | [11 empty] | Type | time (ms) | error |  | Type | mAs | error
37 |  | Large Focal Spot |  |  |  |  | Small Focal Spot |  |  |  |  | 1f | < =4 pulses | 0 | pulses | Cine | < 2 | 15%
38 | [12 empty] | 5-24 pulses | 1 | pulse |  | 2 - 5 | 10%
39 | SDD: | 100 | cm |  |  | SDD: | 100 | cm |  |  |  |  | 25-60 pulses | 2 | pulses |  | > 5 | 5%
40 | mA: |  |  |  |  | mA: | [6 empty] | >60 pulses | 3 | pulses | mAs timer | < 2 | 20%
41 | mAs: | 20 |  |  |  | mAs: | 10 |  |  |  |  | 3f | < 2 | 25% |  |  | 2 - 5 | 15%
42 | time (s): |  |  |  |  | time (s): | [6 empty] | 2 - 10 | 20% |  |  | 5 - 100 | 10%
43 | *SDD=Source to detector distance |  |  |  |  | *SDD=Source to detector distance | [6 empty] | 10 - 50 | 15% |  |  | > 100 | 5%
44 | [12 empty] | 50 - 100 | 10%
45 | Indicated | Measured |  | kV (max) |  | Indicated | Measured |  | kV (max) |  |  |  | > 100 | 5%
46 | kVp | kV (max) | kV (avg) | % error |  | kVp | kV (max) | kV (avg) | % error |  |  | Dental | All | +2 | pulses
47 | 60 |  |  |  |  | 50 | [6 empty] | All | +15%
48 | 80 |  |  |  |  | 70
49 | 100 |  |  |  |  | 90
50 | 120 |  |  |  |  | 110
51 | 140 |  |  |  |  | 130 | [5 empty] | Test Equipment Used: | Piranha CB2-19020491
52
53 | [10 empty] | Large Focus Linearity
54
55
56 |  |  | Acceptable: | TBD |  |  |  | Acceptable: | TBD |  | mA/mAs | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
57 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
58 |  | Criteria: | kVp accurate to within 5% of indicated. | [7 empty] | Adjacent* | 100 | Large |  | 80 | 20
59 | [10 empty] | Adjacent* | 100 | Large |  | 80 | 25
60 | [10 empty] | Non-Adj. | 100 | Large |  | 80 | 2.5
61 | Comments: | [9 empty] | Non-Adj. | 100 | Large |  | 80 | 40
62 | [10 empty] | Non-Adj. | 100 | Large |  | 80 | 12.5
63 | [10 empty] | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
64 | [19 empty] | Acceptable
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Comments: | [7 empty] | Adjacent
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | Do NOT exceed 390 |  | 0 | [5 empty] | TBD
1 | [9 empty] | Medical University of South Carolina |  | New | [7 empty] | Non-Adj.
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | [9 empty] | TBD
3 | [10 empty] | Small Focus Linearity
4
5 |  |  |  |  | Generator Performance | [5 empty] | mA/mAs | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
6 | [10 empty] | Setup | cm | Spot | Sensitivity | kV | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
7 |  | Test Equipment Used: | Piranha CB2-19020491 | [7 empty] | Adjacent* | 100 | Small |  | 80 | 10
8 | [10 empty] | Adjacent* | 100 | Small |  | 80 | 11
9 | kVp and Exposure Reproducibility | [9 empty] | Non-Adj. | 100 | Small |  | 80 | 25
10 | [10 empty] | Non-Adj. | 100 | Small |  | 80 | 2.5
11 |  |  | Large Focal Spot |  |  |  | Small Focal Spot |  |  |  | Non-Adj. | 100 | Small |  | 80 | 16
12 |  |  | SDD*: | 100 | cm |  | SDD*: | 100 | cm |  | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
13 |  |  | mA: |  |  |  | mA: | [12 empty] | Acceptable
14 |  |  | mAs: | 20 |  |  | mAs: | 10 |  |  |  | Comments: | [7 empty] | Adjacent
15 |  |  | time (s): |  |  |  | time (s): |  |  |  |  | Do NOT exceed 390 |  | 0 | [5 empty] | TBD
16 | [11 empty] | New | [7 empty] | Non-Adj.
17 |  |  | Indicated | Measured |  |  | Indicated | Measured | [11 empty] | TBD
18 |  |  | kVp | kVp | mGy |  | kVp | kVp | mGy
19 |  |  | 80 |  |  |  | 90
20
21 | [11 empty] | Coefficient of Linearity | [7 empty] | Combined Focal Spot Linearity
22 | [12 empty] | Adjacent: |  | (mGy/mAs)adj1 - (mGy/mAs)adj2
23 | [14 empty] | (mGy/mAs)adj1 + (mGy/mAs)adj2 |  |  |  |  | Acceptable
24 | [19 empty] | Non-Adj.
25 | [12 empty] | Non-Adjacent: |  | (mGy/mAs)max - (mGy/mAs)min |  |  |  |  | NA
26 | [14 empty] | (mGy/mAs)max + (mGy/mAs)min
27
28 | [11 empty] | Criteria: | Coefficient of linearity <= 10% (Both adjacent and non-adjacent).
29
30 |  |  | mean: |  |  |  | mean:
31 |  |  | std dev: |  |  |  | std dev:
32 |  |  | Coeff. of var. | TBD | TBD |  | Coeff. of var. | TBD | TBD
33 | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA
34 |  |  | Acceptable: | NA | NA |  | Acceptable: | NA | NA | [5 empty] | Generator Functions
35 | [12 empty] | Means provided to indicate technique factors prior to exposure.
36 |  | *SDD = Source to Detector Distance | [10 empty] | All technique indicators functional and are free of excessive wear.
37 | [12 empty] | Technique overload circuits function to block excessive technique
38 |  |  | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%. | [6 empty] | 3 |  | Anode heat calculator/indicator is functioning properly
39 | [12 empty] | Indicator of the individual x-ray tube selected present and operational
40 | [12 empty] | Visible "beam on" indicator present and operational
41 | [12 empty] | Audible exposure end indicator present and operational
42 | [12 empty] | Exposure delay time is not excessive.
43 | Comments:
44
45 | [10 empty] | 2 | Change to DES Standards?
46 | [10 empty] | (Type 1 to base acceptance on DES criteria)
47 | [17 empty] | Measured HVL based on:
48 | [10 empty] | Half Value Layer Measurement | [7 empty] | Avg kV
49
50 | [11 empty] | SDD | Focal | Filter Pack/ |  |  | Set |  | Measured values
51 | [10 empty] | mm Al | cm | Spot | Sensitivity | kV | mA/mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy
52 | [10 empty] | 0 | 100 | Large |  | 80 | 20
53 | [10 empty] | 0 | 100 | Large |  | 80 | 20
54 | [10 empty] | 0 | 100 | Large |  | 80 | 20
55 | [10 empty] | 0 | 100 | Large |  | 80 | 20
56
57
58
59
60
61
62
63
64 | [11 empty] | Comments:
65 | Date: | [6 empty] | Inspector: | Eugene Mah |  |  | Do NOT exceed 390 |  | 0
66 | Room Number: | [6 empty] | Survey ID: |  |  |  | New
1 | [9 empty] | Medical University of South Carolina | Data Processing: |  |  |  | Avg kV
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection | mm Al. | Avg. Exp. | DATA (1/2) | Data Al | kVp |  | HVL @80 kVp (Avg kV) |  | Criteria: MUSC STANDARD |  |  | Last Year:
3 | [10 empty] | 0 |  |  |  | 80 |  | #DIV/0! |  | HVL | MIN. HVL | MAX. HVL
4 | [10 empty] | 0 | [6 empty] | kVp | (mm Al) | (mm Al) | (mm Al)
5 |  | Test Equipment Used: | Piranha CB2-19020491 | [12 empty] | Acceptable (mm Al) |  | 60 | 2.1855 | 1.9 | 2.4
6 | [13 empty] | Data Exp | LN(EXP) | LOW: | 2.6 | 70 | 2.6109 | 2.3 | 2.9
7 | [15 empty] | DESIRED: | 3 | 80 | 3.0364 | 2.6 | 3.4
8 | [5 empty] | Timer performance (Large Focal Spot) | [9 empty] | HIGH: | 3.4 | 90 | 3.4618 | 3 | 3.8
9 | [10 empty] | #DIV/0! | #DIV/0! | [5 empty] | 100 | 3.8873 | 3.4 | 4.3
10 | [15 empty] | Acceptable: | TBD | 110 | 4.3127 | 3.8 | 4.8
11 |  |  | Generator Type: | HF | (High Frequency) | [10 empty] | Meets minimum HVL criteria: | TBD | 120 | 4.7382 | 4.2 | 5.2
12 | [17 empty] | 130 | 5.1636 | 4.6 | 5.7
13 |  |  |  |  | Accuracy |  |  |  | Reproducibility
14
15 |  | SDD: | 100 | cm |  |  |  | SDD: | 100 | cm
16 |  | kVp: | 80 |  |  |  |  | kVp: | 80
17 | [7 empty] | mA: |  |  | Output - Large Focal Spot | [5 empty] | Calculated |  | (Using |  |  | Using | [9 empty] | Using
18 | [7 empty] | mAs: | 20 | [6 empty] | Current | Previous |  | Previous | (Using | (Using | Current | [8 empty] | (Using | Current
19 | [10 empty] | Previous* Measurements |  |  | Current Meas. |  | mGy/mAs | mGy/mAs |  | mGy/mAs | Recorded | Current | mGy/mAs | [8 empty] | Current | mGy/mAs
20 | Measured Values |  |  | Time (sec) |  | Error |  |  | Time (sec) |  | SDD (in.): |  |  | SDD (cm): | 100 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current | [8 empty] | Measured | @ Current
21 | kV (max) | kV (avg) | mA/mAs | Indicated | Measured | (Percent) | Acceptable |  | Indicated | Measured |  | mGy/mAs | mGy/mAs |  | mGy/mAs | kVp | meas. kVp |  | 40" SID) | kVp) | kVp) | kVp | [8 empty] | kVp) | kVp
22 |  |  | 2 | 0.01 |  |  | TBD |  |  |  | kVp | @SDD | @40" | kV (avg) | @SDD | @40" | @40" |  | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40" | [8 empty] | LN(kVp) | @40"
23 |  |  | 4 | 0.02 |  |  | TBD
24 |  |  | 8 | 0.04 |  |  | TBD
25 |  |  | 20 | 0.1 |  |  | TBD
26 |  |  | 50 | 0.25 |  |  | TBD
27 |  |  | 80 | 0.4 |  |  | TBD
28
29
30
31 | [10 empty] | *From previous compliance report |  |  |  | Max | Median | Min.
32 | [14 empty] | NA | NA | NA |  | Acceptable: | NA
33 | [8 empty] | mean: |  |  |  |  | Last Year:
34 | [8 empty] | std dev: |  | Output - Small Focal Spot | [5 empty] | Calculated |  | (Using |  |  | Using
35 | [8 empty] | Coeff. of var.: | TBD | [5 empty] | Current | Previous |  | Previous | (Using | (Using | Current
36 | [10 empty] | Previous* Measurements |  |  | Current Meas. |  | mGy/mAs | mGy/mAs |  | mGy/mAs | Recorded | Current | mGy/mAs
37 | [5 empty] | Acceptable: | NA |  | Acceptable: | NA | SDD (in.): |  |  | SDD (cm): | 100 | @ Current | @ Current |  | Adjusted to | Previous | Measured | @ Current
38 | Comments: | [10 empty] | mGy/mAs | mGy/mAs |  | mGy/mAs | kVp | meas. kVp |  | 40" SID) | kVp) | kVp) | kVp
39 | [10 empty] | kVp | @SDD | @40" | kV (avg) | @SDD | @40" | @40" |  | LN(mGy/mAs) | LN(kVp) | LN(kVp) | @40"
40
41
42
43
44
45 |  | Timer Specifications
46 |  | Generator |  |  |  | Generator |  |  |  |  | *From previous compliance report |  |  |  | Max | Median | Min.
47 |  | Type | time (ms) | error |  | Type | mAs | error | [6 empty] | NA | NA | NA |  | Acceptable: | NA
48 |  | 1f | < =4 pulses | 0 | pulses | Cine | < 2 | 15% | [5 empty] | Last Year:
49 |  |  | 5-24 pulses | 1 | pulses |  | 2 - 5 | 10%
50 |  |  | 25-60 pulses | 2 | pulses |  | > 5 | 5% |  |  | Enter 1 for YES, 2 for NO, 3 for NA
51 |  |  | >60 pulses | 3 | pulses | mAs timer | < 2 | 20% | [6 empty] | Radiation Safety
52 |  | 3f | < 2 | 25% |  |  | 2 - 5 | 15% |  |  |  |  | Continuous pressure on exposure switch required to make an exposure with selected tube
53 |  |  | 2 - 10 | 20% |  |  | 5 - 100 | 10% |  |  | 3 |  | Spot film exposure blocked when there is no cassette
54 |  |  | 10 - 50 | 15% |  |  | > 100 | 5% |  |  | 3 |  | Spot film exposure blocked when serial exposures are completed
55 |  |  | 50 - 100 | 10% | [6 empty] | 3 |  | Spot film device provides acceptable centering to the image receptor
56 |  |  | > 100 | 5% | [8 empty] | Exposure switch for selected tube is secured at a safe location | [6 empty] | NEW
57 |  | Dental | All | +2 | pulses | [7 empty] | Operator has full visibility of the exposure factors and patient | [7 empty] | Half
58 |  |  | All | +15% | [6 empty] | 3 |  | Lead apron available |  | Number: | 0 | Half | 0 | Full |  | Full
59 | [10 empty] | 3 |  | Lead gloves available. |  | Number: | 0 | (Left + Right) |  |  |  | (Left + Right)
60 | [10 empty] | 3 |  | A properly designed and installed apron rack is present.
61 | [10 empty] | 3 |  | Documentation of annual protective apparel integrity inspection is available.
62 | [12 empty] | Patient restraint devices available.
63 | [10 empty] | 3 |  | Gonadal shielding is present and functional.
64 | [10 empty] | 3 |  | Radiation warning light at room entrance is functional with selected tube
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID: | [6 empty] | Physical Inspection
1 | [9 empty] | Medical University of South Carolina |  |  | System clean and free of debris.
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection |  |  | All cables free from excessive wear or restraints.
3 | [12 empty] | System free of missing or damaged components.
4 | [12 empty] | Table movements and locks acceptable
5 |  | Test Equipment Used: | Piranha CB2-19020491 | [9 empty] | Tube crane movements and locks acceptable.
6 | [12 empty] | All other movements and locks functional.
9
8 | Large Focus Linearity |  |  | kVp: | 80 | kV filter:
9 |  |  |  | SDD: | 100 | cm |  |  | Large Focus |  | COMPUTED TOMOGRAPHY
10 | Test |  | Set |  | Measured values |  |  | mGy/mAs | Coeff. of
11 | Setup | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | @ SDD | Linearity | Acceptable | Measurement Protocols
12 | Adjacent* | 20 | [7 empty] | Adjacent
13 | Adjacent* | 25 | [7 empty] | TBD |  |  |  | Abdomen | Pelvis | Head | Abdomen | Pelvis | Head
14 | Non-Adj. | 2.5 | [10 empty] | Protocol name: | 125 kV Half Fan | 125 kV Full fan |  | 125 kV Half Fan | 125 kV Full fan
15 | Non-Adj. | 40 | [7 empty] | Non-Adj. |  |  | kVp: | 125 | 125 |  | 125 | 125
16 | Non-Adj. | 12.5 | [7 empty] | TBD |  |  | mA: | 60 | 60 |  | 60 | 60
17 | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X. | [11 empty] | mAs/pulse: | 1.2 | 1.2 |  | 1.2 | 1.2
18 | [12 empty] | mAs:
19 | Comments: | [11 empty] | Rotation time (s):
20 | [12 empty] | Scan FOV (cm):
21 | [12 empty] | Display FOV (cm):
22 | [12 empty] | Reconstruction algorithm:
23 | Small Focus Linearity |  |  | kVp: | 80 | kV filter: | [6 empty] | Axial (A)/Helical (H): | A | A |  | A | A
24 |  |  |  | SDD: | 100 | cm |  |  | Small Focus |  |  |  | Z-Axis collimation (T): | 280 | 280 |  | 280 | 280
25 | Test |  | Set |  | Measured values |  |  | mGy/mAs | Coeff. of |  |  |  | # of channels (N): | 1 | 1 |  | 1 | 1
26 | Setup | mAs | time(SEC) | kV (max) | kV (avg) | time(SEC) | mGy | @ SDD | Linearity | Acceptable |  |  | Table Increment/speed (I): | 280 | 280 |  | 280 | 280
27 | Adjacent* | 10 | [7 empty] | Adjacent |  |  | Pitch (p): | 1 | 1 |  | 1 | 1
28 | Adjacent* | 11 | [7 empty] | TBD |  |  | Recon scan width (mm):
29 | Non-Adj. | 25 | [10 empty] | Recon scan interval (mm):
30 | Non-Adj. | 2.5 | [7 empty] | Non-Adj. |  |  | Dose reduction techniques:
31 | Non-Adj. | 16 | [7 empty] | TBD |  |  | Indiated CTDIvol (mGy):
32 | *For mAs generators, adjacent means two settings that differ by a factor of approximately 2X.
33 | [10 empty] | CTDI Measurements
34 | Comments: | [10 empty] | Test Equipment Used: | Piranha CB2-19020491
35 | [11 empty] | Chamber length (mm): | 100
36 | [11 empty] | Chamber correction factor: | 1
37
38 | Combined Focal Spot Linearity | [11 empty] | Abdomen | Pelvis | Head |  |  |  |  | Reference | Pass/Fail
39 |  |  |  | Combined | [7 empty] | Indicated CTDIvol (mGy): | [7 empty] | (mGy) | (mGy)
40 |  | mGy/mAs |  | Coeff. of | [7 empty] | Center (mGy) | [6 empty] | Adult Abdomen | 25 | 30
41 |  | @ SDD |  | Linearity |  | Acceptable | [12 empty] | Adult Head: | 75 | 80
42 |  | Max.: | 0 | Non-Adj. |  | Non-Adj. | [12 empty] | Ped Abd (16cm) | 15 | 20
43 |  | Min.: | 0 | NA |  | NA | [5 empty] | Center average (mGy): | [6 empty] | Ped Abd (32cm) | 7.5 | 10
44 | [10 empty] | Periphery (mGy) | Top: | [6 empty] | Pediatric Head: | 35 | 40
45 |  | Coefficient of Linearity | [9 empty] | Bottom:
46 |  |  | Adjacent: |  | (mGy/mAs)adj1 - (mGy/mAs)adj2 | [6 empty] | Left:
47 |  |  |  |  | (mGy/mAs)adj1 + (mGy/mAs)adj2 | [6 empty] | Right:
48 | [11 empty] | Periphery average (mGy): | #DIV/0! | #DIV/0! | #DIV/0!
49 |  |  | Non-Adjacent: |  | (mGy/mAs)max - (mGy/mAs)min | [6 empty] | Planar average dose f(0):
50 |  |  |  |  | (mGy/mAs)max + (mGy/mAs)min | [6 empty] | Phantom size used (cm) | 32 | 32
51 | [11 empty] | Previous year:
52 |  | Criteria: | Coefficient of linearity <= 10% (Both adjacent and non-adjacent). | [8 empty] | Variation from prev year:
53 | [11 empty] | Variation from indicated:
54 | Rule Number |  |  |  | Generator Functions |  |  |  | Compliance |  |  | Acceptable:
55 | DES RHB 4.3.9.1 |  | Means are provided to indicate technique factors prior to exposure. | [5 empty] | TBD |  |  | CTDIvol/100 mAs:
56 |  |  | All technique indicators functional and are free of excessive wear. | [5 empty] | TBD |  |  | Prev CTDIvol/100 mAs:
57 |  |  | Technique overload circuits function to block excessive technique. | [5 empty] | TBD |  |  | Criteria: | CTDI is within 5% of previous year.
58 |  |  | Anode heat calculator/indicator is functioning properly. | [5 empty] | NA |  |  |  | Measured CTDI is within 20% of indicated CTDI
59 | DES RHB 4.3.7 |  | Indicator of the individual x-ray tube selected present and operational. | [5 empty] | TBD
60 | DES RHB 4.3.9.3 |  | Visible "beam on" indicator present and operational. | [5 empty] | TBD
61 | DES RHB 4.7.4.2.4 |  | Audible exposure end indicator present and operational. | [5 empty] | TBD
62 |  |  | Exposure delay time is not excessive. | [5 empty] | TBD
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 |  |  |  |  | HALF VALUE LAYER MEASUREMENT
4
5
6 |  |  | Large Focal Spot
7 |  | SDD*: | 100 | cm |  |  | mA:
8 |  | Set kVp: | 80 |  |  |  | mAs: | 20
9 |  |  |  |  | Set |  | time (s):
10 |  | Al (mm) | mGy | mGy | kVp
11 |  | 0 |  |  | 80 |  | kVp (Avg kV) | 80
12 |  | 0 |  |  |  |  | Meas. HVL: | #DIV/0!
13 | [6 empty] | Meets min HVL: | TBD
14
15 | [6 empty] | Criteria: MUSC STANDARD
16 | [6 empty] | HVL | MIN. HVL | MAX. HVL
17 |  | * SDD = Source to Detector Distance. |  |  |  | kVp | (mm Al) | (mm Al) | (mm Al)
18 | [5 empty] | 60 | 2.1855 | 1.9 | 2.4
19 | [5 empty] | 70 | 2.6109 | 2.3 | 2.9
20 |  |  | **ACCEPTABLE RESULTS (mm Al) |  |  | 80 | 3.0364 | 2.6 | 3.4
21 |  |  |  | Min. acceptable HVL at 80 kvp: | 2.6 | 90 | 3.4618 | 3 | 3.8
22 |  |  |  | Desired HVL at meas. 80 kVp: | 3 | 100 | 3.8873 | 3.4 | 4.3
23 |  |  |  | Max. acceptable HVL at 80 kvp: | 3.4 | 110 | 4.3127 | 3.8 | 4.8
24 | [5 empty] | 120 | 4.7382 | 4.2 | 5.2
25 | [5 empty] | 130 | 5.1636 | 4.6 | 5.7
26 | Comments:
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | Output - Large Focal Spot |  |  | Current | Previous
4 |  |  |  | mGy/mAs | mGy/mAs
5 |  |  | Current | @ Current | @ Current
6 |  |  | Meas. | kVp | kVp | Percent
7 |  |  | kVp | @40" | @40" | Change | Acceptable
8
9
10
11
12
13
14
15 | [7 empty] | Acceptable: | NA
16 | Output - Small Focal Spot |  |  | Current | Previous
17 |  |  |  | mGy/mAs | mGy/mAs
18 |  |  | Current | @ Current | @ Current
19 |  |  | Meas. | kVp | kVp | Percent
20 |  |  | kVp | @40" | @40" | Change | Acceptable
21
22
23
24
25
26 | [7 empty] | Acceptable: | NA
27
28 |  | Criteria: | The x-ray output @40" within +/-25% from previous measurements.
29 | Comments:
30
31
32
33 |  |  | Lrg. Focus |  |  |  |  | Sm. Focus
34 |  |  | Current |  |  |  |  | Current
35 |  |  | mGy/mAs |  |  |  |  | mGy/mAs
36 |  | kVp | @40" |  |  |  | kVp | @40"
37 |  | 55 |  |  |  |  | 45
38 |  | 60 |  |  |  |  | 55
39 |  | 70 |  |  |  |  | 65
40 |  | 80 |  |  |  |  | 75
41 |  | 90 |  |  |  |  | 85
42 |  | 100 |  |  |  |  | 95
43 |  | 110 |  |  |  |  | 105
44 |  | 120 |  |  |  |  | 115
45 |  | 130 |  |  |  |  | 125
46 |  | 140 |  |  |  |  | 135
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
1 | [9 empty] | Medical University of South Carolina
2 |  |  |  |  | Measurement Data |  |  |  |  | Radiation Oncology IGRT Compliance Inspection
3 | Protocols |  | Abdomen | Pelvis | Head
4 |  | Protocol name: | 125 kV Half Fan | 125 kV Full fan
5 |  | kVp: | 125 | 125
6 |  | mA: | 60 | 60
7 |  | mAs/pulse: | 1.2 | 1.2
8 |  | mAs:
9 |  | Rotation time (s):
10 |  | Scan FOV (cm):
11 |  | Display FOV (cm):
12 |  | Reconstruction algorithm:
13 |  | Axial (A)/Helical (H): | A | A
14 |  | Z-Axis collimation (T): | 280 | 280
15 |  | # of channels (N): | 1 | 1
16 |  | Table Increment/speed (I): | 280 | 280
17 |  | Pitch (p): | 1 | 1
18 |  | Recon scan width (mm):
19 |  | Recon scan interval (mm):
20 |  | Dose reduction techniques:
21 |  | Indiated CTDIvol (mGy):
22 | CTDI Measurements
23 |  | Test Equipment Used: | Piranha CB2-19020491
24
25
26 |  |  |  | Abdomen | Pelvis | Head |  |  | Reference | Pass/Fail
27 |  |  | Indicated CTDIvol (mGy): | [5 empty] | (mGy) | (mGy)
28 |  |  | Center (mGy): |  |  |  |  | Adult Abdomen | 25 | 30
29 |  |  | Periphery (mGy): | #DIV/0! | #DIV/0! | #DIV/0! |  | Adult Head: | 75 | 80
30 |  |  | Planar average dose f(0): |  |  |  |  | Ped Abd (16cm) | 15 | 20
31 |  |  | Phantom size used (cm): | 32 | 32 |  |  | Ped Abd (32cm) | 7.5 | 10
32 |  |  | Variation from previous year: |  |  |  |  | Pediatric Head: | 35 | 40
33 |  |  | Variation from indicated CTDI:
34 |  |  | Acceptable:
35 |  |  | f(0)/100 mAs:
36 |  |  | Criteria: | CTDI is within 5% of previous year.
37 |  |  |  | Measured CTDI is within 20% of indicated CTDI
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [6 empty] | Inspector: | Eugene Mah
66 | Room Number: | [6 empty] | Survey ID:
Note on U464: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on H609: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.

## Tables
Huttner Grating |  |  | RADIOGRAPHIC UNITS (MUSC CRITERIA) |  |  |  | RAD / DENTAL UNITS (DHEC CRITERIA)
1.1 | 0.5 |  | Criteria: |  |  |  | Criteria:
1.2 | 0.56 |  | HVL | MIN. HVL | MAX. HVL |  | HVL | MIN. HVL | MAX. HVL
1.3 | 0.63 | kVp | (mm Al) | (mm Al) | (mm Al) | kVp | (mm Al) | (mm Al) | (mm Al)
2.1 | 0.71 | 60 | 2.1855 | 1.9 | 2.4 | 50 | 1.5 | 1.5 | 1.7
2.2 | 0.8 | 70 | 2.6109 | 2.3 | 2.9 | 60 | 1.5 | 1.5 | 1.7
2.3 | 0.9 | 80 | 3.0364 | 2.6 | 3.4 | 70 | 1.5 | 1.5 | 1.7
3.1 | 1 | 90 | 3.4618 | 3 | 3.8 | 80 | 2.3 | 2.3 | 2.7
3.2 | 1.12 | 100 | 3.8873 | 3.4 | 4.3 | 90 | 2.5 | 2.5 | 3
3.3 | 1.25 | 110 | 4.3127 | 3.8 | 4.8 | 100 | 2.7 | 2.7 | 3.2
4.1 | 1.4 | 120 | 4.7382 | 4.2 | 5.2 | 110 | 3 | 3 | 3.5
4.2 | 1.6 | 130 | 5.1636 | 4.6 | 5.7 | 120 | 3.2 | 3.2 | 3.75
4.3 | 1.8
5.1 | 2
5.2 | 2.24
5.3 | 2.5
6.1 | 2.8
6.2 | 3.15
6.3 | 3.55
7.1 | 4
7.2 | 4.5
7.3 | 5
RTI Piranha
Piranha 1 | Piranha 1 CB2-17090320
Piranha 2 | Piranha 2 CB2-19020491
RTI DAP chamber
DAP 1 | DAP 1 2501554
DAP 2 | DAP 2 2501555
Note on E2: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.
Note on J2: To obtain the minimum and maximum acceptable values for the HVL, I first did a trend function on the data in the acceptance test manual for HVL vs kVp, and from that trend data determined the mm Al corresponding to kVp values 10% higher, and 10% lower that the nominal kvp. The calculated HVL values then became the new Min and Max HVL values for that nominal kVp.

## Sheet1
Measurements | [14 empty] | 125 kV Half Fan | 125 kV Full fan | 140 kV Half fan | 100 kV full fan | Measurements
[14 empty] | kV | 125 | 125 | 140 | 100
Set kV (kV) | Set mA (mA) | Set mAs (mAs) | Set time (s) | Added Filt | Focal spot | Tube voltage (kV) | Exposure time (s) | Exposure (mGy) | Exposure rate (mGy/s) | Total filtr. (mm Al) | HVL (mm Al) | Exposure (norm) (mGy) | Exposure rate (norm) (mGy/s) | mAs/pulse | 1.2 | 1.2 | 1.88 | 0.3 | Set kV (kV) | CT phantom type | Set Added filtr. (mm Al) | Exposure time (ms) | Exposure (mGy) | Exposure rate (mGy/s)
60 | 400 | 20 | 0.05 | 0 | Large | [8 empty] | Pulse mA | 60 | 60 | 75 | 15
80 | 400 | 20 | 0.05 | 0 | Large | [8 empty] | Pulse ms | 20 | 20 | 25 | 20
80 | 400 | 20 | 0.05 | 0 | Large | [8 empty] | Pulse rate | 15 | 15 | 15 | 15
80 | 400 | 20 | 0.05 | 0 | Large | [8 empty] | X1 | -24.7 | -14 | -24.7 | -14
80 | 400 | 20 | 0.05 | 0 | Large | [8 empty] | X2 | 3.3 | 14 | 3.3 | 14
80 | 400 | 20 | 0.05 | 0 | Large | [8 empty] | Y1 | -10.7 | -10.7 | -10.7 | -10.7
80 | 400 | 20 | 0.05 | 3 | Large | [8 empty] | Y2 | 10.7 | 10.7 | 10.7 | 10.7
80 | 400 | 20 | 0.05 | 3 | Large
80 | 400 | 20 | 0.05 | 3.5 | Large | [8 empty] | Center
80 | 400 | 20 | 0.05 | 3.5 | Large | [8 empty] | Top
80 | 400 | 20 | 0.05 | 2.5 | Large | [8 empty] | Bottom
80 | 400 | 20 | 0.05 | 2.5 | Large | [8 empty] | Left
80 | 500 | 25 | 0.05 | 0 | Large | [8 empty] | Right
80 | 50 | 2.5 | 0.05 | 0 | Large
80 | 800 | 40 | 0.05 | 0 | Large
80 | 250 | 12.5 | 0.05 | 0 | Large
100 | 400 | 20 | 0.05 | 0 | Large
120 | 400 | 20 | 0.05 | 0 | Large
140 | 400 | 20 | 0.05 | 0 | Large
50 | 100 | 10 | 0.1 | 0 | Small
70 | 100 | 10 | 0.1 | 0 | Small
90 | 100 | 10 | 0.1 | 0 | Small
90 | 100 | 10 | 0.1 | 0 | Small
90 | 100 | 10 | 0.1 | 0 | Small
90 | 100 | 10 | 0.1 | 0 | Small
90 | 100 | 10 | 0.1 | 0 | Small
80 | 100 | 10 | 0.1 | 0 | Small
80 | 110 | 11 | 0.1 | 0 | Small
80 | 250 | 25 | 0.1 | 0 | Small
80 | 50 | 2.5 | 0.1 | 0 | Small
80 | 160 | 16 | 0.1 | 0 | Small
110 | 100 | 10 | 0.1 | 0 | Small
130 | 100 | 10 | 0.1 | 0 | Small
80 | 200 | 2 | 0.01 | 0 | Large
80 | 200 | 4 | 0.02 | 0 | Large
80 | 200 | 8 | 0.04 | 0 | Large
80 | 200 | 20 | 0.1 | 0 | Large
80 | 200 | 50 | 0.25 | 0 | Large
80 | 200 | 80 | 0.4 | 0 | Large
Waveforms

## DataPage
Sheet | RAD
SurveyID
LFS | 0
SFS | 0
Collimation table | 5 | 5
| 10 | 10
LFS Output
SFS Output
Gen Test | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
HVL | #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!
| #REF! | #REF!