## QC Test Summary-Hologic
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Lorad/Hologic
Site Name |  | HCC Breast Imaging Program | [6 empty] | Report Date
Address |  | 86 Jonathan Lucas St Suite 344, Charleston, SC | [6 empty] | Survey Date | 0
Medical Physicist's Name |  |  | Eugene Mah | [5 empty] | Signature
X-Ray Unit Manufacturer |  |  | Lorad/Hologic | [5 empty] | Model | Selenia Dimensions
Date of Installation | [8 empty] | Room ID
QC Manual Version # |  |  | MAN-03706 Rev 007 |  |  |  |  | (use any version applicable to model; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  |  | Model |  | Location |  | QC Manual Version #
|  | Review Workstation* | Sectra |  |  | IDS7 |  |  |  | 24.2
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System (www.accessdata.fda.gov/cdrh_docs/presentations/pghs/Polic_Guidance_Help_System.htm).
Survey Type |  |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [9 empty] | Annual Survey
Features |  |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Mammographic Unit Assembly Evaluation
2. Collimation Assessment
3. Artifact Evaluation
4. kVp Accuracy and Reproducibility
5. Beam Quality Assessment - HVL Measurement
6. Evaluation of System Resolution
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
| Average glandular dose for average breast is ≤3 mGy (300 mrad) (conventional) | [9 empty] | mGy
| Average glandular dose for average breast is ≤3 mGy (300 mrad) (DBT) | [9 empty] | mGy
9. Radiation Output Rate
10. Phantom Image Quality Evaluation | [5 empty] | Fibers | Specks | Masses
Phantom image scores (conventional)
Phantom image scores (DBT)
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
|  | SNR (value) | #DIV/0!
|  | CNR (value) | #DIV/0! | (required for new unit MEE and Annual Survey)
CNR should not vary by more than ±15% (NA for MEE)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
13. DICOM Printer QC (if applicable, MEE only) | [11 empty] | NA
14. Detector Flat Field Calibration (MEE only)
15. Geometry Calibration For Tomosynthsis (DBT MEE only)
16. Compression Thickness Indicator (MEE only)
17. Compression (MEE only)
18. Detector Ghosting (troubleshooting only) | [11 empty] | NA
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M22: Click in boxes to use drop-down lists
Note on I34: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Hologic
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Lorad/Hologic, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | DICOM Printer Quality Control (if applicable) | Weekly | NA
2. | Viewboxes and Viewing Conditions | Weekly
3. | Artifact Evaluation | Weekly
4. | Signal-To-Noise and Contrast-To-Noise Measurements | Weekly
5. | Phantom Image Quality Evaluation | Weekly
6. | Detector Flat-Field Calibration | Weekly
7. | Compression Thickness Indicator | Bi-weekly
8. | Visual Checklist | Monthly
9. | Repeat/Reject Analysis | Quarterly
10. | Compression | Semi-annually
11. | Geometry Calibration (Tomosynthsis Option) (DBT) | Semi-annually
12. | Diagnostic Review Workstation QC (NA if only hardcopy read) | See Hologic QC Manual
13. | Mobile Unit Quality Control (if applicable) | After every move
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use any QC manual version that is applicable to the Lorad model surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for Lorad FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## SBB QC Test Summary
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY
Site Name | HCC Breast Imaging Program |  |  |  |  | Report Date
Address | 86 Jonathan Lucas St, Charleston, SC |  |  |  |  | Survey Date | 0
X-Ray Unit Manufacturer | [5 empty] | Model
Date of Installation | [5 empty] | Room ID
Film (mfr & type) |  | NA |  |  |  | Screen (mfr & type) | NA
Film Processor Mfr |  | NA |  |  |  | Model | NA
Digital Image Receptor Mfr | [5 empty] | Model
Medical Physicist's Name |  | Eugene Mah |  |  |  | Signature
Medical Physicist's QC Tests
PASS/FAIL
1. Stereotactic Breast Biopsy Unit Assembly Evaluation
2. Collimation Assessment
A. X-ray field adequately matches image receptor
B. Biopsy window generally centered over digital image receptor (NA if film used)
3. Focal Spot Performance and System Limiting Resolution
| A. Focal spot performance acceptable (NA if digital used) | [7 empty] | NA
B. Digital system spatial resolution acceptable (NA if film used)
4. kVp Accuracy and Reproducibility
Measured average kVp within ±5% of indicated kVp
kVp coefficient of variation ≤ 0.02
5. Beam Quality Assessment (Half-Value Layer Measurement)
HVL is within acceptable lower and upper limits at all kVp values tested
6. AEC System or Manual Exposure Performance Assessment
Optical density or signal range acceptable
7. Receptor Speed Uniformity
Screen speed uniformity or digital receptor uniformity acceptable
8. Breast Entrance Exposure, Average Glandular Dose and Exposure Reproducibility
Exposure reproducibility is within acceptable limits
| Average glandular dose to a 4.2 cm thick breast is ≤ 3 mGy (300 mrad) | [6 empty] | mGy
9. Image Quality Evaluation
Phantom image quality is acceptable
| Phantom type: | ACR Mammography Accreditation Phantom |  |  |  | Mini Phantom
| Phantom image quality scores: |  | Fibers | 0 | Specks | 0 | Masses | 0
10. Artifact Evaluation
Artifacts were not apparent or not significant:
Artifacts identified:
11. Localization Accuracy Test
Localization and sampling accurate/object captured
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M15: Click in boxes to use drop-down lists

## SBB Tech QC Eval
MEDICAL PHYSICIST'S STEREOTACTIC UNIT
QC TEST SUMMARY (continued)
Evaluation of Site's Technologist QC Program
|  | Frequency | PASS/FAIL
1. | Localization Accuracy Test | Daily
2. | Darkroom Cleanliness (NA if digital used) | Daily | NA
3. | Processor Quality Control (NA if digital used) | Daily | NA
4. | Phantom Images | Weekly
5. | Screen Cleanliness (NA if digital used) | Weekly | NA
6. | Viewboxes and Viewing Conditions (NA if digital used) | Weekly
7. | Hardcopy Output Quality (if hardcopy produced from digital data) | Monthly | NA
8. | Visual Checklist | Monthly
9. | Analysis of Fixer Retention in Film (NA if digital used) | Quarterly | NA
10. | Compression | Semi-annually
11. | Repeat Analysis | Semi-annually
12. | Screen-Film Contact (NA if digital used) | Semi-annually | NA
13. | Darkroom Fog (NA if digital used) | Semi-annually | NA
14. | Zero Alignment Test (if required by manufacturer) | Before each patient | NA
15. | Any additional tests required by manufacturer | As required by manufacturer | Pass
Medical Physicist's Recommendations for Quality Improvement
Comments:
Note on J8: Click in boxes to use drop-down lists

## MQSA Requirements
MEDICAL PHYSICIST'S CHECKLIST MQSA REQUIREMENTS FOR MAMMOGRAPHY EQUIPMENT
Facility Name: | HCC Breast Imaging Program
Unit Manufacturer: |  |  | Model:
Serial number: |  |  | Year Mfr:
Medical Physicist: | Eugene Mah |  | Room ID:
Signature: |  |  | Survey Date: | 0
Feature | FDA Rule Section | Requirement | Applies to | Meets FDA Requirements? (if NA, please explain)
Motion of tube-image receptor assembly | 3(i) | The assembly shall be capable of being fixed in any position where it is designed to operate. Once fixed in any such position, it shall not undergo unintended motion. | S-F & FFDM
| 3(ii) | This mechanism shall not fail in the event of power interruption. | S-F & FFDM
Image receptor sizes | 4(i) | Systems using screen-film image receptors shall provide, at a minimum, for operation with image receptors of 18 x 24 cm and 24 x 30 cm. | S-F
| 4(ii) | Systems using screen-film image receptors shall be equipped with moving grids matched to all image receptor sizes provided. | S-F
| 4(iii) | Systems used for magnification procedures shall be capable of operation with the grid removed from between the source and image receptor. | S-F & FFDM
Beam limitation and light fields | 5(i) | All systems shall have beam limiting devices that allow the useful beam to extend to or beyond the chest wall edge of the image receptor. | S-F & FFDM
| 5(ii) | For any mammography system with a light beam that passes through the X-ray beam-limiting device, the light shall provide an average illumination of not less than 160 lux (15 ft-candles) at 100 cm or the maximum source-image receptor distance (SID), whichever is less. | S-F & FFDM (except Fischer)
Magnification | 6(i) | Systems used to perform noninterventional problem-solving procedures shall have radiographic magnification capability available for use by the operator. | S-F & FFDM
| 6(ii) | Systems used for magnification procedures shall provide, at a minimum, at least one magnification value within the range of 1.4 to 2.0. | S-F & FFDM
Focal spot selection | 7(i) | When more than one focal spot is provided, the system shall indicate, prior to exposure, which focal spot is selected. | S-F & FFDM
| 7(ii) | When more than one target material is provided, the system shall indicate, prior to exposure, the preselected target material. | S-F & FFDM
| 7(iii) | When the target material and/or focal spot is selected by a system algorithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure. | S-F & FFDM
Application of compression | 8(i)(A) | Each system shall provide an initial power-driven compression activated by hands-free controls operable from both sides of the patient. | S-F & FFDM
| 8(i)(B) | Each system shall provide fine adjustment compression controls operable from both sides of the patient. | S-F & FFDM
Compression paddle | 8(ii)(A) | Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system. | S-F & FFDM
| 8(ii)(B) | The compression paddle shall be flat and parallel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when compression is applied. | S-F & FFDM (except Fischer)
| 8(ii)(C) | Equipment intended by the manufacturer’s design to not be flat and parallel to the breast support table during compression shall meet the manufacturer’s design specifications and maintenance requirements. | S-F & FFDM
| 8(ii)(D) | The chest wall edge of the compression paddle shall be straight and parallel to the edge of the image receptor. | S-F & FFDM
| 8(ii)(E) | The chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image. | S-F & FFDM
Technique factor selection and display | 9(i) | Manual selection of mAs or at least one of its component parts (mA and/or time) shall be available. | S-F & FFDM
| 9(ii) | The technique factors (kVp and either mA and seconds or mAs) to be used during an exposure shall be indicated before the exposure begins, except when AEC is used, in which case the technique factors that are set prior to the exposure shall be indicated. | S-F & FFDM
| 9(iii) | Following AEC mode use, the system shall indicate the actual kVp and mAs (or mA and time) used during the exposure. | S-F & FFDM
Automatic exposure control | 10(i) | Each screen-film system shall provide an AEC mode that is operable in all combinations of equipment configuration provided, e.g., grid, non-grid; magnification, nonmagnification; and various target-filter combinations. | S-F
| 10(ii) | The positioning or selection of the detector shall permit flexibility in the placement of the detector under the target tissue. The size and the available positions of the detector shall be clearly indicated at the X-ray input surface of the breast compression paddle. The selected position of the detector shall be clearly indicated. | S-F
| 10(iii) | The system shall provide means for the operator to vary the selected optical density from the normal (zero) setting. | S-F
X-ray film* | 11 | The facility shall use X-ray film for mammography that has been designated by the film manufacturer as appropriate for mammography. | S-F
Intensifying screens* | 12 | The facility shall use intensifying screens for mammography that have been designated by the screen manufacturer as appropriate for mammography and shall use film that is matched to the screen’s spectral output as specified by the manufacturer. | S-F
Film processing solutions* | 13 | For processing mammography films, the facility shall use chemical solutions that are capable of developing the films used by the facility in a manner equivalent to the minimum requirements specified by the film manufacturer. | S-F
Lighting* | 14 | The facility shall make special lights for film illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the interpreting physicians. | S-F & FFDM (for hardcopy comparison)
Film masking devices* | 15 | Facilities shall ensure that film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are available to all interpreting physicians interpreting for the facility. | S-F & FFDM (for hardcopy comparison)
* NA is acceptable for new units at existing facilities if these were previously evaluated and have not changed
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA
Yes
No
NA

## Sheet1
1 | [12 empty] | Revision 2.7-20230320 | [11 empty] | Print Area
2 | [6 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [6 empty] | Page1,Page2,Artifact_Coll,AEC_ImgQual,MGD,Gen_HVL
3 | [6 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [24 empty] | All: | Page1,Page2,Artifact_Coll,AEC_ImgQual,MGD,SBB,Gen_HVL
5 | [6 empty] | Mammography System Compliance Inspection | [10 empty] | Mammography System Compliance Inspection
6 | [24 empty] | Measurement Parameter | Last Year |  | This Year
7 | [12 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah |  | Print Area |  | Change | Page1,Page2,Artifact_Coll,AEC_ImgQual,MGD,Gen_HVL
8 | [6 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [6 empty] | Date
9 |  |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  |  | Inspector |  | Change | Eugene Mah |  |  |  | kV Set | mAs Set | Add Filt | Target | Filter | kVp | ms | Dose (mGy) | Dose rate (mGy/s) | HVL (mm Al) | Corr kV
10 |  |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: | [5 empty] | Facility: | [6 empty] | 24 | 50 | 0
11 |  |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: | [5 empty] | Department: | [6 empty] | 24 | 50 | 0
12 |  |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: | [5 empty] | Area/Division: | [6 empty] | 25 | 50 | 0
13 |  |  |  | Survey ID: |  |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: | [6 empty] | 25 | 50 | 0
14 | [14 empty] | Room Number: |  |  |  | Accession Number: | [5 empty] | Room Number: | [6 empty] | 26 | 50 | 0
15 |  |  | X-Ray Generator | [21 empty] | Site Number: | [6 empty] | 28 | 20 | 0
16 |  |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [10 empty] | Authorized Use: | [6 empty] | 28 | 50 | 0
17 |  |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: | [5 empty] | Date of Installation: | [6 empty] | 28 | 50 | 0
18 |  |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number: | [5 empty] | Number of X-ray Tubes: | [6 empty] | 28 | 50 | 0
19 | [14 empty] | Max kVp: |  |  |  | Max mA: | [5 empty] | Accesssion Number: | [6 empty] | 28 | 50 | 0
20 |  |  | X-Ray Tube 1 | [21 empty] | X-Ray Generator | [6 empty] | 28 | 100 | 0
21 |  |  |  | Tube Designation/Use: |  |  |  |  | Target: |  |  |  |  | X-Ray Tube 1 |  |  |  |  | Target: | [5 empty] | Manufacturer: | [6 empty] | 28 | 300 | 0
22 |  |  |  | Manufacture Date: | [10 empty] | Tube Designation/Use: | [9 empty] | Model: | [6 empty] | 28 | 400 | 0
23 |  |  | Insert | [5 empty] | Filter: | [5 empty] | Manufacture Date: | [9 empty] | Max kVp: | [6 empty] | 30 | 50 | 0
24 |  |  |  | Manufacturer: | [9 empty] | Insert |  |  |  |  | Filter: | [5 empty] | Manufacture Date: | [6 empty] | 32 | 50 | 0
25 |  |  |  | Model: | [10 empty] | Manufacturer: | [9 empty] | Serial Number: | [6 empty] | 32 | 50 | 0
26 |  |  |  | Serial Number: |  |  |  | Focal Spot Sizes (mm) | [6 empty] | Model: | [9 empty] | Max mA: | [6 empty] | 34 | 50 | 0
27 |  |  | Housing | [5 empty] | Large: | [5 empty] | Serial Number: |  |  |  | Focal Spot Sizes (mm) | [5 empty] | X-Ray Tube 1 | [6 empty] | 34 | 50 | 0
28 |  |  |  | Manufacturer: |  |  |  |  | Small: |  |  |  |  | Housing |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: | [6 empty] | 36 | 50 | 0
29 |  |  |  | Model: |  |  |  | Filtration | Inherent: | [5 empty] | Manufacturer: |  |  |  | Small: | [5 empty] | Manufacture Date: | [6 empty] | 38 | 50 | 0
30 |  |  |  | Serial Number: |  |  |  |  | Added: | [5 empty] | Model: | [9 empty] | Manufacturer: | [6 empty] | 28 | 50 | 0
31 | [14 empty] | Serial Number: |  |  |  | Filtration | [5 empty] | Model: | [6 empty] | 30 | 50 | 0
32 | [12 empty] | Enter 1 for YES, 2 for NO | [5 empty] | Inherent: | [5 empty] | Serial Number: | [6 empty] | 32 | 50 | 0
33 | [6 empty] | Technique Chart | [6 empty] | Is this a breast tomosynthesis unit |  |  |  |  | Added: | [5 empty] | Manufacturer: | [6 empty] | 34 | 50 | 0
34 | [13 empty] | Is this unit used for diagnostic/magnification studies? | [10 empty] | Model: | [6 empty] | 36 | 50 | 0
35 |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | Stereotactic breast biopsy unit or add-on present? | [10 empty] | Serial Number: | [6 empty] | 38 | 50 | 0
36 |  | Breast | [11 empty] | Is this a new installation? | [10 empty] | Large: | [6 empty] | 28 | 50 | 0
37 |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | [6 empty] | Inspection Results | [6 empty] | Small: | [6 empty] | 30 | 50 | 0
38 |  | <3 cm | [10 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [11 empty] | Target 1: | [6 empty] | 32 | 50 | 0
39 |  | 3 to 5 cm | [11 empty] | DHEC Registration sticker is present, clearly visible and legible | [10 empty] | Target 2: | [6 empty] | 34 | 50 | 0
40 |  | 5 to 7 cm | [11 empty] | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [10 empty] | Filter 1: | [6 empty] | 36 | 50 | 0
41 |  | >7 cm | [11 empty] | Radiation warning label posted on the generator control panel | [10 empty] | Filter 2: | [6 empty] | 38 | 50 | 0
42 | [13 empty] | Operator manuals are available. | [10 empty] | Filter 3:
43 | [6 empty] | Inspection Results | [6 empty] | Monthly radiation monitoring reports are posted. | [10 empty] | Inherent:
44 |  | Rule Number | [8 empty] | Compliance |  |  | Unit installed as shown on shielding plan | [10 empty] | Added
45 |  | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD | [6 empty] | Mammography Unit Assembly Evaluation | [6 empty] | Technique Chart
46 |  | DHEC RHB 11.2.3 |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  |  | Free-standing unit is mechanically stable | [10 empty] | Fatty Breast
47 |  | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD |  |  | All moving parts move smoothly, without obstructions to motion | [10 empty] | Target/Filter:
48 |  |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | All locks and detents work properly | [10 empty] | kVp:
49 |  |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | Image receptor holder assembly is free from vibrations | [10 empty] | Density:
50 |  |  |  | Unit installed as shown on shielding plan | [6 empty] | TBD |  | 3 | Image receptor slides smoothly into holder assembly | [10 empty] | Target/Filter:
51 | [6 empty] | Mammography Unit Assembly Evaluation | [6 empty] | Image receptor is held securely by assembly in any orientation | [10 empty] | kVp:
52 |  |  |  | Free-standing unit is mechanically stable | [6 empty] | TBD |  |  | Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm | [10 empty] | Density:
53 |  |  |  | All moving parts move smoothly, without obstructions to motion | [6 empty] | TBD |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [10 empty] | Target/Filter:
54 |  |  |  | All locks and detents work properly | [6 empty] | TBD |  |  | Operator technique control charts are posted | [10 empty] | kVp:
55 |  |  |  | Image receptor holder assembly is free from vibrations | [6 empty] | TBD |  |  | Operator protected during exposure by adequate radiation shielding | [10 empty] | Density:
56 |  |  |  | Image receptor slides smoothly into holder assembly | [6 empty] | NA |  |  | All indicator lights working properly | [10 empty] | Target/Filter:
57 |  |  |  | Image receptor is held securely by assembly in any orientation | [6 empty] | TBD |  |  | Auto decompression can be overridden to maintain compression (and status displayed) | [10 empty] | kVp:
58 |  |  |  | Compressed breast thickness scale accurate to ±5mm, reproducible to ±2 mm | [6 empty] | TBD |  |  | Manual emergency compression release can be activated in the event of a power failure | [10 empty] | Density:
59 |  |  |  | Patient or operator is not exposed to sharp or rough edges, or other hazards | [6 empty] | TBD | [13 empty] | 50/50 Breast
60 |  |  |  | Operator technique control charts are posted | [6 empty] | TBD | [6 empty] | MQSA Equipment Standards | [6 empty] | Target/Filter:
61 |  |  |  | Operator protected during exposure by adequate radiation shielding | [6 empty] | TBD |  |  | Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position | [10 empty] | kVp:
62 |  |  |  | All indicator lights working properly | [6 empty] | TBD |  |  | Tube-receptor assembly locking mechanism does not fail during power interruption | [10 empty] | Density:
63 |  |  |  | Autodecompression can be overridden to maintain compression | [6 empty] | TBD |  |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [10 empty] | Target/Filter:
64 |  |  |  | Manual emergency compression release in the event of a power failure | [6 empty] | TBD |  |  | Moving grids are present for all image receptor sizes | [10 empty] | kVp:
65 | [13 empty] | Magnification can be done without grid | [10 empty] | Density:
66 | [13 empty] | Beam limiting device allows useful beam to extend beyond chest wall edge of receptor | [10 empty] | Target/Filter:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | Average illumination of the light field is at least 160 lux | [10 empty] | kVp:
68 |  | Room Number | [8 empty] | Survey ID: |  |  | Magnification mode is available for non-interventional problem solving procedures | [10 empty] | Density:
1 | [11 empty] | Medical University of South Carolina |  | System provides magnification factor between 1.4 – 2.0 | [10 empty] | Target/Filter:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Prior indication of selected focal spot is provided | [10 empty] | kVp:
3 |  | Rule Number |  |  |  |  | MQSA Equipment Standards | [6 empty] | Prior indication of selected target is provided | [10 empty] | Density:
4 |  | 900.12(b)(3i) |  | Tube-receptor can be fixed into any operating position and does not move | [6 empty] | TBD |  |  | Post exposure display of target/focal spot is provided | [10 empty] | Dense Breast
5 |  | 900.12(b)(3ii) |  | Tube-receptor locking mechanism does not fail during power interruption | [6 empty] | TBD |  |  | Initial power driven compression can be activated hands-free from both sides | [10 empty] | Target/Filter:
6 |  | 900.12(b)(4i) |  | Image Receptor Sizes 18 x 24 & 24 x 30 are available | [6 empty] | TBD |  |  | Fine adjustment of compression can be activated from both sides | [10 empty] | kVp:
7 |  | 900.12(b)(4ii) |  | Moving grids are present for all image receptor sizessizes | [6 empty] | TBD |  |  | Compression paddles match the sizes of all image receptors | [10 empty] | Density:
8 |  | 900.12(b)(4iii) |  | Magnification can be done without grid | [6 empty] | TBD |  |  | Compression paddles do not flex by more than 1 cm | [10 empty] | Target/Filter:
9 |  | 900.12(b)(5i) |  | Beam limiting device allows useful beam to extend beyond chest wall edge | [6 empty] | TBD |  |  | Compression paddles are flat and parallel to the breast support table | [10 empty] | kVp:
10 |  | 900.12(b)(5ii) |  | Average illumination of the light field is at least 160 lux | [6 empty] | TBD |  |  | Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor | [10 empty] | Density:
11 |  | 900.12(b)(6i) |  | Magnification mode available for non-interventional problem solving | [6 empty] | TBD |  |  | Chest wall edge may be bent upwards for patient comfort and does not appear in the image | [10 empty] | Target/Filter:
12 |  | 900.12(b)(6ii) |  | System provides magnification factor between 1.4 – 2.0 | [6 empty] | TBD |  |  | X-ray field extends beyond the chest wall edge of the receptor | [10 empty] | kVp:
13 |  | 900.12(b)(7i) |  | Prior indication of selected focal spot is provided | [6 empty] | TBD |  |  | Manual selection of mAs or mA and/or time is available | [10 empty] | Density:
14 |  | 900.12(b)(7i) |  | Prior indication of selected target is provided | [6 empty] | TBD |  |  | Prior indication of selected technique factors is provided | [10 empty] | Target/Filter:
15 |  | 900.12(b)(7iii) |  | Post exposure display of target/focal spot is provided | [6 empty] | TBD |  |  | Post exposure display of technique factors is provided for AEC exposures | [10 empty] | kVp:
16 |  | 900.12(b)(8)(iB) |  | Initial power driven compression can be activated hands-free from both sides | [6 empty] | TBD |  |  | AEC mode is operable in all combinations of equipment configuration | [10 empty] | Density:
17 |  |  |  | Fine adjustment of compression can be activated from both sides | [6 empty] | TBD |  |  | AEC sensor positioning is flexible
18 |  | 900.12(b)(8)(iiA) |  | Compression paddles match the sizes of all image receptors | [6 empty] | TBD |  |  | AEC sensor position and size is indicated on the compression paddle | [10 empty] | 2D MGD:
19 |  | 900.12(b)(8)(iiB) |  | Compression paddles do not flex by more than 1 cm | [6 empty] | TBD |  |  | Selected position of the AEC sensor is clearly indicated | [10 empty] | mGy/mAs:
20 |  | 900.12(b)(8)(iiC) |  | Compression paddles are flat and parallel to the breast support table | [6 empty] | TBD |  |  | AEC density control setting is provided | [10 empty] | mGy/s:
21 |  | 900.12(b)(8)(iiD) |  | Chest wall edge of the compression paddle is straight and parallel to receptor | [6 empty] | TBD |  | 3 | X-ray film used is appropriate for mammography | [10 empty] | 3D MGD:
22 |  | 900.12(b)(8)(iiE) |  | Chest wall edge may be bent upwards and does not appear in the image | [6 empty] | TBD |  | 3 | Intensifying screens used are appropriate for mammography and matched to the film used | [10 empty] | Combo 2D MGD:
23 |  |  |  | X-ray field extends beyond the chest wall edge of the receptor | [6 empty] | TBD |  | 3 | Film processing chemicals are appropriate for the film being used | [10 empty] | Combo 3D MGD:
24 |  | 900.12(b)(9i) |  | Manual selection of mAs or mA and/or time is available | [6 empty] | TBD |  |  | Lights for film illumination i.e. hot lights are available | [10 empty] | Combo 2D+3D MGD:
25 |  | 900.12(b)(9ii) |  | Prior indication of selected technique factors is provided | [6 empty] | TBD |  |  | Film masking devices are available to all interpreting physicians | [10 empty] | SBB Stereo MGD:
26 |  | 900.12(b)(9iii) |  | Post exposure display of technique factors is provided for AEC exposures | [6 empty] | TBD | [13 empty] | SBB Tomo MGD:
27 |  | 900.12(b)(10i) |  | AEC mode is operable in all combinations of equipment configuration | [6 empty] | TBD | [6 empty] | Mammography Phototimer Technique Chart | [6 empty] | System Resolution
28 |  | 900.12(b)(10ii) |  | AEC sensor positioning is flexible | [6 empty] | TBD | [13 empty] | 0
29 |  |  |  | AEC sensor position and size is indicated on the compression paddle | [6 empty] | TBD |  | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | 0
30 |  |  |  | Selected position of the AEC sensor is clearly indicated | [6 empty] | TBD |  | Breast | [11 empty] | 0
31 |  | 900.12(b)(10iii) |  | AEC density control setting is provided | [6 empty] | TBD |  | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs
32 |  | 900.12(b)(11) |  | X-ray film used is appropriate for mammography | [6 empty] | NA |  | <3 cm | [11 empty] | Image Quality (Printer)
33 |  | 900.12(b)(12) |  | Intensifying screens used are appropriate for mammography and matched to film | [6 empty] | NA |  | 3 to 5 cm | [11 empty] | kV: |  |  | 0
34 |  | 900.12(b)(13) |  | Film processing chemicals are appropriate for the film being used | [6 empty] | NA |  | 5 to 7 cm | [11 empty] | mA: |  |  | 0
35 |  | 900.12(b)(14) |  | Lights for film illumination i.e. hot lights are available | [6 empty] | TBD |  | >7 cm | [11 empty] | Bkg Density:
36 |  | 900.12(b)(15) |  | Film masking devices are available to all interpreting physicians | [6 empty] | TBD | [13 empty] | Inside disk:
37 | [6 empty] | Comments | [5 empty] | Compressed | Fatty Breast |  |  | 50% Fatty - 50% Dense Breast |  |  | Dense Breast |  |  |  |  | Outside disk:
38 | [12 empty] | Breast | [11 empty] | Difference:
39 | [12 empty] | Thickness | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs | Target/Filter | kVp | mAs |  |  | Fibers:
40 | [12 empty] | <3 cm | [11 empty] | Specks:
41 | [12 empty] | 3 to 5 cm | [11 empty] | Masses:
42 | [12 empty] | 5 to 7 cm | [11 empty] | Image Quality (Acq Unit)
43 | [12 empty] | >7 cm | [11 empty] | 2D kV: |  |  | 0
44 | [24 empty] | mA: |  |  | 0
45 | [17 empty] | Inspection Results | [6 empty] | Exp Ind: |  |  | 0
46 | [12 empty] | Compression Force Indicator | [11 empty] | Fibers:
47 | [14 empty] | Indicated | Measured | After 20s | Slip | [6 empty] | Specks:
48 | [13 empty] | Power | [10 empty] | Masses:
49 | [13 empty] | Manual | [10 empty] | Mag kV:
50 | [13 empty] | Criteria: | Minimum compression force – 25 lbs | [9 empty] | mA:
51 | [14 empty] | Maximum compression force – 45 lbs | [9 empty] | Exp Ind:
52 | [24 empty] | Fibers:
53 | [12 empty] | Compression Thickness Indicator | [11 empty] | Specks:
54 | [13 empty] | BR-12 (cm) | Indicated | Difference | [8 empty] | Masses:
55 | [13 empty] | 1 | [10 empty] | 3D kV:
56 | [13 empty] | 2 | [10 empty] | mA:
57 | [13 empty] | 4 |  |  | Mean: | [7 empty] | Exp Ind:
58 | [13 empty] | 4 |  |  | StDev: | [7 empty] | Fibers:
59 | [13 empty] | 4 | [10 empty] | Specks:
60 | [13 empty] | 4 |  |  |  | Criteria: | Difference between indicated and actual thickness is ± 5mm | [5 empty] | Masses:
61 | [13 empty] | 6 | [10 empty] | SBB kV:
62 | [13 empty] | 8 | [10 empty] | mA:
63 | [24 empty] | Exp Ind:
64 | [13 empty] | Paddle flex is less than 1 cm |  |  |  | Compression thickness indicator: | [6 empty] | Fibers:
65 | [13 empty] | Overall unit assembly | [10 empty] | Specks:
66 | [24 empty] | Masses:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah | Artifact Evaluation |  | Acceptable: Enter 1 for YES, 2 for NO, 3 for NA | [9 empty] | SBB Tomo kV:
68 |  | Room Number | [8 empty] | Survey ID: |  |  | Target: | [10 empty] | mA:
1 | [11 empty] | Medical University of South Carolina |  | Filter: | [10 empty] | Exp Ind:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Focal spot: | [10 empty] | Fibers:
3 |  | Compression Force Indicator | [11 empty] | kV: | [10 empty] | Specks:
4 |  |  |  | Indicated | Measured | After 20s | Slip | [6 empty] | mAs: | [10 empty] | Masses:
5 |  |  | Power | [10 empty] | EI:
6 |  |  | Manual | [5 empty] | Acceptable: |  |  |  |  | Acceptable: | [10 empty] | SNR: |  | #DIV/0! | #DIV/0!
7 |  |  | Criteria: | Minimum compression force – 25 lbs | [20 empty] | CNR: |  | #DIV/0! | #DIV/0!
8 |  |  |  | Maximum compression force – 45 lbs | [8 empty] | Detector Uniformity
9 | [13 empty] | Target/Filter: | / | / | / | / | / | / |  |  |  |  | Comment Page 1
10 |  | Compression Thickness Indicator | [11 empty] | kV: | [16 empty] | Detector Uniformity deviation
11 |  |  | BR-12 (cm) | Indicated | Difference |  |  |  | Paddle flex is less than 1 cm: |  |  |  |  | mAs: | [16 empty] | / | / | / | / | / | /
12 |  |  | 1 | [5 empty] | Overall unit assembly: |  |  |  |  | ROI 1: | [15 empty] | ROI 1:
13 |  |  | 2 | [5 empty] | Compression thickness indicator: |  |  |  |  | ROI 2: | [15 empty] | ROI 2:
14 |  |  | 4 |  |  | Mean: | [7 empty] | ROI 3: | [15 empty] | ROI 3:
15 |  |  | 4 |  |  | StDev: | [7 empty] | ROI 4: | [15 empty] | ROI 4:
16 |  |  | 4 |  |  | Criteria: | Difference between indicated and actual thickness is ± 5mm | [6 empty] | ROI 5: | [15 empty] | ROI 5:
17 |  |  | 4 | [10 empty] | Average:
18 |  |  | 6 | [10 empty] | Max deviation:
19 |  |  | 8 | [10 empty] | Acceptable:
20
21 |  | Artifact Evaluation | [11 empty] | Criteria: | Maximum deviation from average is <= 7%
22 |  |  | Receptor:
23 |  |  | Target:
24 |  |  | Filter:
25 |  |  | Focal spot:
26 |  |  | kV: | [10 empty] | Comments:
27 |  |  | mAs: | [10 empty] | Do not exceed 390 chars
28 |  |  | EI: | [10 empty] | New:
29 |  |  | Artifact Evaluation Acceptable:
30 |  |  | Detector Uniformity Deviation: | [9 empty] | Light Field Brightness |  |  |  | Enter 1 for YES, 2 for NO, 3 for NA
31 |  |  | Detector Uniformity Acceptable: | [10 empty] | Measurements (lux) |  |  |  |  | Average Illumination
32 |  |  | Criteria: | No visible artifacts in the images | [9 empty] | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux)
33 |  |  |  | Maximum detector uniformity deviation is <= 7% | [17 empty] | Acceptable:
34
35 |  | Light Field Brightness | [10 empty] | Light Field/X-ray Field Alignment
36 |  |  | Measurements (lux) |  |  |  |  | Average Illumination |  |  |  |  | SID (cm): | 65
37 |  |  | 1st quad. | 2nd quad. | 3rd quad. | 4th quad. |  | (lux) | [5 empty] | 24x30 | 18x24 |  |  | 24x29 | Small breast
38 | [10 empty] | Acceptable: |  |  | 24x30 | Left | Center | Right | Tomo | 24x30
39 |  |  | Criteria: | Average light field illumination >= 160 lux. | [8 empty] | Left:
40 | [12 empty] | Right:
41 |  | Light Field/X-ray Field Alignment |  |  |  | X-ray Field/Image Receptor Alignment | [6 empty] | Anterior:
42 |  | SID (cm): | [10 empty] | Chest Wall:
43 |  | 65 | Light Field/X-ray Field % Deviation | [11 empty] | Enter deviation in cm
44 |  |  | 24x30 | 18x24 |  |  | 24x29 | Small breast
45 |  |  | 24x30 | Left | Center | Right | Tomo | Tomo | [5 empty] | % Deviation
46 |  | Left: | [11 empty] | 24x30 | 18x24 |  |  | 24x29 | Small breast
47 |  | Right: | [11 empty] | 24x30 | Left | Center | Right | Tomo | 24x30
48 |  | Anterior: | [10 empty] | Left:
49 |  | Chest Wall: | [10 empty] | Right:
50 |  | Left-Right: | [10 empty] | Anterior:
51 |  | Ant-Chest: | [10 empty] | Chest Wall:
52 |  | Acceptable: | [10 empty] | Left-Right
53 |  |  | X-ray Field/Image Receptor % Deviation | [9 empty] | Ant-Chest
54 |  | Left: | [11 empty] | Criteria: | Sum of left and right edge deviation is less than 2% SID
55 |  | Right: | [12 empty] | Sum of anterior and chest wall edge deviation is less than 2% SID
56 |  | Anterior: | [12 empty] | Enter deviation in cm. Positive deviation indicates x-ray field extends beyond light field
57 |  | Chest Wall: | [10 empty] | X-ray Field/Image Receptor Alignment
58 |  | Paddle:
59 |  | Acceptable: | [11 empty] | 24x30 | 18x24 |  |  | 24x29 | Small breast
60 |  | Criteria: | Sum of left/right and chest wall/anterior deviation is less than 2% SID | [10 empty] | 24x30 | Left | Center | Right | Tomo | 24x30
61 |  |  | X-ray field does not extend beyond image receptor by more than 2% SID | [9 empty] | Left:
62 |  |  | X-ray field extends beyond the chest wall edge of image receptor. | [9 empty] | Right:
63 |  |  | X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side | [9 empty] | Anterior:
64 |  |  | X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side | [9 empty] | Chest Wall:
65 |  |  | Difference between chest wall edge of the receptor and compression paddle is less than 1% SID | [9 empty] | Paddle:
66 | [14 empty] | Enter deviation in cm. Positive deviation indicates x-ray field extends beyond image receptor
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah
68 |  | Room Number | [8 empty] | Survey ID: |  |  | % Deviation
1 | [11 empty] | Medical University of South Carolina |  | 24x30 | 18x24 |  |  | 24x29 | Small breast
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | 24x30 | Left | Center | Right | Tomo | 24x30
3 |  | AEC Thickness Tracking – 2D | [5 empty] | CNR Table: | 4 |  |  |  | Left:
4 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | Mag CNR Table: | 4 |  |  |  | Right:
5 |  |  | Thickness |  |  |  |  | Reported |  |  |  |  | Anterior:
6 |  |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  | Chest Wall:
7 |  |  | 2 | [9 empty] | Paddle:
8 |  |  | 4 | [10 empty] | X-ray field extends beyond the chest wall edge of the receptor
9 |  |  | 4 | 0 | 0 | [8 empty] | Chest wall edge of the compression paddle is not visible in the image
10 |  |  | 4 | 0 | 0
11 |  |  | 4 | 0 | 0 | [8 empty] | Criteria: | X-ray field does not extend beyond image receptor by more than 2% SID
12 |  |  | 6 | [11 empty] | X-ray field extends beyond the chest wall edge of image receptor.
13 |  |  | 8 | [11 empty] | X-ray field does not fall within the image receptor by more than 2% SID on any side
14 |  |  | 4 | [11 empty] | Difference between chest wall edge of the receptor and compression paddle is less than 1% SID
15 | [5 empty] | Average:
16 |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
17 | [12 empty] | System Resolution |  | Large | Small | Large | Small
18 |  | AEC Thickness Tracking – 3D | [11 empty] | Mode (2D/3D):
19 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0 |  |  |  |  | Target:
20 |  |  | Thickness |  |  |  |  | Reported | [5 empty] | kV:
21 |  |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | mAs:
22 |  |  | 2 | [10 empty] | lp/mm:
23 |  |  | 4 | [10 empty] | Acceptable:
24 |  |  | 4 | 0 | 0 | [8 empty] | Previous:
25 |  |  | 4 | 0 | 0
26 |  |  | 4 | 0 | 0 | [8 empty] | Criteria: | Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode
27 |  |  | 6 | [11 empty] | Limiting system resolution must be 3 lp/mm or higher in 3D mode
28 |  |  | 8 | [9 empty] | AEC Thickness Tracking – 2D | [5 empty] | CNR Table: | 4 | Enter -1 if not applicable
29 | [5 empty] | Average: | [6 empty] | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | Mag CNR Table: | 4 | Enter -1 if not applicable
30 |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [9 empty] | Thickness | [5 empty] | Reported
31 | [13 empty] | (cm) | Filter | kV | mAs | EI | EI* | MGD (mGy) | Deviation
32 |  | AEC Density Control | [11 empty] | 2
33 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | kV: | [5 empty] | 4
34 | [5 empty] | Reported | [7 empty] | 4 | 0 | 0
35 |  |  | Density | mAs | EI | MGD (mGy) | Change | Acceptable | [5 empty] | 4 | 0 | 0
36 |  |  | -3 | [10 empty] | 4 | 0 | 0
37 |  |  | -2 | [10 empty] | 6
38 |  |  | -1 | [10 empty] | 8
39 |  |  | 0 | [9 empty] | Mag: | 4
40 |  |  | 1 | [11 empty] | EI* = EI - DC offset (50) |  |  | Average:
41 |  |  | 2 | [10 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
42 |  |  | 3
43 |  |  | 4 | [9 empty] | AEC Thickness Tracking – 3D
44 |  |  | Criteria: | Change in pixel value at each density step is within specified ranges | [8 empty] | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0
45 | [13 empty] | Thickness | [5 empty] | Reported
46 |  | System Resolution |  | Large | Small | Large | Small | [6 empty] | (cm) | Filter | kV | mAs | EI | EI* | MGD (mGy) | Deviation
47 |  |  | Mode (2D/3D): | [10 empty] | 2
48 |  |  | Target: | [10 empty] | 4
49 |  |  | kV: | [5 empty] | Acceptable: |  |  |  |  | 4 | 0 | 0
50 |  |  | mAs: | [10 empty] | 4 | 0 | 0
51 |  |  | lp/mm: | [10 empty] | 4 | 0 | 0
52 |  |  | Criteria: | Limiting system resolution must be 7 lp/mm or higher in conventional/2D mode | [9 empty] | 6
53 | [13 empty] | 8
54 |  | Image Quality Evaluation | [12 empty] | EI* = EI - DC offset (50) |  |  | Average:
55 |  |  | Printer |  |  | Acq Workstation |  |  |  | Review Workstation |  |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
56 |  |  | kV: | 0 |  | Normal | Mag | Tomo |  | Left | Right
57 |  |  | mA: | 0 | kV: | 0 |  |  |  | Workstation:
58 |  |  | Background Density: |  | mA: | 0 |  |  | Fibers: |  |  |  | AEC Density Control
59 |  |  | Inside disk: |  | EI: | 0 |  |  | Specks: |  |  |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 | Reported |  |  |  | Selenia |  | Selenia |  | Selenia Dimensions
60 |  |  | Outside disk: |  | Fibers: |  |  |  | Masses: |  |  |  |  | Density | kV | mAs | EI | MGD (mGy) | EI* | Change |  | Low | High | Low | High | Low | High
61 |  |  | Difference: |  | Specks: |  |  |  |  | Workstation: |  |  |  | -3 | [7 empty] | 0.5 | 0.61 | 0.5 | 0.61 | 0.56 | 0.66
62 |  |  | Fibers: |  | Masses: |  |  |  | Fibers: |  |  |  |  | -2 | [7 empty] | 0.63 | 0.77 | 0.63 | 0.77 | 0.66 | 0.78
63 |  |  | Specks: |  | Criteria: |  |  |  | Specks: |  |  |  |  | -1 | [7 empty] | 0.77 | 0.94 | 0.77 | 0.94 | 0.78 | 0.92
64 |  |  | Masses: |  | 2D – Must be able to see 5 fibers, 4 speck groups, 4 masses |  |  |  | Masses: |  |  |  |  | 0
65 |  |  |  |  | 3D – Must be able to see 4 fibers, 3 speck groups, 3 masses | [8 empty] | 1 | [7 empty] | 1.04 | 1.27 | 1.04 | 1.27 | 1.06 | 1.24
66 | [13 empty] | 2 | [7 empty] | 1.17 | 1.43 | 1.17 | 1.43 | 1.22 | 1.43
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | 3 | [7 empty] | 1.31 | 1.6 | 1.31 | 1.6 | 1.4 | 1.64
68 |  | Room Number | [8 empty] | Survey ID: |  |  | 4 | [7 empty] | 1.44 | 1.76 | 1.44 | 1.76 | 1.61 | 1.89
1 | [11 empty] | Medical University of South Carolina |  |  | EI* = EI - DC offset (50)
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  | Criteria: | Change in pixel value at each density step is within specified ranges
3 | Stereotactic Breast Biopsy
4
5 |  | AEC Thickness Tracking - Stereotactic Breast Biopsy Add-on | [10 empty] | AEC Thickness Tracking - Stereotactic Breast Biopsy Add-on
6 |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0 |  |  |  | AEC Mode: | Auto-filter |  | AEC Position: | 2 |  | CNR Table: | 0
7 |  |  | Thickness |  |  |  |  | Reported | [5 empty] | Thickness | [5 empty] | Reported
8 |  |  | (cm) | Filter | kV | mAs | EI | MGD (mGy) | Deviation |  |  |  |  | (cm) | Filter | kV | mAs | EI | EI* | MGD (mGy) | Deviation
9 |  |  | 2 | [10 empty] | 2
10 |  |  | 4 | [10 empty] | 4
11 |  |  | 4 | 0 | 0 | [8 empty] | 4 | 0 | 0
12 |  |  | 4 | 0 | 0 | [8 empty] | 4 | 0 | 0
13 |  |  | 4 | 0 | 0 | [8 empty] | 4 | 0 | 0
14 |  |  | 6 | [10 empty] | 6
15 |  |  | 8 | [10 empty] | 8
16 | [5 empty] | Average: | [8 empty] | EI* = EI - DC offset (50) |  |  | Average:
17 |  |  | Criteria: | Mean pixel value of each image is within 10% of the average pixel value | [9 empty] | Criteria: | Mean pixel value of each image is within 10% of the average pixel value
18 | [12 empty] | Mean Glandular Dose – 2D
19 |  | Mean Glandular Dose – Stereo biopsy | [11 empty] | Phantom: |  |  | kV:
20 |  |  | Phantom: | 0 |  | kV: | [7 empty] | Target/Filter: |  |  | mA:
21 |  |  | Target/Filter: |  |  | mA: | [10 empty] | Reported |  | Air KERMA
22 | [5 empty] | Reported |  | Air KERMA | [6 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  | Bkg Mean | SD | Disk Mean | SD | SNR | CNR
23 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: | [10 empty] | ESE (mGy): | [7 empty] | #DIV/0! | #DIV/0!
24 | [9 empty] | ESE (mGy): | [10 empty] | DGN (mrad/R): | [7 empty] | #DIV/0! | #DIV/0!
25 | [9 empty] | DGN (mrad/R): | [10 empty] | MGD (mGy): | [7 empty] | #DIV/0! | #DIV/0!
26 | [9 empty] | MGD (mGy): | [10 empty] | Prev MGD (mGy): | [7 empty] | #DIV/0! | #DIV/0!
27 | [9 empty] | Prev MGD (mGy): |  |  |  | Average: | [6 empty] | % Change:
28 |  |  | Average: | [6 empty] | % Change: |  |  |  | Coeff of Var:
29 |  |  | Coeff of Var:
30 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
31 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: |  |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
32
33 |  | Mean Glandular Dose – Tomo biopsy | [10 empty] | Mean Glandular Dose – 3D
34 |  |  | Phantom: | 0 |  | kV: | [7 empty] | Phantom: | 0 |  | kV:
35 |  |  | Target/Filter: |  |  | mA: | [7 empty] | Target/Filter: |  |  | mA:
36 | [5 empty] | Reported |  | Air KERMA | [8 empty] | Reported |  | Air KERMA
37 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
38 | [9 empty] | ESE (mGy): | [10 empty] | ESE (mGy):
39 | [9 empty] | DGN (mrad/R): | [10 empty] | DGN (mrad/R):
40 | [9 empty] | MGD (mGy): | [10 empty] | MGD (mGy):
41 | [9 empty] | Prev MGD (mGy): | [10 empty] | Prev MGD (mGy):
42 |  |  | Average: | [6 empty] | % Change: |  |  |  | Average: | [6 empty] | % Change:
43 |  |  | Coeff of Var: | [10 empty] | Coeff of Var:
44 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
45 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
46 | [14 empty] | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
47 |  | Phantom Image Quality | [12 empty] | MGD for 2D + 3D acquisitions is less than 3 mGy
48 |  |  | Acq Workstation
49 |  |  | SBB | SBB Tomo
50 |  | kV: | [10 empty] | Mean Glandular Dose – Combo
51 |  | mA: | [11 empty] | Phantom: | 0
52 |  | EI:
53 |  | Fibers: | [10 empty] | Combo Mode 2D
54 |  | Specks: | [11 empty] | Target/Filter: |  |  | kV:
55 |  | Masses: | [14 empty] | mA:
56 |  | Criteria: | SBB Stereo - Must be able to see 5 fibers, 4 speck groups, 3.5 masses | [13 empty] | Reported |  | Air KERMA
57 |  |  | SBB Tomo - Must be able to see 3 fibers, 3 speck groups, 2.5 masses | [11 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
58 | [20 empty] | ESE (mGy):
59 | [20 empty] | DGN (mrad/R):
60 | [20 empty] | MGD (mGy):
61 | [20 empty] | Prev MGD (mGy):
62 | [13 empty] | Average: | [6 empty] | % Change:
63 | [13 empty] | Coeff of Var:
64 | [20 empty] | Deviation from Reported MGD:
65 | [12 empty] | Combo Mode 3D
66 | [13 empty] | Target/Filter: |  |  | kV:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  |  |  |  | mA:
68 |  | Room Number | [8 empty] | Survey ID: | [5 empty] | Reported |  | Air KERMA
1 | [11 empty] | Medical University of South Carolina |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection | [8 empty] | ESE (mGy):
3 |  | Mean Glandular Dose – 2D | [18 empty] | DGN (mrad/R):
4 |  |  | Phantom: |  |  | kV: | [14 empty] | MGD (mGy):
5 |  |  | Target/Filter: |  |  | mA: | [14 empty] | Prev MGD (mGy):
6 | [5 empty] | Reported |  | Air KERMA | [5 empty] | Average: | [6 empty] | % Change:
7 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  | Coeff of Var:
8 | [9 empty] | ESE (mGy): | [10 empty] | Deviation from Reported MGD:
9 | [9 empty] | DGN (mrad/R): | [10 empty] | Combo MGD (2D + 3D):
10 | [9 empty] | MGD (mGy): | [10 empty] | Prev Combo MGD:
11 | [9 empty] | Prev MGD (mGy): |  |  | Mean Glandular Dose – Stereo biopsy
12 |  |  | Average: | [6 empty] | % Change: |  |  |  | Phantom: | 0 |  | kV:
13 |  |  | Coeff of Var: | [10 empty] | Target/Filter: |  |  | mA:
14 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | [6 empty] | Reported |  | Air KERMA
15 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: |  |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
16 | [20 empty] | ESE (mGy):
17 |  | Mean Glandular Dose – 3D | [18 empty] | DGN (mrad/R):
18 |  |  | Phantom: | 0 |  | kV: | [14 empty] | MGD (mGy):
19 |  |  | Target/Filter: |  |  | mA: | [14 empty] | Prev MGD (mGy):
20 | [5 empty] | Reported |  | Air KERMA | [5 empty] | Average: | [6 empty] | % Change:
21 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  | Coeff of Var:
22 | [9 empty] | ESE (mGy):
23 | [9 empty] | DGN (mrad/R): |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
24 | [9 empty] | MGD (mGy): |  |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
25 | [9 empty] | Prev MGD (mGy):
26 |  |  | Average: | [6 empty] | % Change: |  |  | Mean Glandular Dose – Tomo biopsy
27 |  |  | Coeff of Var: | [10 empty] | Phantom: | 0 |  | kV:
28 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: |  |  |  | Target/Filter: |  |  | mA:
29 |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy: | [6 empty] | Reported |  | Air KERMA
30 | [14 empty] | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL:
31 |  | Mean Glandular Dose – Combo | [18 empty] | ESE (mGy):
32 |  |  | Phantom: | 0 | [16 empty] | DGN (mrad/R):
33 |  |  |  | Combo Mode 2D Target/Filter: | [16 empty] | MGD (mGy):
34 | [5 empty] | Reported |  | Air KERMA |  | kV: | [10 empty] | Prev MGD (mGy):
35 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  | Average: | [6 empty] | % Change:
36 | [9 empty] | ESE (mGy): |  |  |  | Coeff of Var:
37 | [9 empty] | DGN (mrad/R):
38 | [9 empty] | MGD (mGy): |  |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD:
39 | [9 empty] | Prev MGD (mGy): |  |  |  |  | Deviation from reported MGD is less than15% | [5 empty] | mAs at 3 mGy:
40 |  |  | Average: | [6 empty] | % Change:
41 |  |  | Coeff of Var: | [9 empty] | kVp Accuracy and Output
42 | [9 empty] | Deviation from Reported MGD: |  |  | Meter: | Piranha |  |  | Calibration Date:
43 | [12 empty] | Serial #: |  |  |  | Calibration Due:
44 |  |  |  | Combo Mode 3D Target/Filter:
45 | [5 empty] | Reported |  | Air KERMA |  | kV: |  |  |  | Criteria: | kV accuracy is <= 5% |  |  | Measured Values
46 |  |  |  | mAs | EI | MGD (mGy) | Time (s) | (mGy) |  | HVL: |  |  |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
47 | [9 empty] | ESE (mGy): | [5 empty] | 24 | 50
48 | [9 empty] | DGN (mrad/R): | [5 empty] | 25 | 50
49 | [9 empty] | MGD (mGy): | [5 empty] | 26 | 50
50 | [9 empty] | Prev MGD (mGy): | [5 empty] | 28 | 50
51 |  |  | Average: | [6 empty] | % Change: | [5 empty] | 30 | 50
52 |  |  | Coeff of Var: | [12 empty] | 32 | 50
53 |  |  | Criteria: | MGD is less than 3 mGy | [5 empty] | Deviation from Reported MGD: | [5 empty] | 34 | 50
54 |  |  |  | Deviation from reported MGD is less than15% | [11 empty] | 36 | 50
55 | [9 empty] | Combo MGD: | [5 empty] | 38 | 50
56 | [9 empty] | Prev Combo MGD:
57 | [17 empty] | Measured Values
58 | [13 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
59 |  | Signal to Noise Ratio/Contrast to Noise Ratio | [13 empty] | 28 | 50
60 |  |  | kV: | 0 |  |  | Current | Previous | Change | Acceptable | [5 empty] | 30 | 50
61 |  |  | mAs: | 0 |  | SNR: | #DIV/0! |  | #DIV/0! | #DIV/0! | [5 empty] | 32 | 50
62 |  |  | Target: |  |  | CNR: | #DIV/0! |  | #DIV/0! | NA | [5 empty] | 34 | 50
63 |  |  | Filter: | [12 empty] | 36 | 50
64 |  |  | Criteria: | SNR must be >= 40 | [11 empty] | 38 | 50
65 |  |  |  | CNR does not change by more than 15%
66 | [17 empty] | Measured Values
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
68 |  | Room Number | [8 empty] | Survey ID: |  |  |  |  | 28 | 50
1 | [11 empty] | Medical University of South Carolina |  |  |  | 30 | 50
2 | [6 empty] | Inspection Results |  |  |  |  | Mammography System Compliance Inspection |  |  |  | 32 | 50
3 |  | Meter: | Piranha |  |  |  | Calibration Date: | [8 empty] | 34 | 50
4 |  | SN: |  |  |  |  | Calibration Due: | [8 empty] | 36 | 50
5 |  | kVp Accuracy | [13 empty] | 38 | 50
6 |  | Target: |  |  |  | Target: |  |  |  | Target: |  |  |  | Criteria: | kV accuracy is <= 5%
7 |  | Filter: |  |  |  | Filter: |  |  |  | Filter: |  |  | kV, Exposure, and Output Reproducibility
8 |  | mAs: | 50 |  |  | mAs: | 50 |  |  | mAs: | 50
9 |  | Indicated | Measured |  |  | Indicated | Measured |  |  | Indicated | Measured | [6 empty] | Measured Values
10 |  | kVp | kVp | Error (%) |  | kVp | kVp | Error (%) |  | kVp | kVp | Error (%) |  | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
11 |  | 24 |  |  |  | 28 |  |  |  | 28 | [5 empty] | 28 | 50
12 |  | 25 |  |  |  | 30 |  |  |  | 30 | [5 empty] | 28 | 50
13 |  | 26 |  |  |  | 32 |  |  |  | 32 | [5 empty] | 28 | 50
14 |  | 28 |  |  |  | 34 |  |  |  | 34 | [5 empty] | 28 | 50
15 |  | 30 |  |  |  | 36 |  |  |  | 36 | [6 empty] | Mean:
16 |  | 32 |  |  |  | 38 |  |  |  | 38 | [6 empty] | Std Dev:
17 |  | 34 | [14 empty] | Coeff of Var:
18 |  | 36 |  |  |  |  | Acceptable: |  |  |  | Acceptable: | [5 empty] | Previous:
19 |  | 38 | [11 empty] | Criteria: | Coefficient of variation <= 2%
20 |  |  | Acceptable: |  | Criteria: | kVp accurate to within 5% of indicated | [8 empty] | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
21 | [14 empty] | Minimum acceptable output is 2 mGy/s at 28 kVp for W target
22 |  | kVp and Output Reproducibility | [10 empty] | Exposure Linearity
23 |  | Target: |  | Filter: | [13 empty] | Measured Values
24 |  | mAs: | 50 | [10 empty] | Target: | Filter | kV | mAs | kVp | ms | Dose (mGy) | mGy/mAs | mGy/s
25 |  | Indicated | Measured | [12 empty] | 28 | 20
26 |  | kVp | kVp | mGy | mGy/mAs | mGy/s | [9 empty] | 28 | 50
27 |  | 28 | [13 empty] | 28 | 100
28 | [15 empty] | 28 | 300
29 | [15 empty] | 28 | 400
30 | [13 empty] | Criteria: | Coefficient of linearity is <= 10% |  |  |  |  | Coeff of Linearity:
31 |  | Mean: | [11 empty] | Unit Can Maintain Output for 3 sec at 28 kVp
32 |  | Std Dev:
33 |  | Coeff of Var: | [10 empty] | Half Value Layer – 2D |  | 1 | Compression paddle used?
34 |  | Acceptable: | [11 empty] | Target/Filter | / | / | / | / | / | / | / | /
35 |  |  | Criteria: | Coefficient of variation <= 2% | [9 empty] | kV | 24 | 25 | 28 | 32 | 28 | 30 | 32 | 34
36 |  |  |  | Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target | [9 empty] | Al (mm) | Air Kerma (mGy)
37 |  |  |  | Minimum acceptable output is 2 mGy/s at 28 kVp for W target | [9 empty] | 0
38 | [13 empty] | 0
39 |  | Exposure Linearity | [11 empty] | HVL
40 |  | Target: |  | Filter: | [9 empty] | 0 mm
41 |  | Set mAs | kVp | mGy | mGy/mAs | mGy/s | [7 empty] | Min | 0.27 | 0.28 | 0.31 | 0.35 | 0.31 | 0.33 | 0.35 | 0.37
42 |  | 20 | [11 empty] | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.47 | 0.49 | 0.51 | 0.53
43 |  | 50
44 |  | 100 | [11 empty] | Criteria: | HVL is between the minimum and maximum limits
45 |  | 300
46 |  | 400 | [10 empty] | Half Value Layer – 3D
47 |  |  |  | Coeff of Linearity: | [9 empty] | Target/Filter | / | / | / | / | /
48 |  |  | Criteria: | Coefficient of linearity is <= 10% | [9 empty] | kV | 28 | 30 | 32 | 34 | 36
49 | [13 empty] | Al (mm) | Air Kerma (mGy)
50 |  | Half Value Layer | [11 empty] | 0
51 |  | Target/Filter | / | / | / | / | / | / | / | / |  |  |  | HVL
52 |  | kV | 24 | 25 | 28 | 32 | 28 | 30 | 32 | 34 |  |  |  | Min | 0.31 | 0.33 | 0.35 | 0.37 | 0.39
53 |  | HVL | [11 empty] | Max | 0.4 | 0.42 | 0.44 | 0.46 | 0.55
54 |  | 0 mm
55 |  | Min | 0.27 | 0.28 | 0.31 | 0.35 | 0.31 | 0.33 | 0.35 | 0.37 |  |  |  | Criteria: | Mo/Rh targets – HVL is between the minimum and maximum limits
56 |  | Max | 0.36 | 0.37 | 0.4 | 0.44 | 0.47 | 0.49 | 0.51 | 0.53 |  |  |  |  | W targets – HVL is above the minimum limit
57 |  | Acceptable:
58
59 |  | Target/Filter | / | / | / | / | / | [5 empty] | Image Quality Evaluation
60 |  | kV | 28 | 30 | 32 | 34 | 36 | [5 empty] | Printer | Model: |  | SN: |  |  | QC Manual Version:
61 |  | HVL
62 |  | Min | 0.31 | 0.33 | 0.35 | 0.37 | 0.39 | [8 empty] | Previous |  |  |  | Previous
63 |  | Max | 0.4 | 0.42 | 0.44 | 0.46 | 0.55 | [6 empty] | kV: | 0 |  |  | Fibers:
64 |  | Acceptable: | [11 empty] | mAs: | 0 |  |  | Specks:
65 |  |  | Criteria: | Mo/Rh targets – HVL is between the minimum and maximum limits | [9 empty] | Background Density: |  |  |  | Masses:
66 | [13 empty] | Inside disk: |  |  |  | Acceptable:
67 |  | Date: | [8 empty] | Inspector: | Eugene Mah |  | Outside disk:
68 |  | Room Number | [8 empty] | Survey ID: |  |  | Difference:
Workstation
[15 empty] | Workstation: |  |  |  | Workstation:
[15 empty] | Left | Right |  |  | Left | Right
[14 empty] | Fibers: |  |  |  | Fibers:
[14 empty] | Specks: |  |  |  | Specks:
[14 empty] | Masses: |  |  |  | Masses:
Acquisition Unit
[14 empty] | 2D | Previous | Mag | Previous | 3D | Previous | SBB | Previous | SBB Tomo | Previous
[13 empty] | kV: | 0
[13 empty] | mAs: | 0
[13 empty] | EI: | 0
Fibers:
Specks:
Masses:
[14 empty] | Criteria: | 2D – Must be able to see 5 fibers, 4 speck groups, 4 masses
3D – Must be able to see 4 fibers, 3 speck groups, 3 masses
SBB Stereo - Must be able to see 5 fibers, 4 speck groups, 3.5 masses
SBB Tomo - Must be able to see 3 fibers, 3 speck groups, 2.5 masses
Signal to Noise
[18 empty] | Mean | SD
[14 empty] | kV: | 0 |  | In acrylic: | #DIV/0! | #DIV/0!
[14 empty] | mAs: | 0 |  | Background: | #DIV/0! | #DIV/0!
Target:
[14 empty] | Filter: |  |  |  | Current | Previous | Change | Acceptable
[17 empty] | SNR: | #DIV/0! |  | #DIV/0! | #DIV/0!
[17 empty] | CNR: | #DIV/0! |  | #DIV/0! | NA
[14 empty] | Criteria: | SNR must be >= 40
CNR does not change by more than 15%
Previous Year Comments
Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0
Additional Comments:
[14 empty] | Do not exceed 390 chars |  | 0

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast | [9 empty] | Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
| kVp | [9 empty] | kVp | [9 empty] | kVp
HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | HVL | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
0.23 | 116 | [8 empty] | 0.23 | [9 empty] | 0.23
0.24 | 121 | 124 | [7 empty] | 0.24 | [9 empty] | 0.24
0.25 | 126 | 129 | 131 | [6 empty] | 0.25 | [9 empty] | 0.25
0.26 | 130 | 133 | 135 | 138 | [5 empty] | 0.26 | [9 empty] | 0.26
0.27 | 135 | 138 | 140 | 142 | 143 |  |  |  |  | 0.27 | [9 empty] | 0.27
0.28 | 140 | 142 | 144 | 146 | 147 | 149 |  |  |  | 0.28 |  |  | 149 | 151 | 154 |  |  |  |  | 0.28 |  |  | 150 | 155 | 159
0.29 | 144 | 146 | 148 | 150 | 151 | 153 | 154 |  |  | 0.29 |  |  | 154 | 156 | 158 | 159 |  |  |  | 0.29 |  |  | 155 | 160 | 164 | 168
0.3 | 149 | 151 | 153 | 155 | 156 | 157 | 158 | 159 |  | 0.3 |  |  | 158 | 160 | 162 | 162 | 163 |  |  | 0.3 |  |  | 160 | 164 | 168 | 172 | 176
0.31 | 154 | 156 | 157 | 159 | 160 | 161 | 162 | 163 | 164 | 0.31 |  |  | 163 | 164 | 166 | 166 | 167 | 167 |  | 0.31 |  |  | 165 | 168 | 172 | 174 | 180 | 182
0.32 | 158 | 160 | 162 | 163 | 164 | 166 | 167 | 168 | 168 | 0.32 |  |  | 167 | 169 | 171 | 171 | 171 | 172 | 172 | 0.32 |  |  | 169 | 173 | 177 | 181 | 184 | 186 | 188
0.33 | 163 | 165 | 166 | 168 | 169 | 170 | 171 | 173 | 173 | 0.33 |  |  | 171 | 173 | 175 | 176 | 176 | 176 | 176 | 0.33 |  |  | 174 | 178 | 181 | 185 | 188 | 190 | 192
0.34 | 168 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 0.34 |  |  | 176 | 178 | 179 | 179 | 180 | 180 | 180 | 0.34 |  |  | 179 | 183 | 186 | 190 | 193 | 195 | 196
0.35 |  | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 0.35 |  |  | 180 | 181 | 183 | 183 | 184 | 185 | 185 | 0.35 |  |  | 184 | 187 | 190 | 194 | 197 | 199 | 201
0.36 |  |  | 179 | 181 | 182 | 183 | 184 | 185 | 185 | 0.36 |  |  | 185 | 186 | 187 | 187 | 188 | 188 | 189 | 0.36 |  |  | 189 | 192 | 195 | 198 | 201 | 204 | 205
0.37 |  |  |  | 185 | 186 | 187 | 188 | 189 | 190 | 0.37 |  |  | 189 | 190 | 191 | 191 | 192 | 193 | 193 | 0.37 |  |  | 193 | 196 | 199 | 202 | 205 | 207 | 209
0.38 |  |  |  |  | 190 | 191 | 192 | 193 | 194 | 0.38 |  |  | 193 | 194 | 196 | 196 | 197 | 197 | 197 | 0.38 |  |  | 198 | 201 | 204 | 207 | 209 | 211 | 213
0.39 |  | 174 178: 196; 175 179: 197; 176 180: 198; 177 181: 198; 0.34 0.35: 0.39 |  |  | 198 | 199 | 200 | 200 | 201 | 201 | 202 | 0.39 |  |  | 203 | 206 | 208 | 211 | 214 | 216 | 217
0.4 |  | 175 179: 201; 176 180: 202; 177 181: 203; 0.34 0.35: 0.4 |  |  | 202 | 203 | 204 | 204 | 205 | 205 | 206 | 0.4 |  |  | 208 | 211 | 213 | 216 | 218 | 220 | 221
0.41 |  | 176 180: 206; 177 181: 207; 0.34 0.35: 0.41 |  |  | 206 | 207 | 208 | 208 | 209 | 209 | 210 | 0.41 |  |  | 213 | 215 | 217 | 220 | 222 | 224 | 225
0.42 |  | 177 181: 211; 0.34 0.35: 0.42 |  |  | 211 | 211 | 212 | 212 | 213 | 213 | 214 | 0.42 |  |  | 218 | 220 | 222 | 224 | 226 | 228 | 229
Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast | [12 empty] | Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
| kVp | [12 empty] | kVp
HVL | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | HVL | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39
0.3 | 152 | 157 | 163 | 166 | 170 | 173 | 175 | 177 | 179 | 182 | 184 | 187 | 0.4 | 222 | 226 | 229 | 231 | 234 | 236 | 239 | 241 | 244 | 246 | 248 | 250 | 252
0.325 | 163 | 169 | 174 | 177 | 181 | 183 | 186 | 188 | 190 | 192 | 195 | 197 | 0.425 | 233 | 236 | 239 | 242 | 244 | 246 | 248 | 251 | 253 | 256 | 258 | 260 | 262
0.35 | 175 | 180 | 185 | 188 | 191 | 194 | 196 | 198 | 200 | 202 | 205 | 207 | 0.45 | 244 | 247 | 249 | 252 | 254 | 256 | 258 | 260 | 263 | 265 | 267 | 269 | 271
0.375 | 186 | 191 | 196 | 199 | 202 | 205 | 207 | 209 | 211 | 213 | 215 | 218 | 0.475 | 254 | 257 | 260 | 262 | 264 | 266 | 268 | 270 | 273 | 275 | 277 | 279 | 281
0.4 | 198 | 203 | 207 | 210 | 213 | 215 | 217 | 219 | 221 | 223 | 226 | 228 | 0.5 | 265 | 267 | 270 | 272 | 274 | 276 | 278 | 280 | 282 | 284 | 286 | 288 | 290
0.425 | 209 | 214 | 218 | 221 | 224 | 226 | 228 | 230 | 232 | 234 | 236 | 238 | 0.525 | 275 | 278 | 280 | 282 | 284 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 300
0.45 | 221 | 226 | 230 | 232 | 235 | 237 | 238 | 240 | 242 | 244 | 246 | 248 | 0.55 | 286 | 288 | 290 | 292 | 294 | 296 | 298 | 299 | 301 | 303 | 305 | 307 | 309
0.475 | 233 | 237 | 241 | 243 | 245 | 247 | 249 | 251 | 253 | 254 | 256 | 258 | 0.575 | 296 | 298 | 300 | 302 | 304 | 305 | 307 | 309 | 311 | 313 | 315 | 317 | 318
0.5 | 244 | 248 | 252 | 254 | 256 | 258 | 260 | 261 | 263 | 265 | 267 | 269 | 0.6 | 306 | 308 | 310 | 312 | 313 | 315 | 317 | 319 | 320 | 322 | 324 | 326 | 328
0.525 | 256 | 260 | 263 | 265 | 267 | 269 | 270 | 272 | 273 | 275 | 277 | 279 | 0.625 | 316 | 318 | 320 | 322 | 323 | 325 | 326 | 328 | 330 | 332 | 333 | 335 | 337
0.55 | 267 | 271 | 274 | 276 | 278 | 279 | 281 | 282 | 284 | 285 | 287 | 289 | 0.65 | 326 | 328 | 330 | 331 | 333 | 334 | 336 | 338 | 339 | 341 | 343 | 344 | 346
0.575 | 279 | 282 | 285 | 287 | 288 | 290 | 291 | 292 | 294 | 296 | 297 | 299 | 0.675 | 336 | 338 | 339 | 341 | 342 | 344 | 345 | 347 | 349 | 350 | 352 | 354 | 355
0.6 | 290 | 293 | 296 | 297 | 299 | 300 | 301 | 303 | 304 | 306 | 308 | 310 | 0.7 | 346 | 348 | 349 | 350 | 352 | 353 | 355 | 356 | 358 | 359 | 361 | 363 | 364
0.625 | 301 | 304 | 306 | 308 | 309 | 310 | 312 | 313 | 315 | 316 | 318 | 320 | 0.725 | 356 | 357 | 358 | 360 | 361 | 362 | 364 | 365 | 367 | 368 | 370 | 372 | 373
0.65 | 312 | 314 | 317 | 318 | 320 | 321 | 322 | 323 | 325 | 326 | 328 | 330 | 0.75 | 365 | 367 | 368 | 369 | 370 | 372 | 373 | 375 | 376 | 378 | 379 | 381 | 382
0.675 | 322 | 325 | 327 | 328 | 330 | 331 | 333 | 333 | 335 | 336 | 338 | 340 | 0.775 | 374 | 376 | 377 | 378 | 379 | 381 | 382 | 383 | 385 | 386 | 388 | 390 | 391
0.7 | 333 | 335 | 337 | 339 | 340 | 341 | 342 | 343 | 345 | 346 | 348 | 350 | 0.8 | 384 | 385 | 386 | 387 | 388 | 390 | 391 | 392 | 394 | 395 | 397 | 398 | 400
0.725 | 342 | 345 | 347 | 348 | 349 | 351 | 352 | 353 | 354 | 356 | 358 | 360 | 0.825 | 393 | 394 | 395 | 396 | 397 | 399 | 400 | 401 | 403 | 404 | 406 | 407 | 408
0.75 | 352 | 355 | 357 | 358 | 359 | 360 | 361 | 363 | 364 | 366 | 368 | 369 | 0.85 | 402 | 403 | 404 | 405 | 406 | 407 | 409 | 410 | 411 | 413 | 414 | 416 | 417
0.775 | 361 | 365 | 367 | 368 | 369 | 370 | 371 | 372 | 374 | 375 | 377 | 379 | 0.875 | 410 | 411 | 412 | 413 | 415 | 416 | 417 | 418 | 420 | 421 | 423 | 424 | 425
0.8 | 369 | 374 | 376 | 377 | 378 | 379 | 380 | 382 | 383 | 385 | 387 | 389 | 0.9 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 427 | 428 | 429 | 431 | 432 | 434
Exposure to Dose Conversion Factors W/Al for 4.2cm 50/50 Breast
kVp
HVL | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49
0.2 | 120 | 125 | 133 | 138 | 143 | 148 | 156 | 160 | 165 | 169 | 177 | 181 | 185 | 188 | 195 | 198 | 200 | 203 | 208 | 216 | 219 | 222 | 225 | 231 | 234
0.25 | 143 | 148 | 155 | 160 | 165 | 169 | 176 | 181 | 185 | 189 | 196 | 200 | 204 | 207 | 213 | 216 | 219 | 221 | 226 | 233 | 236 | 239 | 241 | 247 | 250
0.3 | 166 | 171 | 177 | 182 | 187 | 191 | 197 | 202 | 206 | 210 | 216 | 220 | 223 | 226 | 232 | 235 | 237 | 239 | 244 | 250 | 253 | 256 | 258 | 264 | 266
0.35 | 189 | 195 | 200 | 204 | 209 | 213 | 218 | 222 | 226 | 230 | 236 | 239 | 243 | 246 | 250 | 253 | 255 | 257 | 262 | 267 | 270 | 273 | 275 | 280 | 282
0.4 | 212 | 217 | 222 | 226 | 231 | 235 | 239 | 243 | 247 | 251 | 255 | 258 | 262 | 265 | 269 | 271 | 273 | 275 | 280 | 285 | 287 | 290 | 292 | 296 | 298
0.45 | 234 | 239 | 244 | 248 | 252 | 256 | 260 | 264 | 268 | 271 | 275 | 278 | 281 | 284 | 287 | 290 | 292 | 294 | 298 | 302 | 304 | 307 | 309 | 313 | 315
0.5 | 256 | 261 | 265 | 269 | 273 | 277 | 280 | 284 | 288 | 291 | 294 | 297 | 300 | 303 | 306 | 308 | 310 | 312 | 316 | 319 | 321 | 324 | 326 | 329 | 331
0.55 | 278 | 282 | 286 | 290 | 293 | 297 | 300 | 304 | 307 | 310 | 313 | 316 | 319 | 322 | 324 | 326 | 328 | 330 | 334 | 336 | 339 | 341 | 343 | 345 | 347
0.6 | 300 | 303 | 307 | 310 | 313 | 317 | 320 | 323 | 326 | 329 | 332 | 335 | 337 | 340 | 342 | 344 | 346 | 348 | 351 | 354 | 356 | 358 | 360 | 362 | 363
0.65 | 321 | 324 | 327 | 330 | 333 | 336 | 339 | 342 | 345 | 347 | 350 | 353 | 355 | 358 | 360 | 362 | 363 | 365 | 368 | 370 | 372 | 374 | 376 | 378 | 380
0.7 | 341 | 344 | 347 | 350 | 352 | 355 | 358 | 361 | 363 | 366 | 368 | 370 | 373 | 375 | 377 | 379 | 380 | 382 | 385 | 387 | 389 | 390 | 392 | 394 | 395
0.75 | 360 | 363 | 366 | 369 | 370 | 374 | 376 | 379 | 381 | 383 | 386 | 388 | 390 | 392 | 394 | 395 | 397 | 398 | 401 | 403 | 404 | 406 | 408 | 409 | 411
0.8 | 379 | 382 | 385 | 387 | 389 | 392 | 394 | 396 | 398 | 401 | 403 | 405 | 407 | 408 | 410 | 412 | 413 | 414 | 417 | 418 | 420 | 421 | 423 | 424 | 426
0.85 | 398 | 400 | 403 | 405 | 407 | 409 | 411 | 413 | 415 | 417 | 419 | 421 | 423 | 425 | 426 | 427 | 429 | 430 | 432 | 434 | 435 | 436 | 438 | 439 | 440
0.9 | 415 | 418 | 420 | 422 | 424 | 426 | 428 | 430 | 432 | 434 | 435 | 437 | 439 | 440 | 442 | 443 | 444 | 445 | 447 | 449 | 450 | 451 | 452 | 454 | 455
0.95 | 432 | 435 | 437 | 438 | 440 | 442 | 444 | 446 | 448 | 449 | 451 | 453 | 454 | 455 | 457 | 458 | 459 | 460 | 462 | 463 | 464 | 465 | 467 | 468 | 469
1 | 448 | 451 | 453 | 454 | 456 | 458 | 460 | 461 | 463 | 465 | 466 | 467 | 469 | 470 | 471 | 472 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482
kV Accuracy/Output |  | 170 174: Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV | [8 empty] | Radcal kV correction for Hologic Selenia/Dimension
Target | kV | mGy/mAs | mGy/s | Target | kV | mGy/mAs | mGy/s | Target | kV | mGy/mAs | mGy/s | [8 empty] | Selenia |  | Dimension
| 24 | 173 177: 28; 177 181: 28 | [8 empty] | Measured | W/Rh | W/Ag | W/Rh | W/Ag | W/Al
Filter | 25 | 172 176: Filter; 173 177: 30; 176 180: Filter; 177 181: 30
| 26 | 173 177: 32; 177 181: 32
| 28 | 173 177: 34; 177 181: 34
| 30 | 173 177: 36; 177 181: 38
| 32 | 173 177: 38
34
36
38
| Slope: | #DIV/0! | #DIV/0! |  | Slope: | #DIV/0! | #DIV/0! |  | Slope: | #DIV/0! | #DIV/0!
| Intercept: | #DIV/0! | #DIV/0! |  | Intercept: | #DIV/0! | #DIV/0! |  | Intercept: | #DIV/0! | #DIV/0!
| R^2: | #DIV/0! | #DIV/0! |  | R^2: | #DIV/0! | #DIV/0! |  | R^2: | #DIV/0! | #DIV/0!
Half Value Layer
Target | kV | HVL | Target | kV | HVL | Target | kV | HVL
| 24 | 172 176: 28; 175 179: 28
Filter | 25 | 171 175: Filter; 172 176: 30; 174 178: Filter; 175 179: 30
| 26 | 170 174: 0; 172 176: 32; 175 179: 32
| 28 | 172 176: 34; 175 179: 34
| 30 | 170 174: 0; 172 176: 36; 173 177: 0; 175 179: 36
| 32 | 172 176: 38; 173 177: 0; 175 179: 38; 176 180: 0
| 34 | 170 174: #DIV/0!
| 36 | 170 174: 0
| 38 | 170 174: 0
| Slope: | 170 174: #DIV/0!; 172 176: Slope:; 173 177: 0; 175 179: Slope:; 176 180: #DIV/0!
| Intercept: | 170 174: #DIV/0!; 172 176: Intercept:; 173 177: 0; 175 179: Intercept:; 176 180: #DIV/0!
| R^2: | 170 174: #DIV/0!; 172 176: R^2:; 173 177: #DIV/0!; 175 179: R^2:; 176 180: #DIV/0!
Target | Filter | 170 174: C; 172 176: Critiera:; 173 177: Min HVL > kV/100 (no paddle)
Mo | Mo | 170 174: 0.12; 173 177: Min HVL > kV/100 + 0.03 (with paddle)
Mo | Rh | 170 174: 0.19; 173 177: Max HVL < kV/100 + C
Rh | Rh | 170 174: 0.22
DGN | #VALUE! | 174 178: Combo
Mo/Mo | Mo/Rh | Rh/Rh | W/Rh | W/Ag | W/Al | W/Rh | W/Ag | W/Al
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Radcal kV Correction for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Rh | < | 27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| >= | 27.586 | -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 | kV^5 | kV^6
W/Ag | < | 30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| >= | 30.1 | -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
DGN Mo/Mo: Radcal kV Correction for Hologic Dimension |  | Barracuda/MPD kVp correction table for Hologic Selenia
| kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | W/Rh (RQ M6) |  |  | W/Ag (RQ M10)
W/Rh | < | 26.9 | 138.8867 | -10.7264 | 0.26216 | -0.00082 |  |  |  | kV | LFS | SFS | kV | LFS | SFS
| >= | 26.9 | -5009.7752 | 605.732 | -27.3019 | 0.54671 | -0.0040986 |  |  | 22 | 0.2 | 0.1 | 22 | 0.2 | -0.2
[10 empty] | 23 | 0.2 | -0.1 | 23 | 0.4 | -0.1
W/Ag | kV |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | 24 | 0.1 | -0.4 | 24 | 0.4 | 0
| < | 28.7 | 296.3419 | -31.6292 | 1.1803 | -0.01417 |  |  |  | 25 | 0.1 | -0.3 | 25 | 0.5 | -0.1
#VALUE! Mo/Rh: >=; #VALUE! Rh/Rh: 28.7; #VALUE! W/Rh: 4.8345; #VALUE! W/Ag: 0.91924 |  | 26 | 0 | -0.2 | 26 | 0.5 | -0.2
[10 empty] | 27 | 0.1 | -0.3 | 27 | 0.7 | -0.2
W/Al |  |  | kV^0 | kV^1 | kV^2 | kV^3 | kV^4 |  |  | 28 | 0.2 | -0.5 | 28 | 0.9 | -0.1
#VALUE! Mo/Rh: <; #VALUE! Rh/Rh: 28.7; #VALUE! W/Rh: 49.3111; #VALUE! W/Ag: -2.9302; #VALUE! W/Al: 0.07379 |  | 29 | 0.4 | -0.2 | 29 | 0.8 | -0.3
#VALUE! Rh/Rh: 28.7 < kV < 30.1; #VALUE! W/Rh: -24.875; #VALUE! W/Ag: 1.8032 |  | 30 | 0.6 | 0 | 30 | 0.8 | -0.4
#VALUE! Mo/Rh: >=; #VALUE! Rh/Rh: 30.1; #VALUE! W/Rh: -4.8346; #VALUE! W/Ag: 1.1571 |  | 31 | 0.8 | 0.1 | 31 | 0.8 | -0.3
[10 empty] | 32 | 1 | 0.2 | 32 | 0.8 | -0.3
DGN Mo/Mo: CNR Correction |  | 33 | 1 | 0.4 | 33 | 0.9 | -0.3
DGN Mo/Mo: Conventional (XX8xxxxx series detectors) |  | 34 | 1.1 | 0.6 | 34 | 1.1 | -0.3
Thickness (cm) | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 |  | 35 | 1.2 | 0.6 | 35 | 1.1 | -0.2
2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1.25 |  | 36 | 1.4 | 0.6 | 36 | 1.2 | -0.2
4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1.2 |  | 37 | 1.5 | 0.7 | 37 | 1.1 | -0.1
6 | 1.04 | 1.02 | 1.01 | 1.08 | 1.35 | 1.35 | 1.31 | 1.35 |  | 38 | 1.6 | 0.9 | 38 | 1.1 | 0
8 | 1.23 | 1.2 | 1.18 | 1.1 | 1.82 | 1.85 | 1.82 | 1.85 |  | 39 | 1.7 | 0.8 | 39 | 1.3 | 0.3
Magnification (XX8xxxxx series detectors)
Thickness (cm) | 0 | 1 | 2 | 3 | 4
2 | 1.15 | 1.15 | 1.15 | 1.15 | 1.5
4 | 1.15 | 1.15 | 1.15 | 1.15 | 1.5
6 | 1.19 | 1.18 | 1.18 | 1.18 | 1.55
8 | 1.28 | 1.24 | 1.22 | 1.29 | 1.67
Stereotactic biopsy (XX8xxxxx series detectors) |  |  |  | CE2D (XX8xxxxx series detectors)
Thickness (cm) | 0 |  |  | Thickness | 0 Low | 0 High
2 | 0.91 |  |  | 2 | 1 | 0.72
4 | 1 |  |  | 4 | 1 | 0.9
6 | 1.32 |  |  | 6 | 1.91 | 1.71
8 | 1.88 |  |  | 8 | 1.81 | 2.22
For software versions Dimensions 1.8.2 - 1.9.x and 3Dimensions 2.0.x and detector SN XX8xxxxx
Tomo 15 Proj STD (XX8xxxxx series detectors) |  |  |  |  | Tomo 15 Proj ENH (XX8xxxxx series detectors)
Thickness (cm) | 0 | 1 | 2 |  | Thickness | 0
2 | 0.7 | 0.7 | 0.88 |  | 2 | 0.57
4 | 0.91 | 0.91 | 0.89 |  | 4 | 0.91
6 | 1.55 | 1.55 | 1.59 |  | 6 | 1.68
8 | 2.78 | 2.78 | 2.68 |  | 8 | 1.93
For software versions Dimensions 1.10.x+ and 3Dimensions 2.1.x+ and detector SN XX8xxxxx
Tomo 15 Proj STD (XX8xxxxx series detectors)
Thickness (cm) | 0 | 1 | 2
2 | 0.64 | 0.7 | 0.88
4 | 0.91 | 0.91 | 0.89
6 | 1.55 | 1.55 | 1.59
8 | 2.78 | 2.78 | 2.68
CNR Correction (Selenia Dimension W/Rh)
Thickness (cm) | 0 |  | 1 |  | 2 |  | 2
| LFS | SFS | LFS | SFS | LFS | SFS | LFS | SFS
0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
3 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
5 | 1 | 1 | 1 | 1 | 1.15 | 1 | 1 | 1
6 | 1 | 1 | 1 | 1.15 | 1.15 | 1 | 1 | 1
7 | 1.1 | 1 | 1.1 | 1.15 | 1.15 | 1 | 1 | 1
8 | 1.15 | 1 | 1.15 | 1.15 | 1.15 | 1 | 1 | 1
9 | 1.1 | 1 | 1.1 | 1.15 | 1.15 | 1 | 1 | 1
10 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
11 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
12 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
13 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
14 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
CNR Correction (Affirm Prone Biopsy)
| Standard biopsy |  |  | Tomo biopsy
Thickness (cm) | 0 | 1 | Thickness (cm) | 0 | 1
2 | 0.71 | 0.69 | 2 | 0.88 | 0.65
4 | 0.94 | 0.98 | 4 | 0.94 | 0.96
6 | 1.3 | 1.09 | 6 | 1.53 | 1.61
8 | 1.71 | 1.53 | 8 | 2.32 | 2.91
Note on A135: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.2, MAN-03706 Rev 10
Note on A141: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.2, MAN-03706 Rev 10
Note on A147: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.3, MAN-03706 Rev 10
Note on E147: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.4, MAN-03706 Rev 10
Note on A155: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.5.1, MAN-03706 Rev 10
Note on F155: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.6, MAN-03706 Rev 10
Note on A163: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Table D.5.2, MAN-03706 Rev 10

## DataPage
Sheet | MAMMO_HOL
SurveyID
Light field
Resolution | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0
2D MGD
3D MGD
Combo MGD
kV Accuracy | / | kV | Output | / | kV | Output | / | kV | Output
Output | 24 |  |  | 28 |  |  | 28
| 25 |  |  | 30 |  |  | 30
| 26 |  |  | 32 |  |  | 32
| 28 |  |  | 34 |  |  | 34
| 30 |  |  | 36 |  |  | 36
| 32 |  |  | 38
34
Linearity | /
20
50
100
300
HVL | / | 24
| / | 25
| / | 28
| / | 32
| / | 28
| / | 30
| / | 32
| / | 34
| / | 28
| / | 30
| / | 32
| / | 34
| / | 36
Acq | 2D | Mag | 3D | Film
Fibers | 0 | 0 | 0 | 0
Specks | 0 | 0 | 0 | 0
Masses | 0 | 0 | 0 | 0
SNR target/filter | /
SNR | #DIV/0!
CNR | #DIV/0!
Note on B4: Eugene Mah: Target/Filter
Note on B5: Eugene Mah: kV
Note on B6: Eugene Mah: Line pairs/mm

## Corrected kV
GENERAL
Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
Read the true corrected kV in the adjacent cell to the right.
If that cell is light green, the corrected kV is accurate.
If the cell is red, the entered kV is out of range and the correction should not be used.
If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.
Tungsten-Rhodium (W / Rh)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
| Radcal kV | True kV
| 29 | 30.2821
Tungsten-Silver (W / Ag)
Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
| Radcal kV | True kV
| 37.1 | 38.9675
27.586 | -8375.0728 | 975.9254 | -37.9137 | 0.49087 | 0 | 0 | 0
| -9984.6168 | 1436.5245 | -82.5051 | 2.3656 | -0.033867 | 0.00019369 | 0
30.1 | -540847.6955 | 100186.2336 | -7418.479 | 274.4766 | -5.0744 | 0.037501 | 0
| -11057.7739 | 1297.2286 | -56.9892 | 1.1116 | -0.0081234 | 0 | 0

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
Selenia
Selenia Dimensions
Affirm Prone SBB