## Sheet1
1					Revision 1.0-20140918					Print Area
2					Medical University of South Carolina					Medical University of South Carolina					Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3					Charleston, South Carolina					Charleston, South Carolina
4					All:	Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5					Mammography System Compliance Inspection					Mammography System Compliance Inspection
6					Measurement Parameter	Last Year		This Year
7					Date:					Inspector:				Print Area		Change	Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8					System Information					Previous Date:					System Information					Date
9			Location					Location			Input Changes Only				Input Changes Only				Inspector					kV Set	mAs Set	Add Filt	Target	Filter	kVp	ms	Dose (mGy)	Dose rate (mGy/s)	Total filtration (mm Al)	HVL (mm Al)
10				Facility:					Site Number:					Facility:				Site Number:					Facility:					24	50	0.0	Mo	Mo	23.47	523.2	3.621
11				Department:					Authorized Use:					Department:				Authorized Use:					Department:					24	50	0.0	Mo	Mo			3.624
12				Area/Division:					Date of Installation:					Area/Division:				Date of Installation:					Area/Division:					24	50	0.2	Mo	Mo			2.207
13				Survey ID:					Number of X-ray Tubes:					Survey ID:				Number of X-ray Tubes:					Survey ID:					24	50	0.2	Mo	Mo			2.207
14					Room Number:				Accession Number:					Room Number:					24	50	0.3	Mo	Mo			1.774
15			X-Ray Generator					Site Number:					24	50	0.3	Mo	Mo			1.774
16				Manufacturer:					Manufacture Date:					X-Ray Generator					Authorized Use:					25	50	0.0	Mo	Mo	25.67	502.7	4.21
17				Model:					Serial Number:					Manufacturer:				Manufacture Date:					Date of Installation:					25	50	0.0	Mo	Mo			4.23
18				Max kVp:					Max mA:					Model:				Serial Number:					Number of X-ray Tubes:					25	50	0.2	Mo	Mo			2.62
19					Max kVp:				Max mA:					Accesssion Number:					25	50	0.2	Mo	Mo			2.62
20			X-Ray Tube 1					X-Ray Generator					25	50	0.3	Mo	Mo			2.122
21				Tube Designation/Use:					Target:	Mo					X-Ray Tube 1					Target:	Mo	Mo				Manufacturer:					25	50	0.3	Mo	Mo			2.122
22				Manufacture Date:					Tube Designation/Use:					Mo	Mo				Model:					26	50	0.0	Mo	Mo	26.83	502.9	4.822
23			Insert					Filter:	Mo					Manufacture Date:					Max kVp:					28	20	0.0	Mo	Mo	29.06	662.3	2.436
24				Manufacturer:					Rh					Insert					Filter:	Mo	Mo				Manufacture Date:					28	50	0.0	Mo	Mo	28.93	502.8	6.134
25				Model:					Manufacturer:					Rh	Rh				Serial Number:					28	50	0.0	Mo	Mo	28.93	502.9	6.133
26				Serial Number:					Filtration					Model:					Max mA:					28	50	0.0	Mo	Mo	28.92	502.2	6.129
27			Housing					Inherent:					Serial Number:				Focal Spot Sizes (mm)					X-Ray Tube 1					28	50	0.0	Mo	Mo	28.96	502.4	6.127
28				Manufacturer:					Added:					Housing					Large:					Tube Designation/Use:					28	50	0.3	Mo	Mo			3.287
29				Model:					Manufacturer:				Small:					Manufacture Date:					28	50	0.3	Mo	Mo			3.287
30				Serial Number:					Model:					Manufacturer:					28	50	0.4	Mo	Mo			2.737
31					Serial Number:				Filtration					Model:					28	50	0.4	Mo	Mo			2.737
32					Inherent:					Serial Number:					28	100	0.0	Mo	Mo	28.94	1006	12.3
33					Technique Chart					Added:					Manufacturer:					28	300	0.0	Mo	Mo	28.94	3370	36.95
34					Model:					30	50	0.0	Mo	Mo	30.44	500.8	7.551
35		Compressed	Fatty Breast			50% Fatty -			Dense Breast					Serial Number:					32	50	0.0	Mo	Mo	32	502.4	9.11
36		Breast				50% Dense Breast					Large:					32	50	0.0	Mo	Mo			9.16
37		Thickness	Target/Filter	kVp	Density	Target/Filter	kVp	Density	Target/Filter	kVp	Density					Inspection Results					Small:					32	50	0.3	Mo	Mo			5.131
38		<3 cm					Enter 1 for YES, 2 for NO, 3 for NA					Micro:					32	50	0.3	Mo	Mo			5.131
39		3 to 5 cm					DHEC Registration sticker is present, clearly visible and legible					Inherent:					32	50	0.4	Mo	Mo			4.337
40		5 to 7 cm					DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					Added:					32	50	0.4	Mo	Mo			4.337
41		>7 cm					Radiation warning label posted on the generator control panel					Technique Chart					34	50	0.0	Mo	Mo
42					Operator manuals are available.					Fatty Breast					28	50	0.0	Mo	Rh	28.24	503.1	3.855
43					Inspection Results					Monthly radiation monitoring reports are posted.					Target/Filter:					28	50	0.0	Mo	Rh			3.871
44		Rule Number					Compliance					kVp:					28	50	0.4	Mo	Rh			1.997
45		DHEC RHB 2.5.1.1		DHEC Registration sticker is present, clearly visible and legible					TBD					Mammography Unit Assembly Evaluation					Density:					28	50	0.4	Mo	Rh			1.997
46		DHEC RHB 10.2.1		DHEC form SC-RAH-20 “Notice to Employees” posted or referenced					TBD				Free-standing unit is mechanically stable					Target/Filter:					28	50	0.5	Mo	Rh			1.719
47		DHEC RHB 4.3.1		Radiation warning label posted on the generator control panel					TBD				All moving parts move smoothly, without obstructions to motion					kVp:					28	50	0.5	Mo	Rh			1.719
48				Operator manuals are available.					TBD				All locks and detents work properly					Density:					30	50	0.0	Mo	Rh	29.45	502.2	4.843
49				Monthly radiation monitoring reports are posted.					TBD				Image receptor holder assembly is free from vibrations					Target/Filter:					30	50	0.0	Mo	Rh			4.873
50					Mammography Unit Assembly Evaluation					Image receptor slides smoothly into holder assembly					kVp:					30	50	0.4	Mo	Rh			2.561
51				Free-standing unit is mechanically stable					TBD				Image receptor is held securely by assembly in any orientation					Density:					30	50	0.4	Mo	Rh			2.561
52				All moving parts move smoothly, without obstructions to motion					TBD				Compressed breast thickness scale accurate to +0.5 cm, reproducible to +2 mm					Target/Filter:					30	50	0.5	Mo	Rh			2.213
53				All locks and detents work properly					TBD				Patient or operator is not exposed to sharp or rough edges, or other hazards					kVp:					30	50	0.5	Mo	Rh			2.213
54				Image receptor holder assembly is free from vibrations					TBD				Operator technique control charts are posted					Density:					32	50	0.0	Mo	Rh	30.85	502.1	5.905
55				Image receptor slides smoothly into holder assembly					TBD				Operator protected during exposure by adequate radiation shielding					50/50 Breast					32	50	0.0	Mo	Rh			5.934
56				Image receptor is held securely by assembly in any orientation					TBD				All indicator lights working properly					Target/Filter:					32	50	0.4	Mo	Rh			3.176
57				Compressed breast thickness scale accurate to 0.5 cm, reproducible to 2 mm					TBD				Auto decompression can be overridden to maintain compression (and status displayed)					kVp:					32	50	0.4	Mo	Rh			3.176
58				Patient or operator is not exposed to sharp or rough edges, or other hazards					TBD				Manual emergency compression release can be activated in the event of a power failure					Density:					32	50	0.5	Mo	Rh			2.749
59				Operator technique control charts are posted					TBD					Target/Filter:					32	50	0.5	Mo	Rh			2.749
60				Operator protected during exposure by adequate radiation shielding					TBD					MQSA Equipment Standards					kVp:					34	50	0.0	Mo	Rh	33.11	663.9	7.005
61				All indicator lights working properly					TBD				Tube-receptor assembly can be fixed into any operating position and does not move once fixed into position					Density:					34	50	0.0	Mo	Rh			7.043
62				Autodecompression can be overridden to maintain compression					TBD				Tube-receptor assembly locking mechanism does not fail during power interruption					Target/Filter:					34	50	0.4	Mo	Rh			3.813
63				Manual emergency compression release in the event of a power failure					TBD				Image Receptor Sizes 18 x 24 & 24 x 30 are available					kVp:					34	50	0.4	Mo	Rh			3.813
64					Moving grids are present for all image receptor sizessizes					Density:					34	50	0.5	Mo	Rh			3.309
65		Date:					Inspector:				Magnification can be done without grid					Target/Filter:					34	50	0.5	Mo	Rh			3.309
66		Room Number					Survey ID:				Beam limiting device allows useful beam to extend beyond chest wall edge of receptor					kVp:					36	50	0.0	Mo	Rh	35.42	663.5	8.138
1					Medical University of South Carolina			Average illumination of the light field is at least 160 lux					Density:					38	50	0.0	Mo	Rh	37.65	713.6	9.411
2					Inspection Results					Mammography System Compliance Inspection			Magnification mode is available for non-interventional problem solving procedures					Dense Breast					28	50	0	Mo	Mo
3		Rule Number					MQSA Equipment Standards					System provides magnification factor between 1.4 – 2.0					Target/Filter:
4		900.12(b)(3i)		Tube-receptor can be fixed into any operating position and does not move					TBD				Prior indication of selected focal spot is provided					kVp:
5		900.12(b)(3ii)		Tube-receptor locking mechanism does not fail during power interruption					TBD				Prior indication of selected target is provided					Density:
6		900.12(b)(4i)		Image Receptor Sizes 18 x 24 & 24 x 30 are available					TBD				Post exposure display of target/focal spot is provided					Target/Filter:
7		900.12(b)(4ii)		Moving grids are present for all image receptor sizessizes					TBD				Initial power driven compression can be activated hands-free from both sides					kVp:
8		900.12(b)(4iii)		Magnification can be done without grid					TBD				Fine adjustment of compression can be activated from both sides					Density:
9		900.12(b)(5i)		Beam limiting device allows useful beam to extend beyond chest wall edge					TBD				Compression paddles match the sizes of all image receptors					Target/Filter:
10		900.12(b)(5ii)		Average illumination of the light field is at least 160 lux					TBD				Compression paddles do not flex by more than 1 cm					kVp:
11		900.12(b)(6i)		Magnification mode available for non-interventional problem solving					TBD				Compression paddles are flat and parallel to the breast support table					Density:
12		900.12(b)(6ii)		System provides magnification factor between 1.4 – 2.0					TBD				Chest wall edge of the compression paddle is straight and parallel to the edge of the receptor					Target/Filter:
13		900.12(b)(7i)		Prior indication of selected focal spot is provided					TBD				Chest wall edge may be bent upwards for patient comfort and does not appear in the image					kVp:
14		900.12(b)(7i)		Prior indication of selected target is provided					TBD				Manual selection of mAs or mA and/or time is available					Density:
15		900.12(b)(7iii)		Post exposure display of target/focal spot is provided					TBD				Prior indication of selected technique factors is provided
16		900.12(b)(8)(iB)		Initial power driven compression can be activated hands-free from both sides					TBD				Post exposure display of technique factors is provided for AEC exposures					MGD:	1.6		1.7138603366
17				Fine adjustment of compression can be activated from both sides					TBD				AEC mode is operable in all combinations of equipment configuration					mGy/mAs:	0.124		0.122615
18		900.12(b)(8)(iiA)		Compression paddles match the sizes of all image receptors					TBD				AEC sensor positioning is flexible					mGy/s:	12.047		12.1986780224
19		900.12(b)(8)(iiB)		Compression paddles do not flex by more than 1 cm					TBD				AEC sensor position and size is indicated on the compression paddle					Image Quality (Printer)
20		900.12(b)(8)(iiC)		Compression paddles are flat and parallel to the breast support table					TBD				Selected position of the AEC sensor is clearly indicated					kV:
21		900.12(b)(8)(iiD)		Chest wall edge of the compression paddle is straight and parallel to receptor					TBD				AEC density control setting is provided					mA:
22		900.12(b)(8)(iiE)		Chest wall edge may be bent upwards and does not appear in the image					TBD				X-ray film used is appropriate for mammography					Bkg Density:
23		900.12(b)(9i)		Manual selection of mAs or mA and/or time is available					TBD				Intensifying screens used are appropriate for mammography and matched to the film used					Inside disk:
24		900.12(b)(9ii)		Prior indication of selected technique factors is provided					TBD				Film processing chemicals are appropriate for the film being used					Outside disk:
25		900.12(b)(9iii)		Post exposure display of technique factors is provided for AEC exposures					TBD				Lights for film illumination i.e. hot lights are available					Difference:
26		900.12(b)(10i)		AEC mode is operable in all combinations of equipment configuration					TBD				Film masking devices are available to all interpreting physicians					Fibers:			5
27		900.12(b)(10ii)		AEC sensor positioning is flexible					TBD					Specks:			4
28				AEC sensor position and size is indicated on the compression paddle					TBD					Masses:			4
29				Selected position of the AEC sensor is clearly indicated					TBD					Mammography Phototimer Technique Chart					Image Quality (Acq Unit)
30		900.12(b)(10iii)		AEC density control setting is provided					TBD					kV:
31		900.12(b)(11)		X-ray film used is appropriate for mammography					TBD			Compressed	Fatty Breast			50% Fatty -			Dense Breast					mA:
32		900.12(b)(12)		Intensifying screens used are appropriate for mammography and matched to film					TBD			Breast				50% Dense Breast					Exp Ind:
33		900.12(b)(13)		Film processing chemicals are appropriate for the film being used					TBD			Thickness	Target/Filter	kVp	Density	Target/Filter	kVp	Density	Target/Filter	kVp	Density			Fibers:
34		900.12(b)(14)		Lights for film illumination i.e. hot lights are available					TBD			<3 cm					Specks:
35		900.12(b)(15)		Film masking devices are available to all interpreting physicians					TBD			3 to 5 cm					Masses:
36					5 to 7 cm					kV:
37					>7 cm					mA:
38					Exp Ind:
39					Compressed	Fatty Breast			50% Fatty -			Dense Breast					Fibers:
40					Breast				50% Dense Breast					Specks:
41					Thickness	Target/Filter	kVp	Density	Target/Filter	kVp	Density	Target/Filter	kVp	Density			Masses:
42					<3 cm
43					3 to 5 cm					SNR:	45.8		45.8
44					5 to 7 cm					CNR:	9.9		9.9
45					>7 cm
46
47					Inspection Results
48					Compression Force Indicator
49					Indicated	Measured	After 20s	Slip
50					Power
51					Manual
52					Criteria:	Minimum compression force – 25 lbs
53					Maximum compression force – 45 lbs
54
55					Compression Thickness Indicator
56					BR-12 (cm)	Indicated	Difference
57					1	1	0
58					2	2	0
59					4	4	0	Mean:	4.00
60					4	4	0	StDev:	0.00
61					4	4	0
62					4	4	0		Criteria:	Difference between indicated and actual thickness is <= 2mm
63					6	6.1	0.1
64					8	8.1	0.1
65		Date:					Inspector:
66		Room Number					Survey ID:				Paddle flex is less than 1 cm			Pass	Compression thickness indicator:
1					Medical University of South Carolina			Overall unit assembly
2					Inspection Results					Mammography System Compliance Inspection
3		Compression Force Indicator					Artifact Evaluation
4				Indicated	Measured	After 20s	Slip					Receptor:	18x24	18x24	24x30	24x30
5			Power					Target:	Mo	Mo	Mo	Mo
6			Manual					Acceptable:					Filter:	Mo	Mo	Rh	Rh
7			Criteria:	Minimum compression force – 25 lbs					Focal spot:	Large	Small	Large	Small
8				Maximum compression force – 45 lbs					kV:
9					mAs:
10		Compression Thickness Indicator					EI:
11			BR-12 (cm)	Indicated	Difference				Paddle flex is less than 1 cm:	TBD					Acceptable:
12			1	1	0				Overall unit assembly:	TBD
13			2	2	0				Compression thickness indicator:	Pass					Comments:
14			4	4	0	Mean:	4.00					Do not exceed 390 chars
15			4	4	0	StDev:	0.00					New:
16			4	4	0	Criteria:	Difference between indicated and actual thickness is <= 2mm
17			4	4	0					Light Field Brightness
18			6	6.1	0.1					Measurements (lux)					Average Illumination
19			8	8.1	0.1					1st quad.	2nd quad.	3rd quad.	4th quad.		(lux)
20					321	290	340	334		321.3
21		Artifact Evaluation
22			Receptor:	18x24	18x24	24x30	24x30					Light Field/X-ray Field Alignment
23			Target:	Mo	Mo	Mo	Mo					SID (cm):	65					% Deviation
24			Filter:	Mo	Mo	Rh	Rh					24x30	18x24			24x30	18x24
25			Focal spot:	Large	Small	Large	Small					24x30	Left	Center	Right	24x30	Left	Center	Right
26			kV:					Left:	0	0.4	0.1	0.2	0.00%	0.62%	0.15%	0.31%
27			mAs:					Right:	0	0	0.1	0	0.00%	0.00%	0.15%	0.00%
28			EI:					Anterior:	0	0.2	0.2	-0.3	0.00%	0.31%	0.31%	-0.46%
29			Acceptable:					Chest Wall:	0.2	0.2	0.1	0	0.31%	0.31%	0.15%	0.00%
30			Criteria:	No visible artifacts in the images					Enter deviation in cm			Left-Right	0.00%	0.62%	0.31%	0.31%
31			Comments:					Ant-Chest	0.31%	0.62%	0.46%	-0.46%
32					Criteria:	Sum of left and right edge deviation is less than 2% SID
33					Sum of anterior and chest wall edge deviation is less than 2% SID
34
35		Light Field Brightness					X-ray Field/Image Receptor Alignment
36			Measurements (lux)					Average Illumination					% Deviation
37			1st quad.	2nd quad.	3rd quad.	4th quad.		(lux)					24x30	18x24			24x30	18x24
38			321	290	340	334		321.3					24x30	Left	Center	Right	24x30	Left	Center	Right
39			Criteria:	Average light field illumination >= 160 lux.					Left:	0.1	0.2	0.2	0.2	0.15%	0.31%	0.31%	0.31%
40					Right:	0.4	0.4	0.3	0.4	0.62%	0.62%	0.46%	0.62%
41		Light fField/X-ray Field Alignment				X-ray Field/Image Receptor Alignment					Anterior:	0	0.4	0.5	0.5	0.00%	0.62%	0.77%	0.77%
42		SID (cm):					Chest Wall:	0.8	0	0.2	0	1.23%	0.00%	0.31%	0.00%
43		65		Light Field/X-ray Field % Deviation				X-ray Field/Image Receptor % Deviation					Compression Paddle	0.2	0.5	0.5	0.5	0.31%	0.77%	0.77%	0.77%
44				24x30	18x24			24x30	18x24					Enter deviation in cm
45				24x30	Left	Center	Right	24x30	Left	Center	Right				Criteria:	X-ray field does not extend beyond image receptor by more than 2% SID
46			Left:	0.00%	0.62%	0.15%	0.31%	0.15%	0.31%	0.31%	0.31%					X-ray field extends beyond the chest wall edge.
47			Right:	0.00%	0.00%	0.15%	0.00%	0.62%	0.62%	0.46%	0.62%					X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side
48			Anterior:	0.00%	0.31%	0.31%	-0.46%	0.00%	0.62%	0.77%	0.77%					X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side
49			Chest Wall:	0.31%	0.31%	0.15%	0.00%	1.23%	0.00%	0.31%	0.00%					Difference between chest wall edge of the receptor and compression paddle is less than 1% SID
50			Compression Paddle					0.31%	0.77%	0.77%	0.77%
51			Criteria:	Sum of left and right edge deviation is less than 2% SID					System Resolution
52				Sum of anterior and chest wall edge deviation is less than 2% SID					Target:
53				X-ray field does not extend beyond image receptor by more than 2% SID					kV:	28
54				X-ray field extends beyond the chest wall edge.					mAs:	65
55				X-ray field does not fall within the image receptor by more than 2% SID on the right, left or chest wall side					lp/mm:	8
56				X-ray field dose not fall within the image receptor by more than 4% SID on the anterior side					Acceptable:	Pass
57				Difference between chest wall edge of the receptor and compression paddle is less than 1% SID					Criteria:	Limiting system resolution must be 7 lp/mm or higher
58
59		System Resolution					AEC Thickness Tracking
60			Target:					AEC Mode:			AEC Position:
61			kV:	28			Acceptable:	Pass					Thickness					Reported
62			mAs:	65					(cm)	Filter	kV	mAs	Pixel Val	EI	MGD (mGy)	Deviation
63			lp/mm:	8					2	Mo	24	35.7		484	0.651	-3.22%
64			Criteria:	Limiting system resolution must be 7 lp/mm or higher					4	Mo	28	55.1		483	1.26	-3.42%
65		Date:					Inspector:				4	Mo	28	56.2		501	1.29	0.17%
66		Room Number					Survey ID:				4	Mo	28	57		501	1.31	0.17%
1					Medical University of South Carolina			4	Mo	28	56.9		508	1.31	1.57%
2					Inspection Results					Mammography System Compliance Inspection			6	Rh	32	82.6		501	1.85	0.17%
3		AEC Thickness Tracking					8	Rh	33	159.2		504	3.23	0.77%
4		AEC Mode:			AEC Position:					4 cm Mag	Mo	28	43.4		519	3.37	3.77%
5			Thickness					Reported					Average:		500.125
6			(cm)	Filter	kV	mAs	Pixel Val	EI	MGD (mGy)	Deviation					Criteria:	Mean pixel value of each image is within 10% of the average pixel value
7			2	Mo	24	35.7		484	0.651	-3.22%
8			4	Mo	28	55.1		483	1.26	-3.42%				AEC Density Control
9			4	Mo	28	56.2		501	1.29	0.17%				AEC Mode:			AEC Position:			Reported
10			4	Mo	28	57		501	1.31	0.17%					Density	kV	mAs	Pixel Val	EI	MGD (mGy)	Change		Low	High
11			4	Mo	28	56.9		508	1.31	1.57%					-3	28	33.2		297	0.76	0.60		0.5	0.61
12			6	Rh	32	82.6		501	1.85	0.17%					-2		40.2		355	0.92	0.71		0.63	0.77
13			8	Rh	33	159.2		504	3.23	0.77%					-1		49.1		435	1.13	0.87		0.77	0.94
14			4 cm Mag	Mo	28	43.4		519	3.37	3.77%					0		56.3		498.25	1.29
15					Average:		500.125					1		63.9		560	1.47	1.12		1.04	1.27
16			Criteria:	Mean pixel value of each image is within 10% of the average pixel value					2		72.8		632	1.67	1.27		1.17	1.43
17					3		81.7		711	1.88	1.43		1.31	1.6
18		AEC Density Control					4		87.8		757	2.02	1.52		1.44	1.76
19		AEC Mode:			AEC Position:			kV:	28
20					Reported					Criteria:	Change in pixel value at each density step is within specified ranges
21			Density	mAs	Pixel Val	EI	MGD (mGy)	Change
22			-3	33.2		297	0.76	0.60	Pass					Mean Glandular Dose
23			-2	40.2		355	0.92	0.71	Pass					Phantom:			kV:	28
24			-1	49.1		435	1.13	0.87	Pass					Target/Filter:	Mo/Mo		mA:	100
25			0	56.3		498	1.29					Reported		Air KERMA
26			1	63.9		560	1.47	1.12	Pass					mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:	0.34
27			2	72.8		632	1.67	1.27	Pass					69.9	494	1.55	0.699	8.60		ESE (mGy):	8.69
28			3	81.7		711	1.88	1.43	Pass					69.9	500	1.55	0.699	8.60		DGN (mrad/R):	172.8
29			4	87.8		757	2.02	1.52	Pass					71	507	1.57	0.710	8.73		MGD (mGy):	1.71
30			Criteria:	Change in pixel value at each density step is within specified ranges					71.7	515	1.59	0.717	8.82		Prev MGD (mGy):	1.60
31					Average:	70.6	504	1.57	0.706	8.69		% Change:	7.1%
32		Mean Glandular Dose					Coeff of Var:	1.25%	1.80%	1.22%	1.25%	1.25%
33			Phantom:		kV:	28
34			Target/Filter:	Mo/Mo	mA:	100					Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:	9.5%
35					Reported		Air KERMA					Deviation from reported MGD is less than15%					mAs at 3 mGy:	123.6
36				mAs	EI	MGD (mGy)	Time (s)	(mGy)		HVL:	0.34
37				69.9	494	1.55	0.699	8.60		ESE (mGy):	8.69			kVp Accuracy and Output
38				69.9	500	1.55	0.699	8.60		DGN (mrad/R):	172.8			Meter:			Calibration Date:
39				71.0	507	1.57	0.710	8.73		MGD (mGy):	1.71			Serial #:			Calibration Due:
40				71.7	515	1.59	0.717	8.82		Prev MGD (mGy):	1.60
41			Average:	70.6	504	1.565	0.706	8.69		% Change:	7.1%					Measured Values
42			Coeff of Var:	1.25%	1.80%	1.22%	1.25%	1.25%					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
43			Criteria:	MGD is less than 3 mGy					Deviation from Reported MGD:	9.5%				Mo	Mo	24	50	23.47	523.2	3.621	0.072	6.92
44				Deviation from reported MGD is less than15%					mAs at 3 mGy:	123.6				Mo	Mo	25	50	25.67	502.7	4.21	0.084	8.37
45					Mo	Mo	26	50	26.83	502.9	4.822	0.096	9.59
46		Signal to Noise Ratio					Mo	Mo	28	50	28.93	502.8	6.134	0.123	12.20
47			kV:	28			Current	Previous	Change	Acceptable					Mo	Mo	30	50	30.44	500.8	7.551	0.151	15.08
48			mA:	71.7		SNR:	46.8	45.8	2.17%	Pass					Mo	Mo	32	50	32	502.4	9.11	0.182	18.13
49			Target:			CNR:	10.0	9.9	1.38%	Pass					Mo	Mo	34	50
50			Filter:					Mag:	Mo	Mo	28	50
51			Criteria:	SNR must be >= 40					Criteria:	kV accuracy is <= 5%
52				CNR does not change by more than 15%					Measured Values
53					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
54		Image Quality Evaluation					Mo	Rh	28	50	28.24	503.1	3.855	0.077	7.66
55			Printer				Acq Workstation				Review Workstation				Mo	Rh	30	50	29.45	502.2	4.843	0.097	9.64
56			kV:				Normal	Mag			Left	Right			Mo	Rh	32	50	30.85	502.1	5.905	0.118	11.76
57			mA:			kV:					Workstation:				Mo	Rh	34	50	33.11	663.9	7.005	0.140	10.55
58			Background Density:			mA:				Fibers:					Mo	Rh	36	50	35.42	663.5	8.138	0.163	12.27
59			Inside disk:			Exposure Index:				Specks:					Mo	Rh	38	50	37.65	713.6	9.411	0.188	13.19
60			Outside disk:			Fibers:				Masses:					Criteria:	kV accuracy is <= 5%
61			Difference:			Specks:					Workstation:
62			Fibers:	5		Masses:				Fibers:				kV, Exposure, and Output Reproducibility
63			Specks:	4					Specks:
64			Masses:	4					Masses:					Measured Values
65		Date:					Inspector:				Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
66		Room Number					Survey ID:				Mo	Mo	28	50	28.93	502.8	6.134	0.123	12.20
1					Medical University of South Carolina			Mo	Mo	28	50	28.93	502.9	6.133	0.123	12.20
2					Inspection Results					Mammography System Compliance Inspection			Mo	Mo	28	50	28.92	502.2	6.129	0.123	12.20
3					Mo	Mo	28	50	28.96	502.4	6.127	0.123	12.20
4					Mean:	28.94	502.58	6.13	0.123	12.20
5					Std Dev:	0.02	0.33	0.00	0.00	0.00
6					Coeff of Var:	0.06%	0.07%	0.05%	0.05%	0.03%
7					Previous:				0.124	12.047
8					Criteria:	Coefficient of variation <= 2%
9					Minimum acceptable output is 7 mGy/s at 28 kVp for Mo target
10					Minimum acceptable output is 2 mGy/s at 28 kVp for W target
11					Exposure Linearity
12					Measured Values
13					Target:	Filter	kV	mAs	kVp	ms	Dose (mGy)	mGy/mAs	mGy/s
14					Mo	Mo	28	20	29.06	662.3	2.436	0.122	3.68
15					Mo	Mo	28	50	28.935	502.575	6.13075	0.123	12.20
16					Mo	Mo	28	100	28.94	1006	12.3	0.123	12.23
17					Mo	Mo	28	300	28.94	3370	36.95	0.123	10.96
18					Criteria:	Coefficient of linearity is <= 10%					Coeff of Linearity:	0.56%
19
20					1	Unit Can Maintain Output for 3 sec at 28 kVp
21
22					Half Value Layer
23					Target/Filter	Mo/Mo	Mo/Mo	Mo/Mo	Mo/Mo	Mo/Rh	Mo/Rh	Mo/Rh	Mo/Rh
24					kV	24	25	28	32	28	30	32	34
25					Al (mm)	Air Kerma (mGy)
26					0.0	3.621	4.210	6.133	9.110	3.871	4.843	5.905	7.005
27					0.2	2.207	2.620
28					0.3	1.774	2.122	3.287	5.131
29					0.4			2.737	4.337	1.997	2.561	3.176	3.813
30					0.5					1.719	2.213	2.749	3.309
31					0.0	3.624	4.230	6.129	9.160	3.871	4.873	5.934	7.043
32					HVL	0.29	0.30	0.34	0.37	0.42	0.44	0.45	0.46
33					0 mm	0.08%	0.48%	0.07%	0.55%	0.00%	0.62%	0.49%	0.54%
34					Min	0.24	0.25	0.28	0.32	0.28	0.30	0.32	0.34
35					Max	0.36	0.37	0.40	0.44	0.47	0.49	0.51	0.53
36					Pass	Pass	Pass	Pass	Pass	Pass	Pass	Pass
37					Criteria:	HVL is between the minimum and maximum limits
38
39					Image Quality Evaluation
40					Printer	Model:		SN:			QC Manual Version:
41
42					Previous				Previous
43					kV:				Fibers:	5
44					mA:				Specks:	4
45					Background Density:				Masses:	4
46					Inside disk:				Acceptable:	Pass
47					Outside disk:				Criteria:	Must be able to see 5 fibers, 4 speck groups, 4 masses
48					Difference:
49
50					Workstation
51					Workstation:				Workstation:
52					Left	Right			Left	Right
53					Fibers:				Fibers:
54					Specks:				Specks:
55					Masses:				Masses:
56
57					Acquisition Unit
58					Previous		Magnification		Previous
59					kV:				kV:
60					mA:				mA:
61					Exposure Index:				Exposure Index:
62					Fibers:				Fibers:
63					Specks:				Specks:
64					Masses:				Masses:
65		Date:					Inspector:
66		Room Number					Survey ID:			Signal to Noise
				Mean	SD
				kV:	28		In acrylic:	415.38	8.73
				mA:	71.7		Background:	515.14	9.94
				Target:
				Filter:				Current	Previous	Change	Acceptable
				SNR:	46.8	45.8	2.17%	Pass
				CNR:	10.0	9.9	1.38%	Pass
				Criteria:	SNR must be >= 40
				CNR does not change by more than 15%

## QC Test Summary-Lorad
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Lorad

Site Name					Report Date
Address					Survey Date	30-Dec-1899
Medical Physicist's Name					Signature
X-Ray Unit Manufacturer					Model
Date of Installation					Room ID

QC Manual Version #					(use any version applicable to model; contact mfr if questions)

Accessory Equipment			Manufacturer			Model		Location		QC Manual Version #
		Review Workstation*
		Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type			Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist)					Annual Survey

Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
					PASS/FAIL
1. Mammographic Unit Assembly Evaluation
2. Collimation Assessment
3. Artifact Evaluation
4. kVp Accuracy and Reproducibility
5. Beam Quality Assessment - HVL Measurement
6. Evaluation of System Resolution
7. Automatic Exposure Control (AEC) Function Performance (NA for systems without AEC)
8. Breast Entrance Exposure, AEC Reproducibility and Average Glandular Dose
		Average glandular dose for average breast is ≤3 mGy (300 mrad)					2	mGy
9. Radiation Output Rate
10. Phantom Image Quality Evaluation
		Phantom image scores:			Fibers		Specks		Masses
11. Signal-To-Noise Ratio and Contrast-To-Noise Ratio Measurements (values required for all tests)
		SNR (value)	46.7947686117
		CNR (value)	10.0362173038	(Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
		CNR should not vary by more than ±15% (NA for Equipment Evaluation)
12. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
13. DICOM Printer QC (Mammography Equipment Evaluations only)
14. Detector Flat Field Calibration (Mammography Equipment Evaluations only)
15. Compression Thickness Indicator (Mammography Equipment Evaluations only)
16. Compression (Mammography Equipment Evaluations only)

*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***

## Tech QC Eval-Lorad
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Lorad, continued)

Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.

					FREQUENCY		PASS/FAIL
1.	DICOM Printer Quality Control					Weekly
2.	Viewboxes and Viewing Conditions					Weekly
3.	Artifact Evaluation					Weekly
4.	Signal-To-Noise and Contrast-To-Noise Measurements					Weekly
5.	Phantom Image Quality Evaluation					Weekly
6.	Detector Flat-Field Calibration					Weekly
7.	Compression Thickness Indicator					Bi-weekly
8.	Visual Checklist					Monthly
9.	Repeat/Reject Analysis					Quarterly
10.	Compression					Semi-annually
11.	Diagnostic Review Workstation QC (NA if only hardcopy read)					See Lorad QC Manual

Medical Physicist's Recommendations for Quality Improvement

Important: / 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. / 2. Use any QC manual version that is applicable to the Lorad model surveyed. / 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) / 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) / 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day / than the survey date, note the date above. / 6. See the FDA-approved alternative standard for Lorad FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.

## Tables
Exposure to Dose Conversion Factors Mo/Mo for 4.2cm 50/50 Breast					Exposure to Dose Conversion Factors Mo/Rh for 4.2 cm 50/50 Breast					Exposure to Dose Conversion Factors Rh/Rh for 4.2cm 50/50 Breast
	kVp					kVp					kVp
HVL	23	24	25	26	27	28	29	30	31	HVL	23	24	25	26	27	28	29	30	31	HVL	23	24	25	26	27	28	29	30	31
0.23	116					0.23					0.23
0.24	121	124					0.24					0.24
0.25	126	129	131					0.25					0.25
0.26	130	133	135	138					0.26					0.26
0.27	135	138	140	142	143					0.27					0.27
0.28	140	142	144	146	147	149				0.28			149	151	154					0.28			150	155	159
0.29	144	146	148	150	151	153	154			0.29			154	156	158	159				0.29			155	160	164	168
0.3	149	151	153	155	156	157	158	159		0.3			158	160	162	162	163			0.3			160	164	168	172	176
0.31	154	156	157	159	160	161	162	163	164	0.31			163	164	166	166	167	167		0.31			165	168	172	174	180	182
0.32	158	160	162	163	164	166	167	168	168	0.32			167	169	171	171	171	172	172	0.32			169	173	177	181	184	186	188
0.33	163	165	166	168	169	170	171	173	173	0.33			171	173	175	176	176	176	176	0.33			174	178	181	185	188	190	192
0.34	168	170	171	172	173	174	175	176	177	0.34			176	178	179	179	180	180	180	0.34			179	183	186	190	193	195	196
0.35		174	175	176	177	178	179	180	181	0.35			180	181	183	183	184	185	185	0.35			184	187	190	194	197	199	201
0.36			179	181	182	183	184	185	185	0.36			185	186	187	187	188	188	189	0.36			189	192	195	198	201	204	205
0.37				185	186	187	188	189	190	0.37			189	190	191	191	192	193	193	0.37			193	196	199	202	205	207	209
0.38					190	191	192	193	194	0.38			193	194	196	196	197	197	197	0.38			198	201	204	207	209	211	213
0.39					196	197	198	198	0.39			198	199	200	200	201	201	202	0.39			203	206	208	211	214	216	217
0.4					201	202	203	0.4			202	203	204	204	205	205	206	0.4			208	211	213	216	218	220	221
0.41					206	207	0.41			206	207	208	208	209	209	210	0.41			213	215	217	220	222	224	225
0.42					211	0.42			211	211	212	212	213	213	214	0.42			218	220	222	224	226	228	229

Exposure to Dose Conversion Factors W/Rh for 4.2cm 50/50 Breast					Exposure to Dose Conversion Factors W/Ag for 4.2cm 50/50 Breast
	kVp					kVp
HVL	22	23	24	25	26	27	28	29	30	31	32	33	HVL	27	28	29	30	31	32	33	34	35	36	37	38	39
0.3	152	157	163	166	170	173	175	177	179	182	184	187	0.3	222	226	229	231	234	236	239	241	244	246	248	250	252
0.325	163	169	174	177	181	183	186	188	190	192	195	197	0.325	233	236	239	242	244	246	248	251	253	256	258	260	262
0.35	175	180	185	188	191	194	196	198	200	202	205	207	0.35	244	247	249	252	254	256	258	260	263	265	267	269	271
0.375	186	191	196	199	202	205	207	209	211	213	215	218	0.375	254	257	260	262	264	266	268	270	273	275	277	279	281
0.4	198	203	207	210	213	215	217	219	221	223	226	228	0.4	265	267	270	272	274	276	278	280	282	284	286	288	290
0.425	209	214	218	221	224	226	228	230	232	234	236	238	0.425	275	278	280	282	284	286	288	290	292	294	296	298	300
0.45	221	226	230	232	235	237	238	240	242	244	246	248	0.45	286	288	290	292	294	296	298	299	301	303	305	307	309
0.475	233	237	241	243	245	247	249	251	253	254	256	258	0.475	296	298	3090	302	304	305	307	309	311	313	315	317	318
0.5	244	248	252	254	256	258	260	261	263	265	267	269	0.5	306	308	310	312	313	315	317	319	320	322	324	326	328
0.525	256	260	263	265	267	269	270	272	273	275	277	279	0.525	316	318	320	322	323	325	326	328	330	332	333	335	337
0.55	267	271	274	276	278	279	281	282	284	285	287	289	0.55	326	328	330	331	333	334	336	338	339	341	343	344	346
0.575	279	282	285	287	288	290	291	292	294	296	297	299	0.575	336	338	339	341	342	344	345	347	349	350	352	354	355
0.6	290	293	296	297	299	300	301	303	304	306	308	310	0.6	346	348	349	350	352	353	355	356	358	359	361	363	364
0.625	301	304	306	308	309	310	312	313	315	316	318	320	0.625	356	357	358	360	361	362	364	365	367	368	370	372	373
0.65	312	314	317	318	320	321	322	323	325	326	328	330	0.65	365	367	368	369	370	372	373	375	376	378	379	381	382
0.675	322	325	327	328	330	331	333	333	335	336	338	340	0.675	374	376	377	378	379	381	382	383	385	386	388	390	391
0.7	333	335	337	339	340	341	342	343	345	346	348	350	0.7	384	385	386	387	388	390	391	392	394	395	397	398	400
0.725	342	345	347	348	349	351	352	353	354	356	358	360	0.725	393	394	395	396	397	399	400	401	403	404	406	407	408
0.75	352	355	357	358	359	360	361	363	364	366	368	369	0.75	402	403	404	405	406	407	409	410	411	413	414	416	417
0.775	361	365	367	368	369	370	371	372	374	375	377	379	0.775	410	411	412	413	415	416	417	418	420	421	423	424	425
0.8	369	374	376	377	378	379	380	382	383	385	387	389	0.8	419	420	421	422	423	424	425	427	428	429	431	432	434

Siemens Inspiration
	PMMA thickness (mm)	Equiv Breast thickness (mm)	HVL (mm Al)
			0.25	0.3	0.35	0.4	0.45	0.5	0.55	0.6	Slope	Intercept
g-factor	45	53	0.13	0.155	0.177	0.198	0.22	0.245	0.272	0.295	0.4681	0.0126
c-factor	45	53		1.109	1.105	1.102	1.099	1.096	1.091	1.088	-0.0693	1.1298
g-factor	50	60	0.112	0.135	0.154	0.172	0.192	0.214	0.236	0.261	0.4162	0.0076
c-factor	50	60		1.164	1.160	1.151	1.150	1.144	1.139	1.134	-0.0993	1.1935

	Mo/Mo	Mo/Rh	W/Rh
s-factor	1	1.017	1.042

kV Accuracy/Output		Calculate Radiation output (mGy/mAs and mGy/s) as a function of kV^2
Target	kV	kV^2	mGy/mAs	mGy/s		Target	kV	kV^2	mGy/mAs	mGy/s
Mo	24	576	0.072	6.92		Mo	28	784	0.077	7.66
Filter	25	625	0.084	8.37		Filter	30	900	0.097	9.64
Mo	26	676	0.096	9.59		Rh	32	1024	0.118	11.76
	28	784	0.123	12.20			34	1156	0.140	10.55
	30	900	0.151	15.08			36	1296	0.163	12.27
	32	1024	0.182	18.13			38	1444	0.188	13.19
	34	1156					Slope:	0.0001677841	0.007345926
		Slope:	0.0002447246	0.0247704731				Intercept:	-0.0541510178	2.7597965249
		Intercept:	-0.0688504035	-7.2130108093

Half Value Layer
Target	kV	HVL	Target	kV	HVL		Target	Filter	C
Mo	24	0.29	Mo	28	0.42		Mo	Mo	0.12
Filter	25	0.30	Filter	30	0.44		Mo	Rh	0.19
Mo	28	0.34	Rh	32	0.45		Rh	Rh	0.22
	32	0.37		34	0.46		Critiera:	Min HVL > kV/100 (no paddle)
	Slope:	0.0098616023		Slope:	0.0062006919			Min HVL > kV/100 + 0.03 (with paddle)
	Intercept:	0.0563895044		Intercept:	0.2487642822			Max HVL < kV/100 + C

DGN	DGN values (mrad/R) for 28 kV and HVL=0.34 mm Al
Mo/Mo	Mo/Rh	Rh/Rh	W/Rh	W/Ag	Siemens Inspiration
172.8	177.5	186.2	186	236	0.1233

## Corrected kV
GENERAL
	Enter Radcal kV reading in the light-blue cell under the appropriate anode-filter combination.
	Read the true corrected kV in the adjacent cell to the right.
	If that cell is light green, the corrected kV is accurate.
	If the cell is red, the entered kV is out of range and the correction should not be used.
	If the cell is light yellow, the entered value is in a reduced-accuracy range; try the other filter.

Tungsten-Rhodium (W / Rh)
	Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.03 kV.
	Reduced accuracy between true values of 24 to 27 kV. Max Radcal value is 40.47 kV.
	Radcal kV	True kV
	25.03	21.81

Tungsten-Silver (W / Ag)
	Minimum true value is 21.8 kV, corresponding to a Radcal value near 25.04 kV.
	Poor accuracy between true values of 25.35 to 31.85 kV. Max Radcal value is 37.1 kV.
	Radcal kV	True kV
	37.1	38.97

27.586	-8375.0727645926	975.9254356043	-37.913729682	0.4908658347	0	0	0
	-9984.6167916494	1436.5245457141	-82.5051021853	2.3655908176	-0.0338672434	0.0001936869	0

30.1	-540847.695500773	100186.233642731	-7418.4790179813	274.4766092958	-5.0743695436	0.0375005748	0
	-11057.7739361992	1297.2285673767	-56.9891889897	1.1115828564	-0.0081233997	0	0

## dropdowns
Pass/Fail
Pass
Fail

Pass/Fail/NA
Pass
Fail
NA

Phantom Scores
6.0
5.5
5.0
4.5
4.0
3.5
3.0
2.5
2.0
1.5
1.0
0.5