## Fluoro
1 | [11 empty] | Revision 1.3-20180816 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina | [5 empty] | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
4 | [22 empty] | All: | Page1,HVLPage,ExpChart,ImgQuality,Compg1,Compg2,OutputGraphFl,OutputGraphDig,LeedsTO10
5 | [5 empty] | Fluoroscopy System Compliance Inspection | [10 empty] | Fluoroscopy System Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10 |  | Page1,HVLPage,ExpChart,ImgQuality,Compg1,GraphAcryl,LeedsTO10
8 | [5 empty] | System Information | [5 empty] | Previous Date: | 2019-05-02 |  |  |  | System Information | [5 empty] | Date | 2019-05-02 | Change
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector | Eugene Mah |  | Eugene Mah
10 |  |  | Facility: | MUH |  |  |  | Site Number: | 400-432942 |  |  |  |  | Facility: | MUH |  |  | Site Number: | 400-432942 |  |  |  | Facility: | MUH |  | MUH
11 |  |  | Department: | Radiology |  |  |  | Authorized Use: | Fluoroscopy |  |  |  |  | Department: | Radiology |  |  | Authorized Use: | Fluoroscopy |  |  |  | Department: | Radiology |  | Radiology
12 |  |  | Area/Division: | OR |  |  |  | Date of Installation: | 2011-03-01 |  |  |  |  | Area/Division: | OR |  |  | Date of Installation: | 2011-03-01 |  |  |  | Area/Division: | OR |  | OR
13 |  |  | Survey ID: | 2470 |  |  |  | Number of X-ray Tubes: | 1 |  |  |  |  | Survey ID: | 2470 |  |  | Number of X-ray Tubes: | 1 |  |  |  | Survey ID: | 2470 |  | 2470
14 | [13 empty] | Room Number: | OR C-Arm 19 |  |  | Accession Number: | CARM19_2019 |  |  |  | Room Number: | OR C-Arm 19 |  | OR C-Arm 19
15 |  | X-Ray Generator | [20 empty] | Site Number: | 400-432942 |  | 400-432942
16 |  |  | Manufacturer: | GE |  |  |  | Manufacture Date: | 2011-03-01 |  |  |  | X-Ray Generator | [9 empty] | Authorized Use: | Fluoroscopy |  | Fluoroscopy
17 |  |  | Model: | 9900 Elite |  |  |  | Serial Number: | E2-2032 |  |  |  |  | Manufacturer: | GE |  |  | Manufacture Date: | 2011-03-01 |  |  |  | Date of Installation: | 2011-03-01 |  | 2011-03-01
18 |  |  | Max kVp: | 120 |  |  |  | Max mA: | 6 |  |  |  |  | Model: | 9900 Elite |  |  | Serial Number: | E2-2032 |  |  |  | Number of X-ray Tubes: | 1 |  | 1
19 | [7 empty] | Software Version: | [5 empty] | Max kVp: | 120 |  |  | Max mA: | 6 |  |  |  | Accesssion Number: | CARM19_2019 |  | CARM19_2019
20 |  | X-Ray Tube 1 | [15 empty] | Software Version: |  |  |  |  | X-Ray Generator
21 |  |  | Tube Designation/Use: | Fluoroscopy |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | X-Ray Tube 1 | [9 empty] | Manufacturer: | GE |  | GE
22 |  |  | Manufacture Date: | 2011-01-01 |  |  |  | Large: | 0.6 |  |  |  |  | Tube Designation/Use: | Fluoroscopy |  |  | Focal Spot Sizes (mm) |  |  |  |  | Model: | 9900 Elite |  | 9900 Elite
23 |  | Insert | [5 empty] | Small: | 0.3 |  |  |  |  | Manufacture Date: | 2011-01-01 |  |  | Large: | 0.6 |  |  |  | Max kVp: | 120 |  | 120
24 |  |  | Manufacturer: | Varian |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: | 0.3 |  |  |  | Manufacture Date: | 2011-03-01 |  | 2011-03-01
25 |  |  | Model: | Rad-99B | [9 empty] | Manufacturer: | Varian |  |  | Micro: |  |  |  |  | Serial Number: | E2-2032 |  | E2-2032
26 |  |  | Serial Number: | 91787-N1 |  |  |  | Filtration | [5 empty] | Model: | Rad-99B | [7 empty] | Max mA: | 6 |  | 6
27 |  | Housing | [5 empty] | Inherent: | [5 empty] | Serial Number: | 91787-N1 |  |  | Filtration |  |  |  |  | Software Version:
28 |  |  | Manufacturer: | Varian |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
29 |  |  | Model: | ORIIIB | [9 empty] | Manufacturer: | Varian |  |  | Added: |  |  |  |  | Tube Designation/Use: | Fluoroscopy |  | Fluoroscopy
30 |  |  | Serial Number: | H91787 | [9 empty] | Model: | ORIIIB | [7 empty] | Manufacture Date: | 2011-01-01 |  | 2011-01-01
31 | [13 empty] | Serial Number: | H91787 | [7 empty] | Manufacturer: | Varian |  | Varian
32 |  | X-Ray Tube 2 | [20 empty] | Model: | Rad-99B |  | Rad-99B
33 |  |  | Tube Designation/Use: |  |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | X-Ray Tube 2 | [9 empty] | Serial Number: | 91787-N1 |  | 91787-N1
34 |  |  | Manufacture Date: |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Manufacturer: | Varian |  | Varian
35 |  | Insert | [5 empty] | Small: | [5 empty] | Manufacture Date: |  |  |  | Large: |  |  |  |  | Model: | ORIIIB |  | ORIIIB
36 |  |  | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Serial Number: | H91787 |  | H91787
37 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Micro: |  |  |  |  | Large: | 0.6 |  | 0.6
38 |  |  | Serial Number: |  |  |  |  | Filtration | [5 empty] | Model: | [8 empty] | Small: | 0.3 |  | 0.3
39 | [7 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | Filtration |  |  |  |  | Micro:
40 |  |  | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | Inherent:
41 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Added: |  |  |  |  | Added:
42 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | X-Ray Tube 2
43 | [13 empty] | Serial Number: | [8 empty] | Tube Designation/Use:
44 | [22 empty] | Manufacture Date:
45 | [22 empty] | Manufacturer:
46 | [16 empty] | Inspection Results | [5 empty] | Model:
47 | [11 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Serial Number:
48 | [16 empty] | Labels, Notices, Postings |  | Tube Number: | 1 |  |  | Manufacturer:
49 | [12 empty] | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | **Input only one of the following |  |  | Model:
50 | [12 empty] | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | numbers or names |  |  | Serial Number:
51 | [12 empty] | Radiation warning label posted on the generator control panel | [6 empty] | 1 | 1 |  | Large:
52 | [12 empty] | Tube loading charts for each focal spot are posted | [6 empty] | 2 | 2 |  | Small:
53 | [12 empty] | Tube loading charts are correct for installed tube(s). | [6 empty] | 3 | 3 |  | Micro:
54 | [12 empty] | Operator manuals are available. | [6 empty] | AP | AP |  | Inherent:
55 | [12 empty] | Monthly radiation monitoring reports are posted. | [6 empty] | Front | Front |  | Added:
56 | [16 empty] | Fluoroscopy Systems |  |  | Frontal | Frontal |  | Tube Number: | 1 |  | 1
57 | [12 empty] | Interlock systems function to prevent fluoro while II is parked | [6 empty] | Lat | Lat |  | Tube Designation: | 1 |  | 1
58 | [12 empty] | The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [6 empty] | Lateral | Lateral |  |  | 2 |  | 2
59 | [12 empty] | Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels) | [10 empty] | 3 |  | 3
60 | [12 empty] | The useful beam is intercepted by the image assembly at any SID | [10 empty] | AP |  | AP
61 | [12 empty] | Fluoroscopy patient dose charts are posted | [10 empty] | Front |  | Front
62 | [11 empty] | 1 | Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [10 empty] | Frontal |  | Frontal
63 | [11 empty] | 1 | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [10 empty] | Lat |  | Lat
64 | [11 empty] | 3 | High contrast resolution within MUSC specifications | [10 empty] | Lateral |  | Lateral
65 | Date: | [8 empty] | Inspector: | Eugene Mah | 3 | Low contrast resolution within MUSC specifications
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470 | 1 | Transmission through II < 2 uGy/hr for each mGy/min at patient entrance
1 | [10 empty] | Medical University of South Carolina |  | Beam limiting device available (functioning) to restrict the x-ray field | [9 empty] | Receptor Size
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | The useful beam is centered on the image receptor | [9 empty] | Normal: | 30.48 |  | 30.48
3 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance |  |  | Collimator edges are visible at full open | [9 empty] | Mag 1: | 22.86 |  | 22.86
4 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD |  |  | Elapsed timer (5 min maximum) functional | [9 empty] | Mag 2: | 15.24 |  | 15.24
5 | DHEC RHB 10.2.1 |  | DHEC form SC-RAH-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  |  | Fluoroscopic kV and mA meter present and functional | [9 empty] | Mag 3:
6 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD | [6 empty] | Radiation Safety | [5 empty] | Mag 4:
7 |  |  | Tube loading charts for each focal spot are posted | [6 empty] | TBD |  |  | Continuous pressure on exposure switch required to make an exposure with selected tube | [9 empty] | Pt Dose Measured Fluoro
8 |  |  | Tube loading charts are correct for installed tube(s). | [6 empty] | TBD |  |  | Operator has full visibility of the exposure factors and patient | [9 empty] | Manual/Auto:
9 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | Lead aprons available | [9 empty] | Lucite/Al: | 1 |  | 1
10 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | Lead gloves available | [9 empty] | # dose rates: | 2 |  | 2
11 | [12 empty] | Properly designed and installed apron rack is present | [9 empty] | High level present: | 1 |  | 1
12 |  |  |  | Radiation Safety | [8 empty] | Documentation of annual protective apparel integrity inspection is available | [9 empty] | cm/in: | cm |  | cm
13 |  |  | Continuous pressure on exposure switch required to make an exposure with selected tube | [6 empty] | TBD |  |  | Patient restraint devices available | [9 empty] | Auto/mA: | Auto |  | Auto
14 |  |  | Operator has full visibility of the exposure factors and patient | [6 empty] | TBD |  |  | Gonadal shielding is present and functional | [9 empty] | Technique 1: | Normal |  | Normal
15 |  |  | Lead aprons available | [6 empty] | TBD |  |  | Radiation warning light at room entrance is functional with selected tube | [9 empty] | Technique 2: | Low Dose |  | Low Dose
16 |  |  | Lead gloves available | [6 empty] | TBD | [6 empty] | Physical Inspection | [5 empty] | Technique 3: | HLF |  | HLF
17 |  |  | Properly designed and installed apron rack is present | [6 empty] | TBD |  |  | System clean and free of debris | [9 empty] | Maximum Dose Rates Fluoro
18 |  |  | Documentation of annual protective apparel integrity inspection is available | [6 empty] | TBD |  |  | All cables free from excessive wear or restraints | [9 empty] | kV: | 120 | Change | 0
19 |  |  | Patient restraint devices available | [6 empty] | TBD |  |  | System free of missing or damaged components | [9 empty] | mA: | 6.43 | Change | 0
20 |  |  | Gonadal shielding is present and functional | [6 empty] | TBD |  |  | Tube support movements and locks are acceptable | [9 empty] | mGy/min: | 59.07 | Change | 0
21 |  |  | Radiation warning light at room entrance is functional with selected tube | [6 empty] | TBD |  |  | All other movements and locks functional | [9 empty] | kV: | 120 | Change | 0
22 | [16 empty] | Automatic Brightness Control / Display Monitor | [5 empty] | mA: | 3.54 | Change | 0
23 |  |  |  | Fluoroscopy Systems | [8 empty] | Automatic Brightness Control maintains image brightness for varying attenuation | [9 empty] | mGy/min: | 31.83 | Change | 0
24 | DHEC RHB 4.9.2.1.2 |  | Interlock systems function to prevent fluoro while II is parked | [6 empty] | TBD |  |  | Display monitor is free of artifacts, shadows, bright spots, and burn-in | [9 empty] | kV: | 120 | Change | 0
25 | DHEC RHB 4.9.1 |  | The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [6 empty] | TBD |  |  | Display monitor is free of distortion | [9 empty] | mA: | 12.67 | Change | 0
26 | DHEC RHB 4.9.8 |  | Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels) | [6 empty] | TBD |  |  | Display monitor is free of excessive persistence (after-image) | [9 empty] | mGy/min: | 117.4 | Change | 0
27 | DHEC RHB 4.9.2.1.2 |  | The useful beam is intercepted by the image assembly at any SID | [6 empty] | TBD |  |  | Display monitor is free of static or rolling noise bands | [9 empty] | Maximum Dose Rates Pulse/Dig
28 | DHEC RHB 4.9.4.3.6 |  | Fluoroscopy patient dose charts are posted | [6 empty] | TBD |  |  | The last image hold (LIH) display is free of jitter or abnormal appearance | [9 empty] | kV: | 120 | Change | 0
29 | DHEC RHB 4.9.4.2.1 |  | Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [6 empty] | YES | [12 empty] | mA: | 6.43 | Change | 0
30 | DHEC RHB 4.9.4.2.1.2 |  | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [6 empty] | YES | [12 empty] | mGy/min: | 12.11 | Change | 0
31 |  |  | High contrast resolution within MUSC specifications | [6 empty] | NA | [6 empty] | Inspection Results | [5 empty] | kV: | 120 | Change | 0
32 |  |  | Low contrast resolution within MUSC specifications | [6 empty] | NA | [12 empty] | mA: | 3.54 | Change | 0
33 | DHEC RHB 4.9.5.1 |  | Transmission through II < 2 uGy/hr for each mGy/min at patient entrance | [6 empty] | YES |  | Entrance Exposure Rate |  |  |  | Type 1 for manual, leave blank for auto |  |  | Auto | Auto | Auto/mA mode |  | mGy/min: | 6.76 | Change | 0
34 | DHEC RHB 4.9.2.2.1 |  | Beam limiting device available (functioning) to restrict the x-ray field | [6 empty] | TBD |  |  |  | 1 | 1 | Type 1 if using Lucite, leave blank if using Al | [6 empty] | kV: | 120 | Change | 0
35 |  |  | The useful beam is centered on the image receptor | [6 empty] | TBD |  | 2 | 2 | How many dose rate settings are present (exclusing High Level)? | [8 empty] | mA: | 35 | Change | 0
36 | DHEC RHB 4.9.2.2.1 |  | Collimator edges are visible at full open | [6 empty] | TBD |  | 1 | 1 | Is High Level provided? (1=Yes, 2=No) | [8 empty] | mGy/min: | 64.61 | Change | 0
37 | DHEC RHB 4.9.7 |  | Elapsed timer (5 min maximum) functional | [6 empty] | TBD |  | Blank | [10 empty] | HVL
38 | DHEC RHB 4.9.6 |  | Fluoroscopic kV and mA meter present and functional | [6 empty] | TBD |  | cm | cm | Unit for image receptor size | [8 empty] | HVL Measured kV: | 80 |  | 80
39 | DHEC RHB 4.3.5 |  | The Measured HVL meeds DHEC minimum requirements | [6 empty] | TBD |  | Patient Entrance Exposure Rate (Fluoroscopy)* | [10 empty] | HVL Measured: | 4.44 | Change | TBD
40 |  |  | Kerma area product display is accurate to within 40% | [6 empty] | NA |  | Image | Acrylic | Normal |  |  | Low Dose |  |  | HLF |  |  | Pt Dose Measured Dig Acq
41 |  |  | Air kerma display is accurate to within 40% | [6 empty] | NA |  | Mode | Attenuator | Normal |  |  | Low Dose |  |  | HLF |  |  | Technique 1: | Normal 8fps |  | Normal 8fps
42 |  |  | Air kerma rate display is accurate to within 40% | [6 empty] | NA |  | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | Technique 2: | Low Dose 8fps |  | Low Dose 8fps
43 |  |  | kVp accurate to within 5% of indicated. | [6 empty] | NA |  | Normal | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Technique 3: | HLC 8FPS |  | HLC 8FPS
44 |  |  | kV reproducibility coeff. Of var. < 5%. | [6 empty] | NA |  | 30.48 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Receptor Entr Exp Fluoro
45 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Normal Normal | 13.8333 | Change
46 |  |  |  | Automatic Brightness Control / Display Monitor | [7 empty] | Mag 1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Normal Mag 1 | 24.3889 | Change
47 |  |  | Automatic Brightness Control maintains image brightness for varying attenuation | [6 empty] | TBD |  | 22.86 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Normal Mag 2 | 54 | Change
48 |  |  | Display monitor is free of artifacts, shadows, bright spots, and burn-in | [6 empty] | TBD |  |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Normal Mag 3
49 |  |  | Display monitor is free of distortion | [6 empty] | TBD |  | Mag 2 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Normal Mag 4
50 |  |  | Display monitor is free of excessive persistence (after-image) | [6 empty] | TBD |  | 15.24 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Low Dose Normal | 5.6667 | Change
51 |  |  | Display monitor is free of static or rolling noise bands | [6 empty] | TBD |  |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Low Dose Mag 1 | 7.7222 | Change
52 |  |  | The last image hold (LIH) display is free of jitter or abnormal appearance | [6 empty] | TBD |  | Mag 3 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Low Dose Mag 2 | 19.3889 | Change
53 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Low Dose Mag 3
54 |  |  |  | Physical Inspection | [8 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Low Dose Mag 4
55 |  |  | System clean and free of debris | [6 empty] | TBD |  | Mag 4 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | HLF Normal | 22.1111 | Change
56 |  |  | All cables free from excessive wear or restraints | [6 empty] | TBD |  |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | HLF Mag 1 | 43.7222 | Change
57 |  |  | System free of missing or damaged components | [6 empty] | TBD |  |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | HLF Mag 2 | 117.8333 | Change
58 |  |  | Tube support movements and locks are acceptable | [6 empty] | TBD |  |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | HLF Mag 3
59 |  |  | All other movements and locks functional | [6 empty] | TBD |  |  | Last Year: | 120 | 6.43 | 59.07 | 120 | 3.54 | 31.83 | 120 | 12.67 | 117.4 | HLF Mag 4
60 | [12 empty] | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES | Receptor Entr Exp Dig Acq
61 | [12 empty] | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | Normal 8fps Normal | 11 | Change
62 | [12 empty] | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point) | [9 empty] | Normal 8fps Mag 1 | 18 | Change
63 | [12 empty] | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [9 empty] | Normal 8fps Mag 2 | 45 | Change
64 | [12 empty] | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point) | [9 empty] | Normal 8fps Mag 3
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [8 empty] | Normal 8fps Mag 4
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470 |  |  | Max Entrance Exposure Rate < = 176 mGy/min (high level) | [8 empty] | Low Dose 8fps Normal | 5 | Change
1 | [10 empty] | Medical University of South Carolina |  | Comments: | [9 empty] | Low Dose 8fps Mag 1 | 7 | Change
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | Do not exceed 390 chars |  | 0 | [7 empty] | Low Dose 8fps Mag 2 | 18 | Change
3 | Half Value Layer Measurement | [11 empty] | New: | [9 empty] | Low Dose 8fps Mag 3
4 | Set kVp: | 80 | [20 empty] | Low Dose 8fps Mag 4
5 | [11 empty] | Patient Entrance Exposure Rate – Digital Acquisition |  |  |  |  | Mode: | 2 | Enter 1 for Pulse fluoro, 2 for digital acquisition mode |  |  |  | HLC 8FPS Normal | 143 | Change
6 | Al (mm) | mGy/min | mGy/min |  | Measured HVL: | TBD |  |  | ACCEPTABLE RESULTS (mm Al) |  |  | Image | Acrylic | Normal 8fps |  |  | Low Dose 8fps |  |  | HLC 8FPS |  |  | HLC 8FPS Mag 1 | 263 | Change
7 | 0 |  |  |  | HVL Acceptable: | TBD |  |  | Min acceptable HVL at 80 kVp: | 2.6 |  | Mode | Attenuator | Normal 8fps |  |  | Low Dose 8fps |  |  | HLC 8FPS |  |  | HLC 8FPS Mag 2 | 709 | Change
8 | [8 empty] | Desired HVL at 80 kVp: | 3.1 |  | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | HLC 8FPS Mag 3
9 | [7 empty] | Criteria: | MUSC Standard |  |  | Normal | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | HLC 8FPS Mag 4
10 | [11 empty] | 30.48 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Fluoro Scatter/Transmission
11 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Transmission: | 0 |  | 0
12 | [11 empty] | Mag 1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (eyes): | 432 | Change
13 | [11 empty] | 22.86 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (waist): | 720 | Change
14 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (eyes):
15 | [11 empty] | Mag 2 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (waist):
16 | [11 empty] | 15.24 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dig Acq Scatter
17 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (eyes): | 618 | Change
18 | [11 empty] | Mag 3 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (waist): | 1062 | Change
19 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (eyes):
20 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (waist):
21 | [11 empty] | Mag 4 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Leeds Phantom
22 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | kVp: | 70 |  | 70
23 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Huttner Normal: | 1.6 | Change
24 | [12 empty] | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Huttner Mag 1: | 2.24 | Change
25 | [12 empty] | Last Year: | 120 | 6.43 | 12.11 | 120 | 3.54 | 6.76 | 120 | 35 | 64.61 | Huttner Mag 2: | 2.8 | Change
26 | [12 empty] | *Measured at: |  | Acceptable: | NA |  | Acceptable: | NA |  | Acceptable: | NA | Huttner Mag 3:
27 | [12 empty] | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | Huttner Mag 4:
28 | [12 empty] | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point) | [9 empty] | N3 Normal: | 0.0246 | Change
29 | [12 empty] | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [9 empty] | N3 Mag 1: | 0.0215 | Change
30 | [12 empty] | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point) | [9 empty] | N3 Mag 2 | 0.0172 | Change
31 | [12 empty] | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [8 empty] | N3 Mag 3
32 | [13 empty] | Max Entrance Exposure Rate < = 176 mGy/min (high level) | [8 empty] | N3 Mag 4:
33 | [12 empty] | Comments: | [9 empty] | Field Size: | 1 |  | 1
34 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | SID: | 100 |  | 100
35 | [12 empty] | New: | [9 empty] | Front: | -0.9 | Change
36 | [22 empty] | Back: | -0.7 | Change
37 | [11 empty] | Half Value Layer | [10 empty] | Left: | -0.8 | Change
38 | [11 empty] | 2 | Use DHEC standards? (Type 1 to base acceptance on DHEC criteria) | [9 empty] | Right: | -0.9 | Change
39 | [22 empty] | TO.10.A | 6 |  | 6
40 | [11 empty] | kV | mA | mm Al | mGy/min |  | Data Processing | [5 empty] | TO.10.B | 6 |  | 6
41 | [11 empty] | 80 |  | 0 |  |  | mm Al | Avg Exp | DATA(1/2) | Data Al | HVL |  | TO.10.C | 5 |  | 5
42 | [13 empty] | 0 |  |  | 0 |  |  |  | TBD |  | TO.10.D | 6 |  | 6
43 | [22 empty] | TO.10.E | 6 |  | 6
44 | [22 empty] | TO.10.F | 5 |  | 5
45 | [19 empty] | Data Exp | LN(EXP) |  | TO.10.G | 5 |  | 5
46 | [22 empty] | TO.10.H | 5 |  | 5
47 | [22 empty] | TO.10.J | 4 |  | 4
48 | [16 empty] | TBD | TBD |  |  |  |  | TO.10.K | 4 |  | 4
49 | [16 empty] | Acceptable (mm Al) |  |  | HVL Acceptable: | TBD |  | TO.10.L | 2 |  | 2
50 | [16 empty] | Low | 2.6 |  |  |  |  | TO.10.M | 1 |  | 1
51 | [16 empty] | Desired | 3.1 |  | Previous HVL @ 80 kV: | 4.44 |  | TO.10.A | 6 |  | 6
52 | [16 empty] | High | 3.5 |  |  |  |  | TO.10.B | 6 |  | 6
53 | [22 empty] | TO.10.C | 6 |  | 6
54 | [11 empty] | Image Receptor Entrance Exposure Rate – Fluoroscopy | [10 empty] | TO.10.D | 6 |  | 6
55 | [22 empty] | TO.10.E | 6 |  | 6
56 | [11 empty] | Dose | Receptor |  |  |  |  | Previous Year |  |  |  |  | TO.10.F | 6 |  | 6
57 | [11 empty] | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | nGy/frame |  |  |  |  | TO.10.G | 6 |  | 6
58 | [11 empty] | Normal | 30.48 |  |  |  |  | 13.8333 |  | Measure using pancake probe placed against |  |  | TO.10.H | 6 |  | 6
59 | [12 empty] | 22.86 |  |  |  |  | 24.3889 |  | image receptor input window |  |  | TO.10.J | 5 |  | 5
60 | [12 empty] | 15.24 |  |  |  |  | 54 |  | Use 1.5mm Cu filtration |  |  | TO.10.K | 5 |  | 5
61 | [22 empty] | TO.10.L | 3 |  | 3
62 | [22 empty] | TO.10.M | 2 |  | 2
63 | [11 empty] | Low Dose | 30.48 |  |  |  |  | 5.6667 |  |  |  |  | TO.10.A | 7 |  | 7
64 | [12 empty] | 22.86 |  |  |  |  | 7.7222 |  |  |  |  | TO.10.B | 7 |  | 7
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | 15.24 |  |  |  |  | 19.3889 |  |  |  |  | TO.10.C | 6 |  | 6
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470 | [11 empty] | TO.10.D | 7 |  | 7
1 | [10 empty] | Medical University of South Carolina | [11 empty] | TO.10.E | 7 |  | 7
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection | HLF | 30.48 |  |  |  |  | 22.1111 |  |  |  |  | TO.10.F | 6 |  | 6
3 | Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate | [11 empty] | 22.86 |  |  |  |  | 43.7222 |  |  |  |  | TO.10.G | 7 |  | 7
4 |  |  | Indicated KAP µGy·m² | Measured KAP µGy·m² | KAP Deviation | Ind AK @ Ref Pt (mGy) | Meas AK @ Ref Pt (mGy) | AK Deviation | Ind AKR | Meas AKR | AKR Deviation |  | 15.24 |  |  |  |  | 117.8333 |  |  |  |  | TO.10.H | 6 |  | 6
5 | kV | mA | [6 empty] | mGy/min | mGy/min | [12 empty] | TO.10.J | 6 |  | 6
6 | [22 empty] | TO.10.K | 6 |  | 6
7 | [22 empty] | TO.10.L | 5 |  | 5
8 | [12 empty] | Comments: | [9 empty] | TO.10.M | 3 |  | 3
9 | [14 empty] | 0 | [7 empty] | TO.10.A
10 | [12 empty] | New: | [9 empty] | TO.10.B
11 | [22 empty] | TO.10.C
12 | [11 empty] | Image Receptor Entrance Exposure Rate – Digital Acquisition | [10 empty] | TO.10.D
13 | [22 empty] | TO.10.E
14 |  |  |  | Acceptable: |  |  | Acceptable: |  |  | Acceptable: |  | Dose | Receptor |  |  |  |  | Previous Year |  |  |  |  | TO.10.F
15 |  | Criteria: | Measured KAP/AK/AKR is within 40% of the displayed values | [8 empty] | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | nGy/frame |  |  |  |  | TO.10.G
16 | [11 empty] | Normal 8fps | 30.48 |  |  |  |  | 11 |  | Measure using pancake probe placed against |  |  | TO.10.H
17 | [12 empty] | 22.86 |  |  |  |  | 18 |  | image receptor input window |  |  | TO.10.J
18 | kV Accuracy |  |  |  |  | kV Reproducibility | [6 empty] | 15.24 |  |  |  |  | 45 |  | Use 1.5mm Cu filtration |  |  | TO.10.K
19 | [22 empty] | TO.10.L
20 |  | Test Equipment Used: | Piranha CB2-17090320 | [19 empty] | TO.10.M
21 | [11 empty] | Low Dose 8fps | 30.48 |  |  |  |  | 5 |  |  |  |  | TO.10.A
22 |  | Measured |  |  |  |  | Measured | [5 empty] | 22.86 |  |  |  |  | 7 |  |  |  |  | TO.10.B
23 | kV Set | kV | % error |  |  | kV Set | kV | [5 empty] | 15.24 |  |  |  |  | 18 |  |  |  |  | TO.10.C
24 | 60 |  |  |  |  | 80 | [16 empty] | TO.10.D
25 | 80 | [21 empty] | TO.10.E
26 | 100 | [10 empty] | HLC 8FPS | 30.48 |  |  |  |  | 143 |  |  |  |  | TO.10.F
27 | 120 | [11 empty] | 22.86 |  |  |  |  | 263 |  |  |  |  | TO.10.G
28 |  | Acceptable: | [10 empty] | 15.24 |  |  |  |  | 709 |  |  |  |  | TO.10.H
29 |  | Criteria: | kVp accurate to within 5% of indicated. |  |  | Mean: | [16 empty] | TO.10.J
30 | [5 empty] | Std Dev: | [16 empty] | TO.10.K
31 | [5 empty] | Coeff of Var | [16 empty] | TO.10.L
32 | [5 empty] | Acceptable: | [6 empty] | Comments: | [9 empty] | TO.10.M
33 | [6 empty] | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%. |  |  |  |  | Do not exceed 390 chars |  | 0 | [7 empty] | Comments
34 | [12 empty] | New: | [9 empty] | Pt Entr Exp Fluoro:
35 | [22 empty] | Pt Entr Exp Dig Acq:
36 | [11 empty] | Scatter and Transmission | [10 empty] | Receptor Entr Exp Fluoro:
37 | [15 empty] | Pt Entrance | [6 empty] | Receptor Entr Exp Dig Acq:
38 | [13 empty] | kVp | mA | mGy/min | uGy/hr | Last Year |  | Acceptable |  |  | Scatter/Transmission:
39 | [12 empty] | Transmission* @ 10cm | 0 | 0 | 0 | 0 | 0 |  | YES |  |  | Leeds Phantom:
40 | [15 empty] | Unshielded | Shielded | Unshielded | Shielded | Unshielded | Shielded |  | TO.10:
41 | [15 empty] | uGy/min | uGy/min | uGy/hr | uGy/hr | uGy/hr | uGy/hr |  | Beam Limitation:
42 | [11 empty] | Scatter – Fluoro | Eye level | 0 | 0 |  |  |  |  | 432 |  |  | Comment Page 1
43 | [12 empty] | Waist level | 0 | 0 |  |  |  |  | 720
44 | [11 empty] | Scatter – Digital acq | Eye level | 0 | 0 |  |  |  |  | 618
45 | [12 empty] | Waist level | 0 | 0 |  |  |  |  | 1062
46 | [12 empty] | *Measured with 8” Lucite in beam
47 | [12 empty] | Largest field size mode, highest dose rate mode
48 | [12 empty] | Criteria: | Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure
49 | [12 empty] | Comments:
50 | [12 empty] | Do not exceed 390 chars |  | 0
51 | [12 empty] | New:
52
53 | [11 empty] | Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate
54 | [11 empty] | 30 | Source-PDC distance (cm) |  |  |  | Source-Isocenter
55 | [11 empty] | 30 | Source-Ref Pt distance (cm) |  |  |  | Table-Isocenter
56 | [11 empty] | 1 | CAK Reference Point correction factor |  |  |  | RefPt-Isocenter
57 | [11 empty] | 1 | Apply criteria
58 | [13 empty] | Start KAP | End KAP |  | Start AK | End AK |  | Ind AKR | Meas AKR
59 | [11 empty] | kV | mA | µGy·m² |  | KAP | mGy |  | CAK | mGy/min | mGy/min
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection | Radcal PDC |  |  |  | Calibration date:
3 | Patient Entrance Exposure Rate – Fluoroscopy | [10 empty] | kV | mA | mGy | mGy/Min | µGy·m² | µGy·m²/Min
4 | Image | Acrylic
5 | Mode | Attenuator | Normal |  |  | Low Dose |  |  | HLF
6 | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min
7 | 30.48 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
8 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10 | 22.86 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
12 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Indicated KAP µGy·m² | Measured KAP µGy·m² |  | Ind AK @ Ref Pt (mGy) | Meas AK @ Ref Pt (mGy) |  | Ind AKR | Meas AKR
13 | 15.24 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | KAP Deviation |  |  | CAK Deviation | mGy/min | mGy/min | Deviation |  |  | Comment Page 2
14 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
17 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Acceptable: |  |  | Acceptable: |  |  | Acceptable:
23
24 |  | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES |  | Criteria: | Measured KAP/AK/AKR is within 40% of the displayed values
25 |  | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | kV Accuracy and Reproducibility
26 |  | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27 |  | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [10 empty] | Test Equipment Used: | Piranha CB2-17090320
28 |  | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29 |  | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [9 empty] | Measured values
30 |  |  | Max Entrance Exposure Rate < = 176 mGy/min (high level) | [8 empty] | kV set | kV | Time (ms) | HVL | Tot Filt | kV % error
31 | Comments: | [10 empty] | 60
32 | [11 empty] | 80
33 | [11 empty] | 100
34 | [11 empty] | 120
35 | Image Receptor Entrance Exposure Rate – Fluoroscopy | [14 empty] | Acceptable:
36 | [12 empty] | Criteria: | kVp accurate to within 5% of indicated.
37 | Dose | Receptor
38 | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | [6 empty] | Measured values
39 | Normal | 30.48 | [5 empty] | Measure using pancake probe placed against |  |  |  | kV Set: | kV | Time (ms) | HVL | Tot Filt
40 |  | 22.86 | [5 empty] | image receptor input window |  |  |  | 80
41 |  | 15.24 | [5 empty] | Use 1.5mm Cu filtration
42
43
44 | Low Dose | 30.48
45 |  | 22.86 | [9 empty] | Mean:
46 |  | 15.24 | [9 empty] | Std Dev:
47 | [11 empty] | Coeff of Var
48 | [11 empty] | Acceptable:
49 | HLF | 30.48 | [10 empty] | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%.
50 |  | 22.86
51 |  | 15.24
52 | [11 empty] | Leeds Phantom
53
54 | Comments: | [10 empty] | Receptor Size | MS1 | MS3 | MS4 | Huttner | N3 | kVp: | 70 | 70
55 | [11 empty] | 30.48 cm | [6 empty] | MS1-4 |  | Huttner: Enter col.row
56 | [11 empty] | 22.86 cm | [6 empty] | 1 | Clearly visible | N3: Enter # of visible disks
57 | Scatter and Transmission – Fluoroscopy |  |  |  | Pt Entrance | [6 empty] | 15.24 cm | [6 empty] | 2 | Visible
58 |  |  | kVp | mA | mGy/min | uGy/hr | Unshielded | Shielded |  | Acceptable | [8 empty] | 3 | Barely visible
59 |  | Transmission* @ 10cm | 0 | 0 | 0 | 0 | uGy/hr | uGy/hr |  | YES | [8 empty] | 4 | Not visible
60 | Scatter – Fluoro | Eye level | 0 | 0 | [10 empty] | Acceptable: | NA | NA
61 |  | Waist level | 0 | 0 | [8 empty] | Comments:
62 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
63 | [12 empty] | New:
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah | TO.10 |  | Details Detected
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470 |  | Row | 30.48 cm | 22.86 cm | 15.24 cm
1 | [10 empty] | Medical University of South Carolina |  | A | 6 | 6 | 7
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | B | 6 | 6 | 7
3 | Patient Entrance Exposure Rate – Digital Acquisition | [11 empty] | C | 5 | 6 | 6
4 | Image | Acrylic | [10 empty] | D | 6 | 6 | 7
5 | Mode | Attenuator | Normal 8fps |  |  | Low Dose 8fps |  |  | HLC 8FPS |  |  |  | E | 6 | 6 | 7
6 | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min |  | F | 5 | 6 | 6
7 | 30.48 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | G | 5 | 6 | 7
8 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | H | 5 | 6 | 6
9 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | J | 4 | 5 | 6
10 | 22.86 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | K | 4 | 5 | 6
11 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | L | 2 | 3 | 5
12 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | M | 1 | 2 | 3
13 | 15.24 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Comments:
15 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Do not exceed 390 chars |  | 0
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | New:
17 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Field Size |  | 30.48 | 22.86 | 15.24 |  |  |  | 1
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Horizontal | [6 empty] | Type 1 for circular field, 2 for square field
20 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Vertical
21 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Diagonal
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Acceptable
23 | [13 empty] | Criteria: | Measured field size is at least 85% of nominal field size
24 | Comments: | [10 empty] | Useful Beam Limitation
25 | [11 empty] | 100 |  | SID (cm)
26
27 | [12 empty] | Measured | Difference | Sum | Last Year |  |  | T | B | L | R
28 | [11 empty] | Top |  |  |  | -0.9 |  | Visible
29 | [11 empty] | Bottom |  |  |  | -0.7 |  | X-ray
30 | [11 empty] | Left |  |  |  | -0.8 | Acceptable | SID | 100
31 | [11 empty] | Right |  |  |  | -0.9 |  | SOD
32 | Image Receptor Entrance Exposure Rate – Digital Acquisition | [10 empty] | Criteria: | Difference between x-ray field and visible image area is less than 2% SID |  |  |  |  | SID/SOD
33 | [12 empty] | Sum of length and width differences is less than 4% SID
34 | Dose | Receptor | [10 empty] | Comments:
35 | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | [6 empty] | Do not exceed 390 chars |  | 0
36 | Normal 8fps | 30.48 | [5 empty] | Measure using pancake probe placed against |  |  |  |  | New:
37 |  | 22.86 | [5 empty] | image receptor input window
38 |  | 15.24 | [5 empty] | Use 1.5mm Cu filtration
39 | [16 empty] | Comments Page 1
40 | [15 empty] | Previous Year Comments
41 | Low Dose 8fps | 30.48 | [10 empty] | Comments:
42 |  | 22.86 | [10 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
43 |  | 15.24 | [10 empty] | Additional Comments:
44 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
45 | [12 empty] | Additional Comments:
46 | HLC 8FPS | 30.48 | [10 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
47 |  | 22.86 | [10 empty] | Additional Comments:
48 |  | 15.24 | [10 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
49 | [12 empty] | Additional Comments:
50 | [12 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
51 | Comments: | [11 empty] | Additional Comments:
52 | [12 empty] | Do not exceed 390 chars |  | 0
53 | [12 empty] | Additional Comments:
54 | [12 empty] | Do not exceed 390 chars |  | 0
55 | Scatter and Transmission – Digital Acquisition | [11 empty] | Additional Comments:
56 |  |  |  |  | Unshielded | Shielded | [6 empty] | Do not exceed 390 chars |  | 0
57 |  |  | kVp | mA | uGy/hr | uGy/hr | [6 empty] | Additional Comments:
58 | Scatter – Digital acq | Eye level | 0 | 0 | [8 empty] | Do not exceed 390 chars |  | 0
59 |  | Waist level | 0 | 0 | [8 empty] | Additional Comments:
60 | [12 empty] | Do not exceed 390 chars |  | 0
61 | [12 empty] | Additional Comments:
62 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
63 | [12 empty] | Additional Comments:
64 | [12 empty] | Do not exceed 390 chars |  | 0
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Additional Comments:
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470 |  | Do not exceed 390 chars |  | 0
1 | [10 empty] | Medical University of South Carolina |  | Additional Comments:
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | Do not exceed 390 chars |  | 0
3 | Leeds Test Objects |  |  |  |  | High contrast resolution |  | Low contrast resolution |  |  |  |  | Additional Comments:
4 |  | Receptor Size | MS1 | MS3 | MS4 | lp/mm | Acceptable | N3 | Acceptable |  |  |  | Do not exceed 390 chars |  | 0
5 |  | 30.48 cm | [10 empty] | Additional Comments:
6 |  | 22.86 cm | [10 empty] | Do not exceed 390 chars |  | 0
7 |  | 15.24 cm | [10 empty] | Additional Comments:
8 | [12 empty] | Do not exceed 390 chars |  | 0
9 | [12 empty] | Additional Comments:
10 |  | Criteria: | Resolution: >1.0 lp/mm Normal mode, >1.25 lp/mm Mag 1, >1.4 lp/mm Mag 2, > 2 lp/mm Mag 3, > 2.5 lp/mm Mag 4 | [9 empty] | Do not exceed 390 chars |  | 0
11 |  |  | N3: Minimum contrast level < 0.04 in all II modes | [9 empty] | Additional Comments:
12 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
13
14
15 | [16 empty] | Comments Page 2
16 | TO.10 Contrast Resolution | [14 empty] | Previous Year Comments
17 |  | Contrast Detail |  |  |  |  | Threshold Index | [5 empty] | Comments:
18 | Row | 30.48 cm | 22.86 cm | 15.24 cm |  |  | 30.48 cm | 22.86 cm | 15.24 cm |  |  |  | Do not exceed 390 chars |  | 0
19 | A | 0.0322 | 0.0322 | 0.0215 |  |  | 3.157 | 3.157 | 4.7282 |  |  |  | Additional Comments:
20 | B | 0.0322 | 0.0322 | 0.0215 |  |  | 4.4358 | 4.4358 | 6.6434 |  |  |  | Do not exceed 390 chars |  | 0
21 | C | 0.045 | 0.0322 | 0.0322 |  |  | 4.4777 | 6.2576 | 6.2576 |  |  |  | Additional Comments:
22 | D | 0.045 | 0.045 | 0.0322 |  |  | 6.2688 | 6.2688 | 8.7607 |  |  |  | Do not exceed 390 chars |  | 0
23 | E | 0.045 | 0.045 | 0.0322 |  |  | 8.9554 | 8.9554 | 12.5153 |  |  |  | Additional Comments:
24 | F | 0.0674 | 0.045 | 0.045 |  |  | 8.3708 | 12.5375 | 12.5375 |  |  |  | Do not exceed 390 chars |  | 0
25 | G | 0.128 | 0.0876 | 0.0674 |  |  | 6.2968 | 9.2007 | 11.9582 |  |  |  | Additional Comments:
26 | H | 0.128 | 0.0876 | 0.0876 |  |  | 8.8155 | 12.881 | 12.881 |  |  |  | Do not exceed 390 chars |  | 0
27 | J | 0.167 | 0.128 | 0.0876 |  |  | 9.6525 | 12.5935 | 18.4015 |  |  |  | Additional Comments:
28 | K | 0.36 | 0.238 | 0.167 |  |  | 6.2688 | 9.4822 | 13.5135 |  |  |  | Do not exceed 390 chars |  | 0
29 | L | 0.726 | 0.573 | 0.238 |  |  | 4.4407 | 5.6264 | 13.546 |  |  |  | Additional Comments:
30 | M | 0.954 | 0.726 | 0.573 |  |  | 4.7311 | 6.217 | 7.877 |  |  |  | Do not exceed 390 chars |  | 0
31 | Comments: | [11 empty] | Additional Comments:
32 | [12 empty] | Do not exceed 390 chars |  | 0
33 | [12 empty] | Additional Comments:
34 | [12 empty] | Do not exceed 390 chars |  | 0
35 | Field Size | [11 empty] | Additional Comments:
36 |  |  | 30.48 cm | 22.86 cm | 15.24 cm | [7 empty] | Do not exceed 390 chars |  | 0
37 |  | Horizontal | [10 empty] | Additional Comments:
38 |  | Vertical | [10 empty] | Do not exceed 390 chars |  | 0
39 |  | Diagonal | [10 empty] | Additional Comments:
40 |  | Acceptable | [10 empty] | Do not exceed 390 chars |  | 0
41 |  | Criteria: | Measured field size is at least 85% of nominal field size | [9 empty] | Additional Comments:
42 | [12 empty] | Do not exceed 390 chars |  | 0
43 | Useful Beam Limitation | [11 empty] | Additional Comments:
44 | 100 | SID (cm) | [10 empty] | Do not exceed 390 chars |  | 0
45 |  |  | Measured | Difference | Sum | [7 empty] | Additional Comments:
46 |  | Top | [10 empty] | Do not exceed 390 chars |  | 0
47 |  | Bottom |  |  |  |  | Acceptable | [5 empty] | Additional Comments:
48 |  | Left | [10 empty] | Do not exceed 390 chars |  | 0
49 |  | Right | [10 empty] | Additional Comments:
50 |  | Criteria: | Difference between x-ray field and visible image area is less than 2% SID | [9 empty] | Do not exceed 390 chars |  | 0
51 |  |  | Sum of length and width differences is less than 4% SID | [9 empty] | Additional Comments:
52 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
53 | [12 empty] | Additional Comments:
54 | [12 empty] | Do not exceed 390 chars |  | 0
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Comments Page 1 |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Comments Page 2 |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Radiation Output – Fluoroscopy |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Radiation Output – Digital Acquisition |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Contrast Resolution |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | OR C-Arm 19 Tube 1 | [7 empty] | Survey ID: | 2470

## ExpCharts
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection
3 | Patient Entrance Exposure Rate – Fluoroscopy
4 | Image | Acrylic
5 | Mode | Attenuator | Normal | 0 | 0 | Low Dose | 0 | 0 | HLF | 0 | 0
6 | (cm) | (cm) | kVp | mA | R/min | kVp | mA | R/min | kVp | mA | R/min
7 | 30.48 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
8 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10 | 22.86 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
12 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
13 | 15.24 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
17 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
23
24 |  | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES
25 | (cm): Undertable fluoroscopy units: 1 cm above table top
26 | (cm): Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27 | (cm): C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point
28 | (cm): Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29 | (cm): Criteria:; kVp: Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)
30 | kVp: Max Entrance Exposure Rate < = 176 mGy/min (high level)
31 | (cm): Comments:
32
33
34
35 | (cm): Image Receptor Entrance Exposure Rate – Fluoroscopy
36
37 | (cm): Dose; (cm): Receptor
38 | (cm): Level; (cm): Size (cm); kVp: kVp; mA: mA; R/min: mR/min; kVp: nGy/frame
39 | (cm): Normal; (cm): 30.48; R/min: Measure using pancake probe placed against
40 | (cm): 0; (cm): 22.86; R/min: image receptor input window
41 | (cm): 0; (cm): 15.24; R/min: Use 1.5mm Cu filtration
42 | (cm): 0
43 | (cm): 0
44 | (cm): Low Dose; (cm): 30.48
45 | (cm): 0; (cm): 22.86
46 | (cm): 0; (cm): 15.24
47 | (cm): 0
48 | (cm): 0
49 | (cm): HLF; (cm): 30.48
50 | (cm): 0; (cm): 22.86
51 | (cm): 0; (cm): 15.24
52 | (cm): 0
53 | (cm): 0
54 | (cm): Comments:
55
56
57 | (cm): Scatter and Transmission – Fluoroscopy; R/min: Pt Entrance; kVp: Unshielded; mA: Shielded
58 | kVp: kVp; mA: mA; R/min: mGy/min; kVp: uGy/hr; mA: uGy/hr; kVp: Acceptable
59 | (cm): Transmission* @ 10cm; kVp: 0; mA: 0; R/min: 0; kVp: YES
60 | (cm): Scatter – Fluoro; (cm): Eye level; kVp: 0; mA: 0
61 | (cm): 0; (cm): Waist level; kVp: 0; mA: 0
62 | (cm): Comments:
63
64
65 | (cm): Date:; mA: Inspector:; R/min: Eugene Mah
66 | (cm): Room Number; (cm): OR C-Arm 19 Tube 1; mA: Survey ID:; R/min: 2470
1 | R/min: Medical University of South Carolina
2 | kVp: Inspection Results; R/min: Fluoroscopy System Compliance Inspection
3 | (cm): Patient Entrance Exposure Rate – Digital Acquisition
4 | (cm): Image; (cm): Acrylic
5 | Mode | Attenuator | Normal 8fps | 0 | 0 | Low Dose 8fps | 0 | 0 | HLC 8FPS | 0 | 0
6 | (cm) | (cm) | kVp | mA | R/min | kVp | mA | R/min | kVp | mA | R/min
7 | 30.48 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
8 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10 | 22.86 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
12 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
13 | 15.24 cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
17 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
23
24 | (cm): Comments:
25
26
27
28
29
30
31
32 | (cm): Image Receptor Entrance Exposure Rate – Digital Acquisition
33
34 | (cm): Dose; (cm): Receptor
35 | (cm): Level; (cm): Size (cm); kVp: kVp; mA: mA; R/min: mR/min; kVp: nGy/frame
36 | (cm): Normal 8fps; (cm): 30.48; R/min: Measure using pancake probe placed against
37 | (cm): 0; (cm): 22.86; R/min: image receptor input window
38 | (cm): 0; (cm): 15.24; R/min: Use 1.5mm Cu filtration
39 | (cm): 0
40 | (cm): 0
41 | (cm): Low Dose 8fps; (cm): 30.48
42 | (cm): 0; (cm): 22.86
43 | (cm): 0; (cm): 15.24
44 | (cm): 0
45 | (cm): 0
46 | (cm): HLC 8FPS; (cm): 30.48
47 | (cm): 0; (cm): 22.86
48 | (cm): 0; (cm): 15.24
49 | (cm): 0
50 | (cm): 0
51 | (cm): Comments:
52
53
54
55 | (cm): Scatter and Transmission – Digital Acquisition
56 | R/min: Unshielded; kVp: Shielded
57 | kVp: kVp; mA: mA; R/min: uGy/hr; kVp: uGy/hr
58 | (cm): Scatter – Digital acq; (cm): Eye level; kVp: 0; mA: 0
59 | (cm): 0; (cm): Waist level; kVp: 0; mA: 0
60
61
62 | (cm): Comments:
63
64
65 | (cm): Date:; mA: Inspector:; R/min: Eugene Mah
66 | (cm): Room Number; (cm): OR C-Arm 19 Tube 1; mA: Survey ID:; R/min: 2470

## Tables
Leeds MS | [13 empty] | Patient Entrance Exposure Rate (Fluoroscopy)* |  |  | Sorted by kV |  |  |  | Patient Entrance Exposure Rate – Digital Acquisition |  |  | Sorted by kV
1 | Clearly visible | [12 empty] | kV | Fluoro | High Level | kV | Fluoro | High Level |  | kV | Fluoro | High Level | kV | Fluoro | High Level
2 | Visible | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3 | Barely visible | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
4 | Not visible | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
[14 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
Huttner Grating |  |  | Leeds N3 |  | 65 | 70 | 75 | 80 | Nominal |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
1.1 | 0.5 |  | 1 | 0.167 | 0.188 | 0.167 | 0.149 | 0.135 | 0.16 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
1.2 | 0.56 |  | 2 | 0.148 | 0.167 | 0.148 | 0.132 | 0.12 | 0.145 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
1.3 | 0.63 |  | 3 | 0.128 | 0.145 | 0.128 | 0.114 | 0.104 | 0.123 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
2.1 | 0.71 |  | 4 | 0.109 | 0.124 | 0.109 | 0.097 | 0.088 | 0.108 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
2.2 | 0.8 |  | 5 | 0.0876 | 0.0993 | 0.0876 | 0.078 | 0.0705 | 0.086 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
2.3 | 0.9 |  | 6 | 0.0749 | 0.0847 | 0.0749 | 0.067 | 0.0609 | 0.076 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3.1 | 1 |  | 7 | 0.0674 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | 0.066 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3.2 | 1.12 |  | 8 | 0.0525 | 0.0593 | 0.0525 | 0.047 | 0.0427 | 0.055 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3.3 | 1.25 |  | 9 | 0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | 0.045 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
4.1 | 1.4 |  | 10 | 0.0371 | 0.04 | 0.0371 | 0.0347 | 0.0329 | 0.039 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
4.2 | 1.6 |  | 11 | 0.0322 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | 0.033 |  |  |  | Max | 0 | #DIV/0! | #N/A |  |  |  | Max | 0 | #DIV/0! | #N/A
4.3 | 1.8 |  | 12 | 0.0246 | 0.0269 | 0.0246 | 0.0233 | 0.022 | 0.027 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
5.1 | 2 |  | 13 | 0.0215 | 0.0232 | 0.0215 | 0.0201 | 0.019 | 0.023 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
5.2 | 2.24 |  | 14 | 0.0172 | 0.0186 | 0.0172 | 0.0161 | 0.0153 | 0.018 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
5.3 | 2.5 |  | 15 | 0.0155 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | 0.016 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
6.1 | 2.8 |  | 16 | 0.013 | 0.014 | 0.013 | 0.0122 | 0.0115 | 0.013 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
6.2 | 3.15 |  | 17 | 0.011 | 0.0119 | 0.011 | 0.0103 | 0.0097 | 0.011 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
6.3 | 3.55 |  | 18 | 0.0086 | 0.0093 | 0.0086 | 0.0081 | 0.0077 | 0.009 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
7.1 | 4 |  | 19 | 0.0066 | 0.0071 | 0.0066 | 0.0062 | 0.0059 | 0.007 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
7.2 | 4.5 | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
7.3 | 5 | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
[14 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
TO.10 kV Adj | 65 | 70 | 75 | 80 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.012 | 0.0119 | 0.011 | 0.0103 | 0.0097 |  |  |  |  | . |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.016 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.023 | 0.0232 | 0.0215 | 0.0201 | 0.019 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.032 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | [8 empty] | Max | 0 | #DIV/0! | #N/A |  |  |  | Max | 0 | #DIV/0! | #N/A
0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.066 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.086 | 0.0993 | 0.0876 | 0.078 | 0.0705 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.123 | 0.145 | 0.128 | 0.114 | 0.104 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.16 | 0.188 | 0.167 | 0.149 | 0.135 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.23 | 0.267 | 0.238 | 0.214 | 0.194 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.35 | 0.399 | 0.36 | 0.327 | 0.299 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.5 | 0.618 | 0.573 | 0.533 | 0.497 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.66 | 0.769 | 0.726 | 0.686 | 0.649 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
0.93 | 0.969 | 0.954 | 0.936 | 0.915 | [9 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
[14 empty] | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
TO.10 | Diameter (mm) | sqrt(Area) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |  |  | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
A | 11.1 | 9.8371 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012 |  |  | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
B | 7.9 | 7.0012 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012 |  |  | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
C | 5.6 | 4.9629 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012 |  |  | 0 | #N/A | #N/A |  |  |  |  | 0 | #N/A | #N/A
D | 4 | 3.5449 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 |  | Max | 0 | #N/A | #N/A |  |  |  | Max | 0 | #N/A | #N/A
E | 2.8 | 2.4814 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
F | 2 | 1.7725 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
G | 1.4 | 1.2407 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
H | 1 | 0.88623 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
J | 0.7 | 0.62036 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
K | 0.5 | 0.44311 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
L | 0.35 | 0.31018 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
M | 0.25 | 0.22156 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
TO.10 Values | 0 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9
A | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
B | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
C | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
D | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
E | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
F | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
G | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
H | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
J | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
K | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
L | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
M | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
Half Value Layer
| DHEC |  | 2: MUSC; 6: HVL Graph
| Min HVL |  | 2: HVL; 3: Min HVL; 4: Max HVL; 6: mm Al; 7: Avg Exp
kVp | (mm Al) |  | kVp | (mm Al) | (mm Al) | (mm Al) |  | TBD | TBD
30 | 0.3 |  | 60 | 2.2 | 1.8 | 2.4 |  | TBD | TBD
40 | 0.4 |  | 70 | 2.6 | 2.2 | 2.9 |  | TBD | TBD
50 | 0.5 |  | 80 | 3 | 2.6 | 3.5 |  | TBD | TBD
51 | 1.2 |  | 90 | 3.5 | 3.1 | 4 |  | TBD | TBD
60 | 1.3 |  | 100 | 3.9 | 3.5 | 4.5 |  | TBD | TBD
70 | 1.5 |  | 1: 110; 2: 4.3; 3: 3.9; 4: 5.1
71 | 2.1 |  | 1: 120; 2: 5.7; 3: 4.4; 4: 5.6
80 | 2.3 |  | 1: 130; 2: 5.2; 3: 4.8; 4: 6.1
90 | 2.5
100 | 2.7
110 | 3
120 | 3.2
130 | 3.5
140 | 3.8
150 | 4.1
Fluoro Patient Entrance Exposure Rate
Acrylic | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min
Attenuator | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail.
(cm)
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dig Acq Patient Entrance Exposure Rate
Acrylic | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min
Attenuator | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail.
(cm)
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KAP Conversion Table
µGy·m²
cGy·cm² | 1
dGy·cm² | 10
Gy·cm² | 100
mGy·cm² | 0.01
µGy·m² | 1
Note on A30: Nominal contrast values as a function of kVp
Note on A46: TO.10 nominal contrast values
Note on A60: TO.10 contrast values at the kVp specified in G55

## DataEntry
Patient Entrance Exposure Rate (Fluoroscopy)*
Image | Aluminum | Normal |  |  | Low Dose |  |  | HLF
Mode | Attenuator
(cm) | (cm) | kVp | mA | uGy/pls | kVp | mA | uGy/pls | kVp | mA | uGy/pls
Normal | 10
30.48 | 20
30
Mag 1 | 10
22.86 | 20
30
Mag 2 | 10
15.24 | 20
30
Mag 3 | 10
20
30
Mag 4 | 10
20
30
max
Patient Entrance Exposure Rate (Digital Acquisition)*
Image | Aluminum | Normal 8fps |  |  | Low Dose 8fps |  |  | HLC 8FPS
Mode | Attenuator
(cm) | (cm) | kVp | mA | uGy/pls | kVp | mA | uGy/pls | kVp | mA | uGy/pls
Normal | 10
30.48 | 20
30
Mag 1 | 10
22.86 | 20
30
Mag 2 | 10
15.24 | 20
30
Mag 3 | 10
20
30
Mag 4 | 10
20
30
max

## DataPage
Sheet | FLUORO
SurveyID | 2470
HVL | 80 | TBD
Field size | 30.48
22.86
15.24
Dose mode | Normal
Low Dose
HLF
SEE | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Max SEE | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dose mode | Normal 8fps
Low Dose 8fps
HLC 8FPS
SEE | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Max SEE | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
IEER | 30.48 | 0 | 0 | 0 | 0
| 22.86 | 0 | 0 | 0 | 0
| 15.24 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
| 30.48 | 0 | 0 | 0 | 0
| 22.86 | 0 | 0 | 0 | 0
| 15.24 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
| 30.48 | 0 | 0 | 0 | 0
| 22.86 | 0 | 0 | 0 | 0
| 15.24 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
IEER | 30.48 | 0 | 0 | 0 | 0
| 22.86 | 0 | 0 | 0 | 0
| 15.24 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
| 30.48 | 0 | 0 | 0 | 0
| 22.86 | 0 | 0 | 0 | 0
| 15.24 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
| 30.48 | 0 | 0 | 0 | 0
| 22.86 | 0 | 0 | 0 | 0
| 15.24 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
TO.N3 | 30.48
22.86
15.24
TO.10 | Field size | A | B | C | D | E | F | G | H | J | K | L | M
| 30.48 | 0.0322 | 0.0322 | 0.045 | 0.045 | 0.045 | 0.0674 | 0.128 | 0.128 | 0.167 | 0.36 | 0.726 | 0.954
| 22.86 | 0.0322 | 0.0322 | 0.0322 | 0.045 | 0.045 | 0.045 | 0.0876 | 0.0876 | 0.128 | 0.238 | 0.573 | 0.726
| 15.24 | 0.0215 | 0.0215 | 0.0322 | 0.0322 | 0.0322 | 0.045 | 0.0674 | 0.0876 | 0.0876 | 0.167 | 0.238 | 0.573
| 30.48 | 3.157 | 4.4358 | 4.4777 | 6.2688 | 8.9554 | 8.3708 | 6.2968 | 8.8155 | 9.6525 | 6.2688 | 4.4407 | 4.7311
| 22.86 | 3.157 | 4.4358 | 6.2576 | 6.2688 | 8.9554 | 12.5375 | 9.2007 | 12.881 | 12.5935 | 9.4822 | 5.6264 | 6.217
| 15.24 | 4.7282 | 6.6434 | 6.2576 | 8.7607 | 12.5153 | 12.5375 | 11.9582 | 12.881 | 18.4015 | 13.5135 | 13.546 | 7.877
Resolution | 30.48
22.86
15.24