## 4_1
Suggested occupancy factors (for use as a guide in planning shielding where other occupancy data are not available)

Location	Occupancy (T)
Administrative or clerical offices; laboratories, pharmacies and other work areas fully occupied by an individual; receptionist areas, attended waiting rooms, children's indoor play areas, adjacent x-ray rooms, film reading areas, nurse's stations, x-ray control rooms	1
Rooms used for patient examinations and treatments	0.5
Corridors, patient rooms, employee lounges, staff rest rooms	0.2
Corridor doors	0.125
Public toilets, unattended vending areas, storage rooms, outdoor areas with seating, unattended waiting rooms, patient holding areas	0.05
Outdoor areas with only transient pedestrian or vehicular traffic, unattended parking lots, vehicular drop off areas (unattended), attics, stairways, unattended elevators, janitor's closets	0.025

## 4_2
Operating potential (kVp) distribution of workload (mA-min) normalized per patient, from survey conducted by AAPM TG9 (Simpkin, 1996a)

	Radiography room			Rad/Fluoro
kVp	Rad room (all barriers)	Rad room (chest bucky)	Rad room (floor or other)	Fluoro tube (RF)	Rad tube (RF)	Chest room	Mammo room	Cardiac	Angio
25	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	9.25E-01	0.00E+00	0.00E+00
30	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	4.67E+00	0.00E+00	0.00E+00
35	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	1.10E+00	0.00E+00	0.00E+00
40	1.38E-04	0.00E+00	1.38E-04	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00	0.00E+00
45	7.10E-04	0.00E+00	7.10E-04	0.00E+00	5.78E-04	0.00E+00	0.00E+00	0.00E+00	0.00E+00
50	8.48E-03	6.78E-03	1.70E-03	0.00E+00	7.65E-04	0.00E+00	0.00E+00	3.40E-01	8.94E-02
55	1.09E-02	4.56E-04	1.04E-02	7.02E-02	7.26E-04	0.00E+00	0.00E+00	4.20E-01	3.98E-02
60	9.81E-02	8.96E-03	8.91E-02	1.13E-01	1.52E-02	0.00E+00	0.00E+00	1.96E+00	6.99E-01
65	1.04E-01	3.42E-02	7.00E-02	1.87E-01	2.52E-02	0.00E+00	0.00E+00	4.55E+00	1.50E+01
70	4.58E-01	7.25E-02	3.85E-01	1.45E-01	8.89E-02	2.02E-02	0.00E+00	6.03E+00	1.22E+01
75	5.01E-01	9.53E-02	4.05E-01	1.94E-01	2.24E-01	2.36E-03	0.00E+00	8.02E+00	1.53E+01
80	5.60E-01	1.40E-01	4.20E-01	1.72E+00	4.28E-01	0.00E+00	0.00E+00	2.54E+01	1.10E+01
85	3.15E-01	6.62E-02	2.49E-01	2.19E+00	2.18E-01	7.83E-04	0.00E+00	4.03E+01	4.09E+00
90	1.76E-01	1.41E-02	1.62E-01	1.46E+00	5.33E-02	0.00E+00	0.00E+00	2.10E+01	3.43E+00
95	2.18E-02	3.51E-03	1.82E-02	1.15E+00	4.89E-02	0.00E+00	0.00E+00	1.06E+01	6.73E-01
100	1.55E-02	8.84E-04	1.46E-02	1.12E+00	5.87E-02	3.01E-02	0.00E+00	7.40E+00	1.53E+00
105	3.48E-03	1.97E-03	1.51E-03	9.64E-01	1.05E-02	0.00E+00	0.00E+00	7.02E+00	9.27E-02
110	1.05E-02	9.91E-03	5.51E-04	7.47E-01	6.46E-02	2.14E-02	0.00E+00	6.59E+00	3.05E-02
115	4.10E-02	3.74E-02	3.69E-03	1.44E+00	2.90E-02	9.36E-02	0.00E+00	1.38E+01	0.00E+00
120	6.99E-02	5.12E-02	1.87E-02	9.37E-01	1.04E-01	4.74E-02	0.00E+00	3.35E+00	0.00E+00
125	4.84E-02	4.81E-02	3.47E-04	1.38E-01	8.13E-02	0.00E+00	0.00E+00	2.75E+00	0.00E+00
130	1.84E-03	1.71E-03	1.25E-04	1.53E-01	4.46E-02	0.00E+00	0.00E+00	3.10E-02	0.00E+00
135	7.73E-03	7.73E-03	0.00E+00	1.46E-01	9.47E-03	0.00E+00	0.00E+00	0.00E+00	0.00E+00
140	0.00E+00	0.00E+00	0.00E+00	1.92E-02	4.26E-03	0.00E+00	0.00E+00	0.00E+00	0.00E+00
Tot Workload	2.45	0.60	1.85	12.89	1.51	0.22	6.70	159.56	64.17
Pt/week	110			18	23	210	47	19	21

## 4_3
Estimated total workloads in various medical x-ray imaging installations in clinics and hospitals.
The total workload values are for general guidance and are to be used only if the actual workloads are not available.
		Typical # of patients (N) per 40h week		W_tot (mA-min/week)
Room Type	W_norm (mA-min/pt)	Average	Busy	Average	Busy
Rad room (Chest bucky)	0.60	120	160	72	96
Rad room (floor or other)	1.85	120	160	222	296
Chest room	0.22	200	400	43	86
Fluoro tube (RF)	12.89	20	30	258	387
Rad tube (RF)	1.51	25	40	38	60
Mammo	6.70	80	160	536	1071
Cardiac	159.56	20	30	3191	4787
Angio	64.17	20	30	1283	1925

## 4_4
Primary beam use factors (U) for a general radiographic room determined from the survey of clinical sites (Simpkin, 1996a)

Barrier	Use factor (U)
Floor	0.89
Cross-table wall	0.09
Wall #3	0.02
Chest image reecptor	1

## 4_7
Unshielded leakage, scattered and total secondary air kermas (in mGy/pt) for the indicated workload distribution at d_s = d_L = 1m
The workload distributions and total workloads per patient (Wnorm) for the indicated clinical sites are the average per patient surveyed by AAPM TG9 (Simpkin, 1996a)
The primary field size F (in cm^2) is known at primary distance d_F. Side scattered radiation is calculated for 90 degree scatter.
Forward and backscattered radiations are calculated for 135 degree scatter.
Leakage radiation technique factors are 150 kVp at 3.3 mA to achieve 0.876 mGy/h for all tubes except mammography, which assumes leakage radiation technique factors of 50 kVp at 5 mA
				Unshielded Air Kerma (mGy/pt) at 1m
Workload Distribution	Wnorm (mAmin/pt)	F (cm^2)	d_F (m)	Leakage	Side scatter	Leakage and side scatter	Forward/Backscatter	Lekage and forward/back scatter)
Rad room (all barriers)	2.45	1000	1	0.00053	0.034	0.03453	0.048	0.04853
Rad room (chest bucky)	0.60	1535	1.83	0.00039	0.0049	0.00529	0.0069	0.00729
Rad room (floor or other)	1.85	1000	1	0.00014	0.023	0.02314	0.033	0.03314
Fluoro tube (RF)	12.89	730	0.8	0.012	0.31	0.322	0.44	0.452
Rad tube (RF)	1.51	1000	1	0.00094	0.028	0.02894	0.039	0.03994
Chest room	0.22	1535	2	0.00038	0.0023	0.00268	0.0032	0.00358
Mammo room	6.70	720	0.58	0.000011	0.011	0.011011	0.049	0.049011
Cardiac	159.56	730	0.9	0.088	2.6	2.688	3.7	3.788
Angio	64.17	730	0.9	0.0034	0.66	0.6634	0.95	0.9534

## 5_2
Currently suggested defalut DLP values per procedure.
For use as a guide in planning shielding in cases where facility-specific DLP values are not available

Procedure	CTDIvol (mGy)	Scan length (cm)	DLP (mGy-cm)
Head	60	20	1200
Chest	15	35	525
Abdomen	25	25	625
Pelvis	25	20	500
Body avg			550

## A_1
Fits of transmission for broad primary x-ray beams (for lead, concrete, gypsum wall board, steel, plate glass, and wood) to Equation A.2 (thickness x is input in mm)

	Pb (mm)					Concrete (mm)					Gypsum					Steel					Plate Glass					Wood
kVp	α	β	γ	HVL	TVL	α	β	γ	HVL	TVL	α	β	γ	HVL	TVL	α	β	γ	HVL	TVL	α	β	γ	HVL	TVL	α	β	γ	HVL	TVL
25	49.52	194	0.3037	0.014	0.046	0.3904	1.645	0.2757	1.775	5.898	0.1576	0.7175	0.3048	4.398	14.610	9.364	41.25	0.3202	0.074	0.246	0.3804	1.543	0.2869	1.822	6.053	0.0223	0.0434	0.1937	31.083	103.255
30	38.8	178	0.3473	0.018	0.059	0.3173	1.698	0.3593	2.185	7.257	0.1208	0.7043	0.3613	5.738	19.061	7.406	41.93	0.3959	0.094	0.311	0.3061	1.599	0.3693	2.264	7.522	0.02166	0.03966	0.2843	32.001	106.306
35	29.55	164.7	0.3948	0.023	0.078	0.2528	1.807	0.4648	2.742	9.108	0.08878	0.6988	0.4245	7.807	25.936	5.716	43.41	0.4857	0.121	0.403	0.2396	1.694	0.4683	2.893	9.610	0.01901	0.03873	0.3732	36.462	121.125
40					0.1297	0.178	0.2189	5.344	17.753
45					0.1095	0.1741	0.2269	6.330	21.028
50	8.801	27.28	0.2957	0.079	0.262	0.0932	0.1712	0.2324	7.437	24.706	0.03883	0.0873	0.5105	17.851	59.299	1.817	4.84	0.4021	0.381	1.267	0.09721	0.1799	0.4912	7.130	23.687	0.01076	0.001862	1.17	64.419	213.995
55	7.839	25.92	0.3499	0.088	0.294	0.07422	0.1697	0.2454	9.339	31.024	0.03419	0.08315	0.5606	20.273	67.347	1.493	4.515	0.4293	0.464	1.542	0.08552	0.1661	0.5112	8.105	26.925	0.01012	0.001404	1.269	68.493	227.528
60	6.951	24.89	0.4198	0.100	0.331	0.06251	0.1692	0.2733	11.089	36.835	0.02985	0.07961	0.6169	23.221	77.139	1.183	4.219	0.4571	0.586	1.946	0.07452	0.1539	0.5304	9.301	30.899	0.009512	0.0009672	1.333	72.871	242.072
65	6.13	24.09	0.5019	0.113	0.376	0.05528	0.1696	0.3217	12.539	41.653	0.02609	0.07597	0.6756	26.568	88.255	0.9172	3.982	0.4922	0.756	2.510	0.06514	0.1443	0.5582	10.641	35.348	0.00899	0.000647	1.353	77.102	256.127
70	5.369	23.49	0.5881	0.129	0.429	0.05087	0.1696	0.3847	13.626	45.264	0.02302	0.07163	0.7299	30.111	100.025	0.7149	3.798	0.5378	0.970	3.221	0.05791	0.1357	0.5967	11.969	39.761	0.00855	0.000539	1.194	81.070	269.308
75	4.666	22.69	0.6618	0.149	0.493	0.04797	0.1663	0.4492	14.450	48.001	0.02066	0.06649	0.775	33.550	111.451	0.5793	3.629	0.5908	1.197	3.975	0.05291	0.128	0.6478	13.100	43.519	0.005203	0.0006421	1.062	133.221	442.550
80	4.04	21.69	0.7187	0.172	0.570	0.04583	0.1549	0.4926	15.124	50.242	0.01886	0.06093	0.8103	36.752	122.088	0.4921	3.428	0.6427	1.409	4.679	0.04955	0.1208	0.7097	13.989	46.470	0.007903	0.000864	0.9703	87.707	291.356
85	3.504	20.37	0.755	0.198	0.657	0.04398	0.1348	0.4943	15.761	52.355	0.01746	0.05558	0.8392	39.699	131.878	0.4355	3.178	0.6861	1.592	5.287	0.04721	0.114	0.7786	14.682	48.773	0.007686	0.001056	1.015	90.183	299.582
90	3.067	18.83	0.7726	0.226	0.751	0.04228	0.1137	0.469	16.394	54.460	0.01663	0.05039	0.8585	41.681	138.460	0.3971	2.913	0.7204	1.746	5.799	0.0455	0.1077	0.8522	15.234	50.606	0.007511	0.001159	1.081	92.284	306.562
95	2.731	17.07	0.7714	0.254	0.843	0.04068	0.09705	0.4406	17.039	56.602	0.01543	0.04571	0.8763	44.922	149.228	0.3681	2.654	0.7461	1.883	6.255	0.0441	0.1013	0.9222	15.718	52.213	0.007345	0.001133	1.116	94.370	313.490
100	2.5	15.28	0.7557	0.277	0.921	0.03925	0.08567	0.4273	17.660	58.665	0.01466	0.04171	0.8939	47.282	157.066	0.3415	2.42	0.7645	2.030	6.743	0.04278	0.09466	0.9791	16.203	53.824	0.00723	0.0009343	1.309	95.871	318.476
105	2.364	13.41	0.7239	0.293	0.974	0.03808	0.07862	0.4394	18.202	60.467	0.01397	0.03815	0.908	49.617	164.824	0.3135	2.227	0.7788	2.211	7.345	0.04143	0.08751	1.014	16.731	55.578	0.00705	0.0006199	1.365	98.319	326.608
110	2.296	11.7	0.6827	0.302	1.003	0.03715	0.07436	0.4752	18.658	61.981	0.01336	0.03521	0.9244	51.882	172.349	0.2849	2.061	0.7897	2.433	8.082	0.04008	0.08047	1.03	17.294	57.450	0.006921	0.0001976	3.309	100.151	332.695
115	2.265	10.21	0.6363	0.306	1.017	0.03636	0.07201	0.5319	19.063	63.327	0.01283	0.06271	0.9423	54.026	179.469	0.2579	1.922	0.8008	2.688	8.928	0.03878	0.07394	1.033	17.874	59.376	0.006864	-0.0003908	0.6469	100.983	335.458
120	2.246	8.95	0.5873	0.309	1.025	0.03566	0.07109	0.6073	19.438	64.571	0.01235	0.03047	0.9566	56.125	186.444	0.2336	1.797	0.8116	2.967	9.857	0.03758	0.06808	1.031	18.445	61.272	0.006726	-0.0008308	1.006	103.055	342.341
125	2.219	7.923	0.5386	0.312	1.038	0.03502	0.07113	0.6974	19.793	65.751	0.01192	0.02863	0.9684	58.150	193.170	0.213	1.677	0.8217	3.254	10.810	0.03652	0.06304	1.031	18.980	63.050	0.006584	-0.001214	1.192	105.278	349.724
130	2.17	7.094	0.4909	0.319	1.061	0.03445	0.0716	0.7969	20.120	66.838	0.01155	0.02702	0.9802	60.013	199.358	0.1969	1.557	0.8309	3.520	11.694	0.03561	0.05874	1.037	19.465	64.661	0.006472	-0.001539	1.285	107.099	355.776
135	2.102	6.45	0.4469	0.330	1.095	0.03394	0.07263	0.9099	20.423	67.843	0.01122	0.02561	0.9901	61.778	205.221	0.1838	1.44	0.8391	3.771	12.528	0.03481	0.05519	1.049	19.912	66.147	0.006306	-0.001731	1.465	109.919	365.142
140	2.009	5.916	0.4018	0.345	1.146	0.03345	0.07476	1.047	20.722	68.837	0.01088	0.02436	0.9964	63.708	211.635	0.1724	1.328	0.8458	4.021	13.356	0.03407	0.05145	1.057	20.345	67.584	0.006191	-0.001849	1.53	111.960	371.925
145	1.895	5.498	0.358	0.366	1.215	0.03296	0.07875	1.224	21.030	69.860	0.01056	0.02313	0.9987	65.639	218.048	0.1616	1.225	0.8519	4.289	14.249	0.03336	0.04795	1.063	20.778	69.022	0.006115	-0.001869	1.498	113.352	376.547
150	1.757	5.177	0.3156	0.395	1.311	0.03243	0.08599	1.467	21.374	71.002	0.0103	0.02198	1.013	67.296	223.552	0.1501	1.132	0.8566	4.618	15.340	0.03266	0.04491	1.073	21.223	70.502	0.00602	-0.001752	1.483	115.141	382.489

Air kerma at 1m per unit workload (mGy/mA-min)
kVp	Ko**
25	1.484
30	2.466
35	3.449
40	0.719
45	0.872
50	1.066
55	1.297
60	1.563
65	1.863
70	2.194
75	2.553
80	2.938
85	3.347
90	3.777
95	4.226
100	4.692
105	5.172
110	5.664
115	6.166
120	6.675
125	7.189
130	7.705
135	8.222
140	8.736
145	9.246
150	9.750

## W_Rh_Al_Ag
Fits of transmission for broad primary x-ray beams (for lead, concrete, gypsum wall board, steel, plate glass, and wood) to Equation A.2 (thickness x is input in mm) for W/Ag x-ray beams (mammography)
	Pb (mm)			Concrete (mm)			Gypsum			Steel			Plate Glass			Wood
kVp	α	β	γ	α	β	γ	α	β	γ	α	β	γ	α	β	γ	α	β	γ
25	6.700E+01	2.126E+01	2.288E-01	4.288E-01	5.526E-01	4.359E-01	2.174E-01	2.774E-01	3.707E-01	1.115E+01	7.133E+00	1.884E-01	4.812E-01	6.010E-01	4.720E-01	2.201E-02	1.011E-02	3.083E-01
30	4.695E+01	3.373E+01	1.923E-01	2.891E-01	4.375E-01	2.742E-01	1.384E-01	2.394E-01	2.387E-01	8.295E+00	8.655E+00	2.527E-01	3.197E-01	5.001E-01	2.888E-01	1.756E-02	8.786E-03	1.797E-01
35	2.841E+01	6.294E+01	3.201E-01	1.541E-01	5.974E-01	3.090E-01	6.557E-02	3.031E-01	2.593E-01	5.402E+00	1.681E+01	5.086E-01	1.514E-01	6.343E-01	2.725E-01	1.186E-02	1.312E-02	1.603E-01
39	1.544E+01	6.738E+01	2.861E-01	1.292E-01	6.718E-01	4.475E-01	6.427E-02	3.709E-01	4.628E-01	2.956E+00	1.431E+01	3.706E-01	1.443E-01	7.505E-01	4.603E-01	1.071E-02	1.373E-02	2.387E-01

Source:	Li X, Zhang D, Liu B, “Workload and transmission data for the installation of a digital breast tomosynthesis system”, MedPhys 40(6) 063901-1

Fits of transmission for broad primary x-ray beams (for lead, concrete, gypsum wall board, steel, plate glass, and wood) to Equation A.2 (thickness x is input in mm) for W/Al x-ray beams (mammography)
	Pb (mm)			Concrete (mm)			Gypsum			Steel			Plate Glass			Wood
kVp	α	β	γ	α	β	γ	α	β	γ	α	β	γ	α	β	γ	α	β	γ
25	6.306E+01	2.709E+01	1.490E-01	4.283E-01	6.450E-01	4.340E-01	2.180E-01	3.308E-01	3.798E-01	1.144E+01	8.967E+00	2.344E-01	4.733E-01	7.575E-01	4.454E-01	2.190E-02	1.372E-02	3.181E-01
30	2.572E+01	4.912E+01	1.217E-01	2.545E-01	4.988E-01	3.782E-01	1.391E-01	2.945E-01	4.519E-01	5.831E+00	9.782E+00	2.166E-01	2.919E-01	6.105E-01	4.443E-01	1.723E-02	1.196E-02	4.733E-01
35	2.220E+01	4.234E+01	2.152E-01	1.826E-01	4.967E-01	4.828E-01	9.200E-02	2.694E-01	4.893E-01	4.344E+00	9.436E+00	3.310E-01	1.975E-01	5.284E-01	4.656E-01	1.417E-02	9.502E-03	4.630E-01
40	1.792E+01	4.158E+01	3.201E-01	1.379E-01	4.620E-01	5.482E-01	6.727E-02	2.510E-01	5.468E-01	2.957E+00	8.541E+00	3.569E-01	1.492E-01	4.492E-01	5.028E-01	1.244E-02	8.440E-03	5.356E-01
45	1.297E+01	3.985E+01	3.514E-01	1.093E-01	4.065E-01	5.917E-01	5.115E-02	2.210E-01	5.830E-01	2.297E+00	8.032E+00	4.421E-01	1.173E-01	4.391E-01	5.791E-01	1.142E-02	8.285E-03	7.651E-01
49	1.150E+01	3.707E+01	4.101E-01	9.267E-02	4.031E-01	6.295E-01	4.267E-02	2.252E-01	6.361E-01	1.752E+00	8.332E+00	4.704E-01	1.007E-01	4.478E-01	6.430E-01	1.073E-02	8.158E-03	8.549E-01

Source:	Li X, Zhang D, Liu B, “Transmission of broad W/Rh and W/Al (target/filter) x-ray beams operated at 25-29 kVp through common shielding materials”, MedPhys 39(7) 4132

Fits of transmission for broad primary x-ray beams (for lead, concrete, gypsum wall board, steel, plate glass, and wood) to Equation A.2 (thickness x is input in mm) for W/Rh x-ray beams (mammography)
	Pb (mm)			Concrete (mm)			Gypsum			Steel			Plate Glass			Wood
kVp	α	β	γ	α	β	γ	α	β	γ	α	β	γ	α	β	γ	α	β	γ
25	6.794E+01	2.096E+01	8.503E-02	4.906E-01	4.668E-01	3.878E-01	2.432E-01	2.539E-01	3.153E-01	1.371E+01	4.772E+00	1.771E-01	5.308E-01	5.201E-01	3.626E-01	2.429E-02	9.839E-03	3.314E-01
30	4.972E+01	4.585E+01	2.231E-01	2.721E-01	6.599E-01	2.827E-01	1.538E-01	3.565E-01	3.216E-01	7.870E+00	1.280E+01	2.533E-01	3.101E-01	7.473E-01	3.029E-01	1.855E-02	1.307E-02	2.267E-01
35	2.561E+01	8.242E+01	3.307E-01	1.580E-01	8.254E-01	3.908E-01	7.447E-02	4.409E-01	3.739E-01	3.199E+00	1.711E+01	2.814E-01	1.702E-01	9.212E-01	3.908E-01	1.378E-02	2.024E-02	3.624E-01
39	1.796E+01	8.168E+01	3.821E-01	1.376E-01	9.611E-01	5.957E-01	6.656E-02	5.080E-01	5.791E-01	3.134E+00	1.814E+01	4.678E-01	1.485E-01	1.079E+00	5.887E-01	1.154E-02	1.996E-02	4.113E-01

Source:	Li X, Zhang D, Liu B, “Transmission of broad W/Rh and W/Al (target/filter) x-ray beams operated at 25-29 kVp through common shielding materials”, MedPhys 39(7) 4132