## Fluoro
1 | [11 empty] | Revision 2.3-20251003 | [10 empty] | Print Area
2 | [5 empty] | Medical University of South Carolina | [10 empty] | Medical University of South Carolina | [5 empty] | Page1,Page2,HVLPage,KAP_AKR,ExpChartFluoroPage,ExpChartDAPage,ImgQualityPage,Com1Page,Com2Page,OutputGraphFluoroPage,OutputGraphDAPage,LeedsTO10Page
3 | [5 empty] | Charleston, South Carolina | [10 empty] | Charleston, South Carolina
4 | [22 empty] | All: | Page1,Page2,HVLPage,KAP_AKR,ExpChartFluoroPage,ExpChartDAPage,ImgQualityPage,Com1Page,Com2Page,OutputGraphFluoroPage,OutputGraphDAPage,LeedsTO10Page
5 | [5 empty] | Fluoroscopy System Compliance Inspection | [10 empty] | Fluoroscopy System Compliance Inspection
6 | [22 empty] | Measurement Parameter | Last Year |  | This Year
7 | [11 empty] | Date: | [7 empty] | Inspector: | Eugene Mah | Eugene Mah | Print Area |  | Change | Page1,Page2,HVLPage,KAP_AKR,ExpChartFluoroPage,ExpChartDAPage,ImgQualityPage,Com1Page,Com2Page,OutputGraphFluoroPage,OutputGraphDAPage,LeedsTO10Page
8 | [5 empty] | System Information | [5 empty] | Previous Date: |  |  |  |  | System Information | [5 empty] | Date
9 |  | Location | [10 empty] | Location |  |  | Input Changes Only |  |  |  | Input Changes Only |  |  | Inspector |  | Change | Eugene Mah
10 |  |  | Facility: |  |  |  |  | Site Number: | [5 empty] | Facility: |  |  |  | Site Number: |  |  |  |  | Facility:
11 |  |  | Department: |  |  |  |  | Authorized Use: | [5 empty] | Department: |  |  |  | Authorized Use: |  |  |  |  | Department:
12 |  |  | Area/Division: |  |  |  |  | Date of Installation: | [5 empty] | Area/Division: |  |  |  | Date of Installation: |  |  |  |  | Area/Division:
13 |  |  | Survey ID: |  |  |  |  | Number of X-ray Tubes: | [5 empty] | Survey ID: |  |  |  | Number of X-ray Tubes: |  |  |  |  | Survey ID:
14 | [13 empty] | Room Number: |  |  |  | Accession Number: |  |  |  |  | Room Number:
15 |  | X-Ray Generator | [20 empty] | Site Number:
16 |  |  | Manufacturer: |  |  |  |  | Manufacture Date: |  |  |  |  | X-Ray Generator | [9 empty] | Authorized Use:
17 |  |  | Model: |  |  |  |  | Serial Number: | [5 empty] | Manufacturer: |  |  |  | Manufacture Date: |  |  |  |  | Date of Installation:
18 |  |  | Max kVp: |  |  |  |  | Max mA: | [5 empty] | Model: |  |  |  | Serial Number: |  |  |  |  | Number of X-ray Tubes:
19 | [7 empty] | Software Version: | [5 empty] | Max kVp: |  |  |  | Max mA: |  |  |  |  | Accesssion Number:
20 |  | X-Ray Tube 1 | [15 empty] | Software Version: |  |  |  |  | X-Ray Generator
21 |  |  | Tube Designation/Use: |  |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | X-Ray Tube 1 | [9 empty] | Manufacturer:
22 |  |  | Manufacture Date: |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Model:
23 |  | Insert | [5 empty] | Small: | [5 empty] | Manufacture Date: |  |  |  | Large: |  |  |  |  | Max kVp:
24 |  |  | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Manufacture Date:
25 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Micro: |  |  |  |  | Serial Number:
26 |  |  | Serial Number: |  |  |  |  | Filtration | [5 empty] | Model: | [8 empty] | Max mA:
27 |  | Housing | [5 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | Filtration |  |  |  |  | Software Version:
28 |  |  | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | X-Ray Tube 1
29 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Added: |  |  |  |  | Tube Designation/Use:
30 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | Manufacture Date:
31 | [13 empty] | Serial Number: | [8 empty] | Manufacturer:
32 |  | X-Ray Tube 2 | [20 empty] | Model:
33 |  |  | Tube Designation/Use: |  |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | X-Ray Tube 2 | [9 empty] | Serial Number:
34 |  |  | Manufacture Date: |  |  |  |  | Large: | [5 empty] | Tube Designation/Use: |  |  |  | Focal Spot Sizes (mm) |  |  |  |  | Manufacturer:
35 |  | Insert | [5 empty] | Small: | [5 empty] | Manufacture Date: |  |  |  | Large: |  |  |  |  | Model:
36 |  |  | Manufacturer: |  |  |  |  | Micro: |  |  |  |  | Insert |  |  |  |  | Small: |  |  |  |  | Serial Number:
37 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Micro: |  |  |  |  | Large:
38 |  |  | Serial Number: |  |  |  |  | Filtration | [5 empty] | Model: | [8 empty] | Small:
39 | [7 empty] | Inherent: | [5 empty] | Serial Number: |  |  |  | Filtration |  |  |  |  | Micro:
40 |  |  | Manufacturer: |  |  |  |  | Added: |  |  |  |  | Housing |  |  |  |  | Inherent: |  |  |  |  | Inherent:
41 |  |  | Model: | [10 empty] | Manufacturer: |  |  |  | Added: |  |  |  |  | Added:
42 |  |  | Serial Number: | [10 empty] | Model: | [8 empty] | X-Ray Tube 2
43 | [13 empty] | Serial Number: | [8 empty] | Tube Designation/Use:
44 | [22 empty] | Manufacture Date:
45 | [22 empty] | Manufacturer:
46 | [16 empty] | Inspection Results | [5 empty] | Model:
47 | [11 empty] | Enter 1 for YES, 2 for NO, 3 for NA | [10 empty] | Serial Number:
48 | [11 empty] | 2 | 2 | Is this a new installation? |  |  | Labels, Notices, Postings |  | Tube Number: |  |  |  | Manufacturer:
49 | [12 empty] | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | **Input only one of the following |  |  | Model:
50 | [12 empty] | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | numbers or names |  |  | Serial Number:
51 | [12 empty] | Radiation warning label posted on the generator control panel | [6 empty] | 1 | 1 |  | Large:
52 | [12 empty] | Tube loading charts for each focal spot are posted | [6 empty] | 2 | 2 |  | Small:
53 | [12 empty] | Tube loading charts are correct for installed tube(s). | [6 empty] | 3 | 3 |  | Micro:
54 | [12 empty] | Operator manuals are available. | [6 empty] | AP | AP |  | Inherent:
55 | [12 empty] | Monthly radiation monitoring reports are posted. | [6 empty] | Front | Front |  | Added:
56 | [16 empty] | Fluoroscopy Systems |  |  | Frontal | Frontal |  | Tube Number:
57 | [12 empty] | Interlock systems function to prevent fluoro while II is parked | [6 empty] | Lat | Lat |  | Tube Designation: |  | Change | 1
58 | [12 empty] | The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [6 empty] | Lateral | Lateral |  |  |  | Change | 2
59 | [12 empty] | Minimum of 2.5 mm lead equivalent shielding (i.e. drapes or sliding panels) | [11 empty] | Change | 3
60 | [12 empty] | The useful beam is intercepted by the image assembly at any SID | [11 empty] | Change | AP
61 | [24 empty] | Change | Front
62 | [11 empty] | 1 | Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [11 empty] | Change | Frontal
63 | [11 empty] | 1 | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [11 empty] | Change | Lat
64 | [11 empty] | 3 | High contrast resolution within MUSC specifications | [11 empty] | Change | Lateral
65 | Date: | [8 empty] | Inspector: | Eugene Mah | 3 | Low contrast resolution within MUSC specifications
66 | Room Number | [8 empty] | Survey ID: |  | 1 | Transmission through II < 2 uGy/hr for each mGy/min at patient entrance
1 | [10 empty] | Medical University of South Carolina |  | Beam limiting device available (functioning) to restrict the x-ray field | [9 empty] | Receptor Size
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | The useful beam is centered on the image receptor | [9 empty] | Normal:
3 | Rule Number |  |  | Labels, Notices, Postings | [5 empty] | Compliance |  |  | Collimator edges are visible at full open | [9 empty] | Mag 1:
4 | DHEC RHB 2.5.1.1 |  | DHEC Registration sticker is present, clearly visible and legible | [6 empty] | TBD |  |  | Elapsed timer (5 min maximum) functional | [9 empty] | Mag 2:
5 | DHEC RHB 11.2.3 |  | DHEC form SC-RHA-20 “Notice to Employees” posted or referenced | [6 empty] | TBD |  |  | Fluoroscopic kV and mA meter present and functional | [9 empty] | Mag 3:
6 | DHEC RHB 4.3.1 |  | Radiation warning label posted on the generator control panel | [6 empty] | TBD | [6 empty] | Radiation Safety | [5 empty] | Mag 4:
7 |  |  | Tube loading charts for each focal spot are posted | [6 empty] | TBD |  |  | Continuous pressure on exposure switch required to make an exposure with selected tube | [9 empty] | Pt Dose Measured Fluoro
8 |  |  | Tube loading charts are correct for installed tube(s). | [6 empty] | TBD |  |  | Operator has full visibility of the exposure factors and patient | [9 empty] | Manual/Auto:
9 |  |  | Operator manuals are available. | [6 empty] | TBD |  |  | Lead aprons available | [9 empty] | Lucite/Al: |  | Change | 1
10 |  |  | Monthly radiation monitoring reports are posted. | [6 empty] | TBD |  |  | Lead gloves available | [9 empty] | # dose rates:
11 | [12 empty] | Properly designed and installed apron rack is present | [9 empty] | High level present:
12 |  |  |  | Radiation Safety | [8 empty] | Documentation of annual protective apparel integrity inspection is available | [9 empty] | cm/in: |  | Change | cm
13 |  |  | Continuous pressure on exposure switch required to make an exposure with selected tube | [6 empty] | TBD |  |  | Patient restraint devices available | [9 empty] | Auto/mA: |  | Change | Auto
14 |  |  | Operator has full visibility of the exposure factors and patient | [6 empty] | TBD |  |  | Gonadal shielding is present and functional | [9 empty] | Test equipment: |  | Change | Piranha 1 CB2-17090320
15 | DHEC RHB 4.2.12.7 |  | Lead aprons available | [6 empty] | TBD |  |  | Radiation warning light at room entrance is functional with selected tube | [9 empty] | Calibration date:
16 | DHEC RHB 4.2.12.7 |  | Lead gloves available | [6 empty] | TBD | [6 empty] | Physical Inspection | [5 empty] | Calibration due:
17 |  |  | Properly designed and installed apron rack is present | [6 empty] | TBD |  |  | System clean and free of debris | [9 empty] | Technique 1:
18 | DHEC RHB 4.2.8 |  | Documentation of annual protective apparel integrity inspection is available | [6 empty] | TBD |  |  | All cables free from excessive wear or restraints | [9 empty] | Technique 2:
19 |  |  | Patient restraint devices available | [6 empty] | TBD |  |  | System free of missing or damaged components | [9 empty] | Technique 3:
20 | DHEC RHB 4.2.12.7 |  | Gonadal shielding is present and functional | [6 empty] | TBD |  |  | Tube support movements and locks are acceptable | [9 empty] | Maximum Dose Rates Fluoro
21 |  |  | Radiation warning light at room entrance is functional with selected tube | [6 empty] | TBD |  |  | All other movements and locks functional | [9 empty] | kV: |  |  | 0
22 | [22 empty] | mA: |  |  | 0
23 |  |  |  | Fluoroscopy Systems | [12 empty] | Automatic Brightness Control / Display Monitor | [5 empty] | mGy/min: |  |  | 0
24 | DHEC RHB 4.9.2.1.2 |  | Interlock systems function to prevent fluoro while II is parked | [6 empty] | TBD |  |  | Automatic Brightness Control maintains image brightness for varying attenuation | [9 empty] | kV: |  |  | 0
25 | DHEC RHB 4.9.1 |  | The minimum Source Skin Distance > 38 cm (20 cm for surgical applications) | [6 empty] | TBD |  |  | Display monitor is free of artifacts, shadows, bright spots, and burn-in | [9 empty] | mA: |  |  | 0
26 | DHEC RHB 4.9.8.1 |  | Minimum of 0.25 mm lead equivalent shielding (i.e. drapes or sliding panels) | [6 empty] | TBD |  |  | Display monitor is free of distortion | [9 empty] | mGy/min: |  |  | 0
27 | DHEC RHB 4.9.2.1.1 |  | The useful beam is intercepted by the image assembly at any SID | [6 empty] | TBD |  |  | Display monitor is free of excessive persistence (after-image) | [9 empty] | kV: |  |  | 0
28 | [12 empty] | Display monitor is free of static or rolling noise bands | [9 empty] | mA: |  |  | 0
29 | DHEC RHB 4.9.4.2.1 |  | Max Entrance Exposure rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [6 empty] | YES |  |  | The last image hold (LIH) display is free of jitter or abnormal appearance | [9 empty] | mGy/min: |  |  | 0
30 | DHEC RHB 4.9.4.2.1.2 |  | Max Entrance Exposure Rate <= 176 mGy/min (high level) | [6 empty] | YES | [12 empty] | Maximum Dose Rates Pulse/Dig
31 |  |  | High contrast resolution within MUSC specifications | [6 empty] | NA | [6 empty] | Inspection Results | [5 empty] | kV: |  |  | 0
32 |  |  | Low contrast resolution within MUSC specifications | [6 empty] | NA | [12 empty] | mA: |  |  | 0
33 | DHEC RHB 4.9.5 |  | Transmission through II < 2 uGy/hr for each mGy/min at patient entrance | [6 empty] | YES |  | Entrance Exposure Rate |  |  |  | Type 1 for manual, leave blank for auto |  |  | Auto | Auto | Auto/mA mode |  | mGy/min: |  |  | 0
34 | DHEC RHB 4.9.2.2.1 |  | Beam limiting device available (functioning) to restrict the x-ray field | [6 empty] | TBD |  |  |  | 1 | 1 | Type 1 if using Lucite, leave blank if using Al |  |  | Piranha 1 CB2-17090320 | Piranha 1 CB2-17090320 | Test equipment |  | kV: |  |  | 0
35 |  |  | The useful beam is centered on the image receptor | [6 empty] | TBD |  |  |  | How many dose rate settings are present (exclusing High Level)? | [6 empty] | Calibration Date |  | mA: |  |  | 0
36 | DHEC RHB 4.9.2.2 |  | Collimator edges are visible at full open | [6 empty] | TBD |  |  |  | Is High Level provided? (1=Yes, 2=No) | [6 empty] | Calibration due |  | mGy/min: |  |  | 0
37 | DHEC RHB 4.9.7.1 |  | Elapsed timer (5 min maximum) functional | [6 empty] | TBD |  | Blank | [10 empty] | kV: |  |  | 0
38 | DHEC RHB 4.9.6 |  | Fluoroscopic kV and mA meter present and functional | [6 empty] | TBD |  | cm | cm | Unit for image receptor size | [8 empty] | mA: |  |  | 0
39 | DHEC RHB 4.3.5 |  | The Measured HVL meeds DHEC minimum requirements | [6 empty] | TBD |  | Patient Entrance Exposure Rate (Pulse Fluoroscopy) |  |  |  |  | Mode: | 2 | Enter 1 for continuous fluoro, or 2 for pulsed fluoro |  |  |  | mGy/min: |  |  | 0
40 |  |  | Kerma area product display is accurate to within 40% | [6 empty] | #DIV/0! | #DIV/0! | Image | Acrylic | [9 empty] | HVL
41 |  |  | Air kerma display is accurate to within 40% | [6 empty] | NA |  | Mode | Attenuator | Auto 1 | [5 empty] | Auto |  |  | HVL Measured kV: |  | Change | 80
42 |  |  | Air kerma rate display is accurate to within 40% | [6 empty] | #DIV/0! | #DIV/0! | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | HVL Measured: |  | Change | TBD
43 |  |  | kVp accurate to within 5% of indicated. | [6 empty] | NA |  | Normal | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Pt Dose Measured Dig Acq
44 |  |  | kV reproducibility coeff. Of var. < 5%. | [6 empty] | NA |  |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Technique 1:
45 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Technique 2:
46 |  |  |  | Automatic Brightness Control / Display Monitor | [7 empty] | Mag 1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Technique 3:
47 |  |  | Automatic Brightness Control maintains image brightness for varying attenuation | [6 empty] | TBD |  |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Receptor Entr Exp Fluoro
48 |  |  | Display monitor is free of artifacts, shadows, bright spots, and burn-in | [6 empty] | TBD |  |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto 1 Normal |  |  | 0
49 |  |  | Display monitor is free of distortion | [6 empty] | TBD |  | Mag 2 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto 1 Mag 1 |  |  | 0
50 |  |  | Display monitor is free of excessive persistence (after-image) | [6 empty] | TBD |  |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto 1 Mag 2 |  |  | 0
51 |  |  | Display monitor is free of static or rolling noise bands | [6 empty] | TBD |  |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto 1 Mag 3 |  |  | 0
52 |  |  | The last image hold (LIH) display is free of jitter or abnormal appearance | [6 empty] | TBD |  | Mag 3 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto 1 Mag 4 |  |  | 0
53 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Normal |  |  | 0
54 |  |  |  | Physical Inspection | [8 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Mag 1 |  |  | 0
55 |  |  | System clean and free of debris | [6 empty] | TBD |  | Mag 4 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Mag 2 |  |  | 0
56 |  |  | All cables free from excessive wear or restraints | [6 empty] | TBD |  |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Mag 3 |  |  | 0
57 |  |  | System free of missing or damaged components | [6 empty] | TBD |  |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Mag 4 |  |  | 0
58 |  |  | Tube support movements and locks are acceptable | [6 empty] | TBD |  |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto Normal |  |  | 0
59 |  |  | All other movements and locks functional | [6 empty] | TBD |  |  | Last Year: | [9 empty] | Auto Mag 1 |  |  | 0
60 | [12 empty] | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES | Auto Mag 2 |  |  | 0
61 | [12 empty] | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | Auto Mag 3 |  |  | 0
62 | [12 empty] | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point) | [9 empty] | Auto Mag 4 |  |  | 0
63 | [12 empty] | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [9 empty] | Receptor Entr Exp Dig Acq
64 | [12 empty] | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point) | [9 empty] | Auto 1 Normal |  |  | 0
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [8 empty] | Auto 1 Mag 1 |  |  | 0
66 | Room Number | [8 empty] | Survey ID: |  |  |  | Max Entrance Exposure Rate < = 176 mGy/min (high level) | [8 empty] | Auto 1 Mag 2 |  |  | 0
1 | [10 empty] | Medical University of South Carolina |  | Comments: | [9 empty] | Auto 1 Mag 3 |  |  | 0
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | Do not exceed 390 chars |  | 0 | [7 empty] | Auto 1 Mag 4 |  |  | 0
3 | Half Value Layer Measurement | [11 empty] | New: | [9 empty] | Normal |  |  | 0
4 | Set kVp: | 80 | [20 empty] | Mag 1 |  |  | 0
5 | [11 empty] | Patient Entrance Exposure Rate – Pulse Fluroscopy |  |  |  |  | Mode: | 1 | Enter 1 for pulsed fluoro, 2 for digital acquisition mode |  |  |  | Mag 2 |  |  | 0
6 | Al (mm) | mGy/min | mGy/min |  | Measured HVL: | TBD |  |  | ACCEPTABLE RESULTS (mm Al) |  |  | Image | Acrylic | [9 empty] | Mag 3 |  |  | 0
7 | 0 |  |  |  | HVL Acceptable: | TBD |  |  | Min acceptable HVL at 80 kVp: | 2.6 |  | Mode | Attenuator | Auto 1 | [5 empty] | Auto |  |  | Mag 4 |  |  | 0
8 | 0 | [7 empty] | Desired HVL at 80 kVp: | 3.1 |  | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | Auto Normal |  |  | 0
9 | 0 | [6 empty] | Criteria: | MUSC Standard |  |  | Normal | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto Mag 1 |  |  | 0
10 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto Mag 2 |  |  | 0
11 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto Mag 3 |  |  | 0
12 | [11 empty] | Mag 1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Auto Mag 4 |  |  | 0
13 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Fluoro Scatter/Transmission
14 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Transmission: |  |  | 0
15 | [11 empty] | Mag 2 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (eyes):
16 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (waist):
17 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (eyes):
18 | [11 empty] | Mag 3 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (waist):
19 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dig Acq Scatter
20 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (eyes):
21 | [11 empty] | Mag 4 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Unshielded Scatter (waist):
22 | [12 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (eyes):
23 | [12 empty] | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Shielded Scatter (waist):
24 | [12 empty] | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Leeds Phantom
25 | [12 empty] | Last Year: | [9 empty] | kVp: |  | Change | 70
26 | [12 empty] | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES | Huttner Normal:
27 | [12 empty] | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | Huttner Mag 1:
28 | [12 empty] | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point) | [9 empty] | Huttner Mag 2:
29 | [12 empty] | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [9 empty] | Huttner Mag 3:
30 | [12 empty] | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point) | [9 empty] | Huttner Mag 4:
31 | [12 empty] | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [8 empty] | N3 Normal:
32 | [13 empty] | Max Entrance Exposure Rate < = 176 mGy/min (high level) | [8 empty] | N3 Mag 1:
33 | [12 empty] | Comments: | [9 empty] | N3 Mag 2
34 | [12 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | N3 Mag 3
35 | [12 empty] | New: | [9 empty] | N3 Mag 4:
36 | [22 empty] | Field Size:
37 | [11 empty] | Half Value Layer | [10 empty] | SID:
38 | [11 empty] | 2 | Use DHEC standards? (Type 1 to base acceptance on DHEC criteria) | [9 empty] | Front:
39 | [22 empty] | Back:
40 | [11 empty] | kV | mA | mm Al | mGy/min |  | Data Processing | [5 empty] | Left:
41 | [11 empty] | 80 | 0 | 0 |  |  | mm Al | Avg Exp | DATA(1/2) | Data Al | HVL |  | Right:
42 | [13 empty] | 0 |  |  | 0 |  |  |  | TBD |  | TO.10.A
43 | [13 empty] | 0 |  |  | 0 | [5 empty] | TO.10.B
44 | [13 empty] | 0 |  |  | 0 | [5 empty] | TO.10.C
45 | [13 empty] | 0 | [5 empty] | Data Exp | LN(EXP) |  | TO.10.D
46 | [13 empty] | 0 | [8 empty] | TO.10.E
47 | [22 empty] | TO.10.F
48 | [16 empty] | TBD | TBD |  |  |  |  | TO.10.G
49 | [16 empty] | Acceptable (mm Al) |  |  | HVL Acceptable: | TBD |  | TO.10.H
50 | [16 empty] | Low | 2.6 |  |  |  |  | TO.10.J
51 | [16 empty] | Desired | 3.1 |  | Previous HVL @ kV: |  |  | TO.10.K
52 | [16 empty] | High | 3.5 |  |  |  |  | TO.10.L
53 | [22 empty] | TO.10.M
54 | [11 empty] | Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy | [10 empty] | TO.10.A
55 | [22 empty] | TO.10.B
56 | [11 empty] | Dose | Receptor |  |  |  |  | Previous Year |  |  |  |  | TO.10.C
57 | [11 empty] | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | nGy/frame |  |  |  |  | TO.10.D
58 | [11 empty] | Auto 1 |  | 0 | 0 | 0 | 0 |  |  | Measure using pancake probe placed against |  |  | TO.10.E
59 | [13 empty] | 0 | 0 | 0 | 0 |  |  | image receptor input window |  |  | TO.10.F
60 | [13 empty] | 0 | 0 | 0 | 0 |  |  | Use 1.5mm Cu filtration |  |  | TO.10.G
61 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.H
62 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.J
63 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.K
64 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.L
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.M
66 | Room Number | [8 empty] | Survey ID: |  |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.A
1 | [10 empty] | Medical University of South Carolina |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.B
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection | Auto |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.C
3 | Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate | [12 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.D
4 | Test equipment: | Piranha 1 CB2-17090320 |  | Calibration date: |  |  | Calibration due: | [6 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.E
5 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.F
6 |  |  | Ind KAP mGy·cm² | Meas KAP mGy·cm² | KAP Deviation | Ind AK @ Ref Pt (mGy) | Meas AK @ Ref Pt (mGy) | AK Deviation | Ind AKR | Meas AKR | AKR Deviation |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.G
7 | kV | mA | [6 empty] | mGy/min | mGy/min | [12 empty] | TO.10.H
8 | 60 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0! |  | Comments: | [9 empty] | TO.10.J
9 | 80 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0! |  |  |  | 0 | [7 empty] | TO.10.K
10 | 100 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0! |  | New: | [9 empty] | TO.10.L
11 | 120 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0! | [11 empty] | TO.10.M
12 | [11 empty] | Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy | [10 empty] | TO.10.A
13 | [22 empty] | TO.10.B
14 | [11 empty] | Dose | Receptor |  |  |  |  | Previous Year |  |  |  |  | TO.10.C
15 | [11 empty] | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | nGy/frame |  |  |  |  | TO.10.D
16 |  |  |  | Acceptable: | #DIV/0! |  | Acceptable: |  |  | Acceptable: | #DIV/0! | Auto 1 |  | 0 | 0 | 0 | 0 |  |  | Measure using pancake probe placed against |  |  | TO.10.E
17 |  | Criteria: | Measured KAP/AK/AKR is within 40% of the displayed values | [10 empty] | 0 | 0 | 0 | 0 |  |  | image receptor input window |  |  | TO.10.F
18 | [13 empty] | 0 | 0 | 0 | 0 |  |  | Use 1.5mm Cu filtration |  |  | TO.10.G
19 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.H
20 | kV Accuracy | [5 empty] | kV Reproducibility | [6 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.J
21 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.K
22 |  | Test Equipment Used: | Piranha 1 CB2-17090320 |  |  | Calibration Date: |  |  | Calibration Due: |  |  |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.L
23 | [13 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.M
24 |  |  | Measured |  | Exp Rate |  |  | Measured | Exp Rate |  |  |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.A
25 | kV Set | mA Set | kV | % error | (mGy/min) |  | kV Set | kV | (mGy/min) |  |  |  |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.B
26 | 60 | 5 |  |  |  |  | 80 |  |  |  |  | Auto |  | 0 | 0 | 0 | 0 | [5 empty] | TO.10.C
27 | 80 | 5 | [11 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.D
28 | 100 | 5 | [11 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.E
29 | 120 | 5 | [11 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.F
30 |  | Acceptable: | [11 empty] | 0 | 0 | 0 | 0 | [5 empty] | TO.10.G
31 |  | Criteria: | kVp accurate to within 5% of indicated. |  |  |  | Mean: | [15 empty] | TO.10.H
32 | [6 empty] | Std Dev: | [5 empty] | Comments: | [9 empty] | TO.10.J
33 | [6 empty] | Coeff of Var | [5 empty] | Do not exceed 390 chars |  | 0 | [7 empty] | TO.10.K
34 | [6 empty] | Acceptable: | [5 empty] | New: | [9 empty] | TO.10.L
35 | [6 empty] | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%. | [14 empty] | TO.10.M
36 | [11 empty] | Scatter and Transmission | [10 empty] | Comments
37 | Air Kerma Linearity | [14 empty] | Pt Entrance | [6 empty] | Pt Entr Exp Fluoro:
38 | [13 empty] | kVp | mA | mGy/min | uGy/hr | Last Year |  | Acceptable |  |  | Pt Entr Exp Dig Acq:
39 | kV Set | mA | mGy/mA-min | [9 empty] | Transmission* @ 10cm | 0 | 0 | 0 | 0 |  |  | YES |  |  | Receptor Entr Exp Fluoro:
40 | 80 | 1 | [13 empty] | Unshielded | Shielded | Unshielded | Shielded |  |  |  | Receptor Entr Exp Dig Acq:
41 |  | 5 | [13 empty] | uGy/hr | uGy/hr | uGy/hr | uGy/hr |  |  |  | Scatter/Transmission:
42 |  | 10 | [9 empty] | Scatter – Fluoro | Eye level | 0 | 0 | [7 empty] | Leeds Phantom:
43 |  | 15 | [10 empty] | Waist level | 0 | 0 | [7 empty] | TO.10:
44 |  | Coeff of Linearity: | [9 empty] | Scatter - Pulse | Eye level | 0 | 0 | [7 empty] | Beam Limitation:
45 |  | Acceptable: | [10 empty] | Waist level | 0 | 0 | [7 empty] | Comment Page 1
46 | Criteria: | Coefficient of linearity <= 10%. | [10 empty] | *Measured with 8” Lucite in beam
47 | [12 empty] | Largest field size mode, highest dose rate mode
48 | [12 empty] | Criteria: | Transmission through the image receptor is < 2 uGy/hr at 10 cm for each mGy/min of patient entrance exposure
49 | [12 empty] | Comments:
50 | [12 empty] | Do not exceed 390 chars |  | 0
51 | [12 empty] | New:
52
53 | [11 empty] | Kerma Area Product/Cumulative Air Kerma/Air Kerma Rate
54 | [11 empty] | 30 | Source-PDC distance (cm) |  |  |  | Source-Isocenter
55 | [11 empty] | 30 | Source-Ref Pt distance (cm) |  |  |  | Table-Isocenter
56 | [11 empty] | 1 | CAK Reference Point correction factor |  |  |  | RefPt-Isocenter
57 | [11 empty] | 1 | Apply criteria
58 | [13 empty] | Start KAP | End KAP |  | Start AK | End AK |  | Ind AKR | Meas AKR
59 | [11 empty] | kV | mA | mGy·cm² |  | KAP | mGy |  | CAK | mGy/min | mGy/min
60 | [11 empty] | 60 | 0 | 0 | 0 | 0 |  |  |  | 0 | 0
61 | [11 empty] | 80 | 0 | 0 | 0 | 0 |  |  |  | 0 | 0
62 | [11 empty] | 100 | 0 | 0 | 0 | 0 |  |  |  | 0 | 0
63 | [11 empty] | 120 | 0 | 0 | 0 | 0 |  |  |  | 0 | 0
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection | Radcal PDC |  | Test equipment: | Piranha 1 CB2-17090320 |  | Calibration date: |  |  | Calibration due:
3 | Patient Entrance Exposure Rate (Pulse Fluoroscopy) |  |  |  |  | Test equipment: | Piranha 1 CB2-17090320 | Calibration date: |  | Calibration due: |  | kV | mA | mGy | mGy/s | mGy·cm² | mGy·cm²/s
4 | Image | Acrylic | [9 empty] | 60 | 0 |  |  | 0 | 0
5 | Mode | Attenuator | Auto 1 | [5 empty] | Auto |  |  | 80 | 0 |  |  | 0 | 0
6 | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min | 100 | 0 |  |  | 0 | 0
7 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 120 | 0 |  |  | 0 | 0
8 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
12 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ind KAP mGy·cm² | Meas KAP mGy·cm² |  | Ind AK @ Ref Pt (mGy) | Meas AK @ Ref Pt (mGy) |  | Ind AKR | Meas AKR
13 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | KAP Deviation |  |  | CAK Deviation | mGy/min | mGy/min | Deviation
14 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0!
15 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0!
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0! |  |  | Comment Page 2
17 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! |  |  |  | 0 | 0 | #DIV/0!
18 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Acceptable: | #DIV/0! |  | Acceptable: |  |  | Acceptable: | #DIV/0!
23
24 |  | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES |  | Criteria: | Measured KAP/AK/AKR is within 40% of the displayed values
25 |  | Undertable fluoroscopy units: 1 cm above table top | [9 empty] | kV Accuracy and Reproducibility
26 |  | Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27 |  | C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point | [10 empty] | Test Equipment Used: | Piranha 1 CB2-17090320 |  | Calibration Date: |  |  | Calibration Due:
28 |  | Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29 |  | Criteria: | Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual) | [10 empty] | Measured values
30 |  |  | Max Entrance Exposure Rate < = 176 mGy/min (high level) | [8 empty] | kV set | mA Set | kV | Exposure (mGy) | Exp rate (mGy/min) | HVL | Tot Filt | kV % error
31 | Comments: | [10 empty] | 60 | 5
32 | [11 empty] | 80 | 5
33 | [11 empty] | 100 | 5
34 | [11 empty] | 120 | 5
35 | Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy | [16 empty] | Acceptable:
36 | [12 empty] | Criteria: | kVp accurate to within 5% of indicated.
37 | Dose | Receptor
38 | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | [6 empty] | Measured values | [5 empty] | Air Kerma Rate Linearity
39 | Auto 1 |  | 0 | 0 | 0 | 0 |  | Measure using pancake probe placed against |  |  |  | kV Set: | kV | Exposure (mGy) | Exp rate (mGy/min) | HVL | Tot Filt |  | kV Set: | mA | mGy/mA-min
40 |  |  | 0 | 0 | 0 | 0 |  | image receptor input window |  |  |  | 80 | [6 empty] | 80 | 1
41 |  |  | 0 | 0 | 0 | 0 |  | Use 1.5mm Cu filtration | [11 empty] | 5
42 |  |  | 0 | 0 | 0 | 0 | [13 empty] | 10
43 |  |  | 0 | 0 | 0 | 0 | [13 empty] | 15
44 |  |  | 0 | 0 | 0 | 0 | [13 empty] | Coeff of Linearity:
45 |  |  | 0 | 0 | 0 | 0 | [5 empty] | Mean: | [7 empty] | Acceptable:
46 |  |  | 0 | 0 | 0 | 0 | [5 empty] | Std Dev: | [6 empty] | Criteria: | Coefficient of linearity <= 10%.
47 |  |  | 0 | 0 | 0 | 0 | [5 empty] | Coeff of Var
48 |  |  | 0 | 0 | 0 | 0 | [5 empty] | Acceptable:
49 | Auto |  | 0 | 0 | 0 | 0 | [6 empty] | Criteria: | kVp coefficient of variation (std dev./mean) <= 5%.
50 |  |  | 0 | 0 | 0 | 0
51 |  |  | 0 | 0 | 0 | 0
52 |  |  | 0 | 0 | 0 | 0 | [5 empty] | Leeds Phantom
53 |  |  | 0 | 0 | 0 | 0
54 | Comments: | [10 empty] | Receptor Size | MS1 | MS3 | MS4 | Huttner | N3 | kVp: | 70 | 70 |  |  | New installation? |  | Change | 2
55 | [18 empty] | MS1-4 |  | Huttner: Enter col.row
56 | [18 empty] | 1 | Clearly visible | N3: Enter # of visible disks
57 | Scatter and Transmission – Fluoroscopy |  |  |  | Pt Entrance | [13 empty] | 2 | Visible
58 |  |  | kVp | mA | mGy/min | uGy/hr | Unshielded | Shielded |  | Acceptable | [8 empty] | 3 | Barely visible
59 |  | Transmission* @ 10cm | 0 | 0 | 0 | 0 | uGy/hr | uGy/hr |  | YES | [8 empty] | 4 | Not visible
60 | Scatter – Fluoro | Eye level | 0 | 0 | [10 empty] | Acceptable: | NA | NA
61 |  | Waist level | 0 | 0 | [8 empty] | Comments:
62 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
63 | [12 empty] | New:
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah | TO.10 |  | Details Detected
66 | Room Number | [8 empty] | Survey ID: |  |  | Row
1 | [10 empty] | Medical University of South Carolina |  | A
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | B
3 | Patient Entrance Exposure Rate – Pulse Fluroscopy |  |  |  |  | Test equipment: | Piranha 1 CB2-17090320 | Calibration date: |  | Calibration due: |  |  | C
4 | Image | Acrylic | [10 empty] | D
5 | Mode | Attenuator | Auto 1 | [5 empty] | Auto |  |  |  | E
6 | (cm) | (cm) | kVp | mA | mGy/min | kVp | mA | mGy/min | kVp | mA | mGy/min |  | F
7 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | G
8 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | H
9 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | J
10 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | K
11 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | L
12 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | M
13 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Comments:
15 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Do not exceed 390 chars |  | 0
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | New:
17 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Field Size
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Horizontal | [6 empty] | Type 1 for circular field, 2 for square field
20 |  | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Vertical
21 |  | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Diagonal
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | Acceptable | NA | NA | NA | NA | NA
23 | [13 empty] | Criteria: | Measured field size is at least 85% of nominal field size
24 | Comments: | [10 empty] | Useful Beam Limitation
25 | [13 empty] | SID (cm)
26
27 | [12 empty] | Measured | Difference | Sum | Last Year |  |  | T | B | L | R
28 | [11 empty] | Top | [5 empty] | Visible
29 | [11 empty] | Bottom | [5 empty] | X-ray
30 | [11 empty] | Left |  |  |  |  | Acceptable | SID
31 | [11 empty] | Right | [5 empty] | SOD
32 | Image Receptor Entrance Exposure Rate – Pulse Fluoroscopy | [10 empty] | Criteria: | Difference between x-ray field and visible image area is less than 2% SID |  |  |  |  | SID/SOD
33 | [12 empty] | Sum of length and width differences is less than 4% SID
34 | Dose | Receptor | [10 empty] | Comments:
35 | Level | Size (cm) | kVp | mA | uGy/min | nGy/frame | [6 empty] | Do not exceed 390 chars |  | 0
36 | Auto 1 |  | 0 | 0 | 0 | 0 |  | Measure using pancake probe placed against |  |  |  |  | New:
37 |  |  | 0 | 0 | 0 | 0 |  | image receptor input window
38 |  |  | 0 | 0 | 0 | 0 |  | Use 1.5mm Cu filtration
39 |  |  | 0 | 0 | 0 | 0 | [10 empty] | Comments Page 1
40 |  |  | 0 | 0 | 0 | 0 | [9 empty] | Previous Year Comments
41 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Comments:
42 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Do not exceed 390 chars |  | 0 |  |  | **Patient exposure/dose measurements from the previous report were not posted at the time of inspection. The exposure measurements provided with this report should be posted near the operator's console.
43 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Additional Comments:
44 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Do not exceed 390 chars |  | 0 |  |  | **Monthly radiation monitoring reports were not available at the time of inspection. These should be posted in accordance with the requirements of the radiation safety office.
45 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Additional Comments:
46 | Auto |  | 0 | 0 | 0 | 0 | [6 empty] | Do not exceed 390 chars |  | 0 |  |  | A copy of the fluoroscopic skin entrance doses (attached) must be posted on the unit in a location visible to the operator (DHEC RHB 4.9.4.5.6).
47 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Additional Comments:
48 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Do not exceed 390 chars |  | 0 |  |  | All lead aprons must be inspected and documented on an annual basis (DHEC RHB 4.2.14.7). Please provide a copy of documentation of inspection or arrange to have lead aprons inspected.
49 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Additional Comments:
50 |  |  | 0 | 0 | 0 | 0 | [6 empty] | Do not exceed 390 chars |  | 0 |  |  | **The DHEC registration tag was not present on the system or was damaged and should be replaced. Arrangements must be made with the Radiation Safety Office to obtain the required registration tag.
51 | Comments: | [11 empty] | Additional Comments:
52 | [12 empty] | Do not exceed 390 chars |  | 0
53 | [12 empty] | Additional Comments:
54 | [12 empty] | Do not exceed 390 chars |  | 0
55 | Scatter and Transmission – Pulse Fluoroscopy | [11 empty] | Additional Comments:
56 |  |  |  |  | Unshielded | Shielded | [6 empty] | Do not exceed 390 chars |  | 0
57 |  |  | kVp | mA | uGy/hr | uGy/hr | [6 empty] | Additional Comments:
58 | Scatter - Pulse | Eye level | 0 | 0 | [8 empty] | Do not exceed 390 chars |  | 0
59 |  | Waist level | 0 | 0 | [8 empty] | Additional Comments:
60 | [12 empty] | Do not exceed 390 chars |  | 0
61 | [12 empty] | Additional Comments:
62 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
63 | [12 empty] | Additional Comments:
64 | [12 empty] | Do not exceed 390 chars |  | 0
65 | Date: | [8 empty] | Inspector: | Eugene Mah |  | Additional Comments:
66 | Room Number | [8 empty] | Survey ID: |  |  | Do not exceed 390 chars |  | 0
1 | [10 empty] | Medical University of South Carolina |  | Additional Comments:
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection |  | Do not exceed 390 chars |  | 0
3 | Leeds Test Objects |  |  |  |  | High contrast resolution |  | Low contrast resolution |  |  |  |  | Additional Comments:
4 |  | Receptor Size | MS1 | MS3 | MS4 | lp/mm | Acceptable | N3 | Acceptable |  |  |  | Do not exceed 390 chars |  | 0
5 | [12 empty] | Additional Comments:
6 | [12 empty] | Do not exceed 390 chars |  | 0
7 | [12 empty] | Additional Comments:
8 | [12 empty] | Do not exceed 390 chars |  | 0
9 | [12 empty] | Additional Comments:
10 |  | Criteria: | Resolution: >1.0 lp/mm Normal mode, >1.25 lp/mm Mag 1, >1.4 lp/mm Mag 2, > 2 lp/mm Mag 3, > 2.5 lp/mm Mag 4 | [9 empty] | Do not exceed 390 chars |  | 0
11 |  |  | N3: Minimum contrast level < 0.04 in all II modes | [9 empty] | Additional Comments:
12 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
13
14
15 | [16 empty] | Comments Page 2
16 | TO.10 Contrast Resolution | [14 empty] | Previous Year Comments
17 |  | Contrast Detail |  |  |  |  | Threshold Index | [5 empty] | Comments:
18 | Row | [11 empty] | Do not exceed 390 chars |  | 0
19 | A | [11 empty] | Additional Comments:
20 | B | [11 empty] | Do not exceed 390 chars |  | 0
21 | C | [11 empty] | Additional Comments:
22 | D | [11 empty] | Do not exceed 390 chars |  | 0
23 | E | [11 empty] | Additional Comments:
24 | F | [11 empty] | Do not exceed 390 chars |  | 0
25 | G | [11 empty] | Additional Comments:
26 | H | [11 empty] | Do not exceed 390 chars |  | 0
27 | J | [11 empty] | Additional Comments:
28 | K | [11 empty] | Do not exceed 390 chars |  | 0
29 | L | [11 empty] | Additional Comments:
30 | M | [11 empty] | Do not exceed 390 chars |  | 0
31 | Comments: | [11 empty] | Additional Comments:
32 | [12 empty] | Do not exceed 390 chars |  | 0
33 | [12 empty] | Additional Comments:
34 | [12 empty] | Do not exceed 390 chars |  | 0
35 | Field Size | [11 empty] | Additional Comments:
36 | [12 empty] | Do not exceed 390 chars |  | 0
37 |  | Horizontal | [10 empty] | Additional Comments:
38 |  | Vertical | [10 empty] | Do not exceed 390 chars |  | 0
39 |  | Diagonal | [10 empty] | Additional Comments:
40 |  | Acceptable | NA | NA | NA | NA | NA | [5 empty] | Do not exceed 390 chars |  | 0
41 |  | Criteria: | Measured field size is at least 85% of nominal field size | [9 empty] | Additional Comments:
42 | [12 empty] | Do not exceed 390 chars |  | 0
43 | Useful Beam Limitation | [11 empty] | Additional Comments:
44 |  | SID (cm) | [10 empty] | Do not exceed 390 chars |  | 0
45 |  |  | Measured | Difference | Sum | [7 empty] | Additional Comments:
46 |  | Top | [10 empty] | Do not exceed 390 chars |  | 0
47 |  | Bottom |  |  |  |  | Acceptable | [5 empty] | Additional Comments:
48 |  | Left | [10 empty] | Do not exceed 390 chars |  | 0
49 |  | Right | [10 empty] | Additional Comments:
50 |  | Criteria: | Difference between x-ray field and visible image area is less than 2% SID | [9 empty] | Do not exceed 390 chars |  | 0
51 |  |  | Sum of length and width differences is less than 4% SID | [9 empty] | Additional Comments:
52 | Comments: | [11 empty] | Do not exceed 390 chars |  | 0
53 | [12 empty] | Additional Comments:
54 | [12 empty] | Do not exceed 390 chars |  | 0
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Comments Page 1 |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Comments Page 2 |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Radiation Output – Fluoroscopy |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Radiation Output - Pulse Fluoroscopy |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Contrast Resolution |  |  |  |  | Fluoroscopy System Compliance Inspection
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65 | Date: | [8 empty] | Inspector: | Eugene Mah
66 | Room Number | [8 empty] | Survey ID:

## ExpCharts
1 | [10 empty] | Medical University of South Carolina
2 | [5 empty] | Inspection Results |  |  |  |  | Fluoroscopy System Compliance Inspection
3 | Patient Entrance Exposure Rate (Pulse Fluoroscopy) |  |  |  |  | Test equipment: | Piranha 1 CB2-17090320 | Calibration date: |  | Calibration due:
4 | Image | Acrylic
5 | Mode | Attenuator | Auto 1 | 0 | 0 |  | 0 | 0 | Auto | 0 | 0
6 | (cm) | (cm) | kVp | mA | R/min | kVp | mA | R/min | kVp | mA | R/min
7 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
8 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
12 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
13 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
16 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
17 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
19 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
22 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
23
24 |  | *Measured at: |  | Acceptable: | YES |  | Acceptable: | YES |  | Acceptable: | YES
25 | (cm): Undertable fluoroscopy units: 1 cm above table top
26 | (cm): Overhead tube fluoroscopy units: 30 cm above table top (Tube as close as possible to measuring point)
27 | (cm): C-arm fluoroscopy units: 30 cm from image receptor input (Tube at any ID, but spacer no closer than measuring point
28 | (cm): Lateral fluoroscopy units: 15 cm from center line of table and towards the tube (Position table and tube so spacer is as close as possible to measuring point)
29 | (cm): Criteria:; kVp: Max Entrance Exposure Rate <= 88 mGy/min (auto) and <= 44 mGy/min (manual)
30 | kVp: Max Entrance Exposure Rate < = 176 mGy/min (high level)
31 | (cm): Comments:
32
33
34
35 | (cm): Patient Entrance Exposure Rate – Pulse Fluroscopy; kVp: Test equipment:; mA: Piranha 1 CB2-17090320; R/min: Calibration date:; mA: Calibration due:
36 | (cm): Image; (cm): Acrylic
37 | Mode | Attenuator | Auto 1 | 0 | 0 |  | 0 | 0 | Auto | 0 | 0
38 | (cm) | (cm) | kVp | mA | R/min | kVp | mA | R/min | kVp | mA | R/min
39 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
40 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
41 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
42 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
43 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
44 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
45 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
46 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
47 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
48 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
49 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
50 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
51 | cm | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
52 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
53 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
54 |  | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
55
56 | (cm): Comments:
57
58
59
60
61
62
63
64
65 | (cm): Date:; mA: Inspector:; R/min: Eugene Mah
66 | (cm): Room Number; mA: Survey ID:

## Tables
Leeds MS |  |  | RTI Piranha | [10 empty] | Patient Entrance Exposure Rate (Pulse Fluoroscopy) |  |  | Sorted by kV |  |  |  | Patient Entrance Exposure Rate – Pulse Fluroscopy |  |  | Sorted by kV
1 | Clearly visible |  | Piranha 1 | Piranha 1 CB2-17090320 | [9 empty] | kV | Fluoro | High Level | kV | Fluoro | High Level |  | kV | Fluoro | High Level | kV | Fluoro | High Level
2 | Visible |  | Piranha 2 | Piranha 2 CB2-19020491 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3 | Barely visible | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
4 | Not visible | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
[14 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
Huttner Grating |  |  | Leeds N3 |  | 65 | 70 | 75 | 80 | Nominal |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
1.1 | 0.5 |  | 1 | 0.167 | 0.188 | 0.167 | 0.149 | 0.135 | 0.16 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
1.2 | 0.56 |  | 2 | 0.148 | 0.167 | 0.148 | 0.132 | 0.12 | 0.145 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
1.3 | 0.63 |  | 3 | 0.128 | 0.145 | 0.128 | 0.114 | 0.104 | 0.123 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
2.1 | 0.71 |  | 4 | 0.109 | 0.124 | 0.109 | 0.097 | 0.088 | 0.108 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
2.2 | 0.8 |  | 5 | 0.0876 | 0.0993 | 0.0876 | 0.078 | 0.0705 | 0.086 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
2.3 | 0.9 |  | 6 | 0.0749 | 0.0847 | 0.0749 | 0.067 | 0.0609 | 0.076 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3.1 | 1 |  | 7 | 0.0674 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | 0.066 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3.2 | 1.12 |  | 8 | 0.0525 | 0.0593 | 0.0525 | 0.047 | 0.0427 | 0.055 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
3.3 | 1.25 |  | 9 | 0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | 0.045 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
4.1 | 1.4 |  | 10 | 0.0371 | 0.04 | 0.0371 | 0.0347 | 0.0329 | 0.039 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
4.2 | 1.6 |  | 11 | 0.0322 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | 0.033 |  |  |  | Max | 0 | #DIV/0! | #N/A |  |  |  | Max | 0 | #DIV/0! | #N/A
4.3 | 1.8 |  | 12 | 0.0246 | 0.0269 | 0.0246 | 0.0233 | 0.022 | 0.027 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
5.1 | 2 |  | 13 | 0.0215 | 0.0232 | 0.0215 | 0.0201 | 0.019 | 0.023 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
5.2 | 2.24 |  | 14 | 0.0172 | 0.0186 | 0.0172 | 0.0161 | 0.0153 | 0.018 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
5.3 | 2.5 |  | 15 | 0.0155 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | 0.016 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
6.1 | 2.8 |  | 16 | 0.013 | 0.014 | 0.013 | 0.0122 | 0.0115 | 0.013 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
6.2 | 3.15 |  | 17 | 0.011 | 0.0119 | 0.011 | 0.0103 | 0.0097 | 0.011 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
6.3 | 3.55 |  | 18 | 0.0086 | 0.0093 | 0.0086 | 0.0081 | 0.0077 | 0.009 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
7.1 | 4 |  | 19 | 0.0066 | 0.0071 | 0.0066 | 0.0062 | 0.0059 | 0.007 |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
7.2 | 4.5 | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
7.3 | 5 | [12 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
[14 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
TO.10 kV Adj | 65 | 70 | 75 | 80 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.012 | 0.0119 | 0.011 | 0.0103 | 0.0097 |  |  |  |  | . |  |  |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.016 | 0.0168 | 0.0155 | 0.0145 | 0.0138 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.023 | 0.0232 | 0.0215 | 0.0201 | 0.019 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.032 | 0.0347 | 0.0322 | 0.0301 | 0.0285 | [8 empty] | Max | 0 | #DIV/0! | #N/A |  |  |  | Max | 0 | #DIV/0! | #N/A
0.045 | 0.0512 | 0.045 | 0.0399 | 0.036 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.066 | 0.0766 | 0.0674 | 0.0599 | 0.0542 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.086 | 0.0993 | 0.0876 | 0.078 | 0.0705 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.123 | 0.145 | 0.128 | 0.114 | 0.104 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.16 | 0.188 | 0.167 | 0.149 | 0.135 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.23 | 0.267 | 0.238 | 0.214 | 0.194 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.35 | 0.399 | 0.36 | 0.327 | 0.299 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.5 | 0.618 | 0.573 | 0.533 | 0.497 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.66 | 0.769 | 0.726 | 0.686 | 0.649 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
0.93 | 0.969 | 0.954 | 0.936 | 0.915 | [9 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
[14 empty] | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
TO.10 | Diameter (mm) | sqrt(Area) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
A | 11.1 | 9.8371 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012 |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
B | 7.9 | 7.0012 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012 |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
C | 5.6 | 4.9629 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 | 0.012 |  |  | 0 | #DIV/0! | #N/A |  |  |  |  | 0 | #DIV/0! | #N/A
D | 4 | 3.5449 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016 |  | Max | 0 | #DIV/0! | #N/A |  |  |  | Max | 0 | #DIV/0! | #N/A
E | 2.8 | 2.4814 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
F | 2 | 1.7725 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032 | 0.023 | 0.016
G | 1.4 | 1.2407 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
H | 1 | 0.88623 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
J | 0.7 | 0.62036 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066 | 0.045 | 0.032
K | 0.5 | 0.44311 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
L | 0.35 | 0.31018 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
M | 0.25 | 0.22156 | 0.93 | 0.66 | 0.5 | 0.35 | 0.23 | 0.16 | 0.123 | 0.086 | 0.066
TO.10 Values | 0 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9
A | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
B | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
C | #N/A | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155 | 0.01325 | 0.011
D | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
E | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
F | #N/A | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322 | 0.02685 | 0.0215 | 0.0185 | 0.0155
G | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
H | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
J | #N/A | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674 | 0.0562 | 0.045 | 0.0386 | 0.0322
K | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
L | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
M | #N/A | 0.954 | 0.84 | 0.726 | 0.6495 | 0.573 | 0.4665 | 0.36 | 0.299 | 0.238 | 0.2025 | 0.167 | 0.1475 | 0.128 | 0.1078 | 0.0876 | 0.0775 | 0.0674
Half Value Layer
| DHEC |  | 2: MUSC; 6: HVL Graph
| Min HVL |  | 2: HVL; 3: Min HVL; 4: Max HVL; 6: mm Al; 7: Avg Exp
kVp | (mm Al) |  | kVp | (mm Al) | (mm Al) | (mm Al) |  | TBD | TBD
30 | 0.3 |  | 60 | 2.2 | 1.8 | 2.4 |  | TBD | TBD
40 | 0.4 |  | 70 | 2.6 | 2.2 | 2.9 |  | TBD | TBD
50 | 0.5 |  | 80 | 3 | 2.6 | 3.5 |  | TBD | TBD
51 | 1.2 |  | 90 | 3.5 | 3.1 | 4 |  | TBD | TBD
60 | 1.3 |  | 100 | 3.9 | 3.5 | 4.5 |  | TBD | TBD
70 | 1.5 |  | 1: 110; 2: 4.3; 3: 3.9; 4: 5.1
71 | 2.1 |  | 1: 120; 2: 5.7; 3: 4.4; 4: 5.6
80 | 2.3 |  | 1: 130; 2: 5.2; 3: 4.8; 4: 6.1
90 | 2.5
100 | 2.7
110 | 3
120 | 3.2
130 | 3.5
140 | 3.8
150 | 4.1
Fluoro Patient Entrance Exposure Rate
Acrylic | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min
Attenuator | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail.
(cm)
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dig Acq Patient Entrance Exposure Rate
Acrylic | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min | mGy/min
Attenuator | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail. | Not Avail.
(cm)
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KAP Conversion Table
µGy·m²
cGy·cm² | 1
dGy·cm² | 10
Gy·cm² | 100
mGy·cm² | 0.01
µGy·m² | 1
Copy/SortSEE
Note on A30: Eugene Mah: Nominal contrast values as a function of kVp
Note on A46: Eugene Mah: TO.10 nominal contrast values
Note on A60: Eugene Mah: TO.10 contrast values at the kVp specified in Fluoro!V319

## DataEntry
Patient Entrance Exposure Rate (Fluoroscopy)*
Image | #REF!
Mode | Attenuator
(cm) | (cm) | kVp | mA | uGy/pls | kVp | mA | uGy/pls | kVp | mA | uGy/pls
Normal | 10
20
30
Mag 1 | 10
20
30
Mag 2 | 10
20
30
Mag 3 | 10
20
30
Mag 4 | 10
20
30
max
Patient Entrance Exposure Rate (Digital Acquisition)*
Image | #REF!
Mode | Attenuator
(cm) | (cm) | kVp | mA | uGy/pls | kVp | mA | uGy/pls | kVp | mA | uGy/pls
Normal | 10
20
30
Mag 1 | 10
20
30
Mag 2 | 10
20
30
Mag 3 | 10
20
30
Mag 4 | 10
20
30
max

## RTI_kV
kV Continuous | [6 empty] | kV Pulse
Measurements | [6 empty] | Measurements
Set kV (kV) | Set mA (mA) | Tube voltage (kV) | Exposure (mGy) | Exposure rate (mGy/min) | HVL (mm Al) | Total filtr. (mm Al) | Set kV (kV) | Set mA (mA) | Tube voltage (kV) | Exposure (mGy) | Exposure rate (mGy/min) | HVL (mm Al) | Total filtr. (mm Al) | Set frames/s (FPS) | Frames/s (FPS) | Exposure/frame (µGy/frame) | Frame exp. rate (mGy/s) | Pulse width (ms)
60 | 5 | [5 empty] | 60 | 5 | [5 empty] | 15
80 | 5 | [5 empty] | 80 | 5 | [5 empty] | 15
80 | 5 | [5 empty] | 80 | 5 | [5 empty] | 15
80 | 5 | [5 empty] | 80 | 5 | [5 empty] | 15
80 | 5 | [5 empty] | 80 | 5 | [5 empty] | 15
80 | 5 | [5 empty] | 80 | 5 | [5 empty] | 15
100 | 5 | [5 empty] | 100 | 5 | [5 empty] | 15
120 | 5 | [5 empty] | 120 | 5 | [5 empty] | 15
80 | 1 | [5 empty] | 80 | 5 | [5 empty] | 7.5
80 | 10 | [5 empty] | 80 | 5 | [5 empty] | 10
80 | 15 | [5 empty] | 80 | 5 | [5 empty] | 30
[7 empty] | 80 | 1 | [5 empty] | 15
[7 empty] | 80 | 10 | [5 empty] | 15
[7 empty] | 80 | 15 | [5 empty] | 15

## RTI_SEE
Continuous | [5 empty] | Pulse 1 | [9 empty] | Pulse 2 | [9 empty] | Digital Acq
Measurements | [5 empty] | Measurements | [9 empty] | Measurements | [9 empty] | Measurements
Dose mode | Mag mode | Thickness (cm) | kV | mA | Exposure rate (mGy/min) | Dose rate | Mag Mode | Thickness (cm) | Pulse rate | kV | mA | Exposure rate (mGy/min) | Exposure/frame (µGy/frame) | Frames/s (FPS) | Pulse width (ms) | Dose rate | Mag Mode | Thickness (cm) | Pulse rate | kV | mA | Exposure rate (mGy/min) | Exposure/frame (µGy/frame) | Frames/s (FPS) | Pulse width (ms) | Dose rate | Mag Mode | Thickness (cm) | Pulse rate | kV | mA | Exposure rate (mGy/min) | Exposure/frame (µGy/frame) | Frames/s (FPS) | Pulse width (ms)
Dose 1 | Normal | 10 |  |  |  | Dose 1 | Normal | 10 | 7.5 | [6 empty] | Dose 1 | Normal | 10 | 15 | [6 empty] | Dose 1 | Normal | 10 | 15
Dose 1 | Mag 1 | 10 |  |  |  | Dose 1 | Mag 1 | 10 | 7.5 | [6 empty] | Dose 1 | Mag 1 | 10 | 15 | [6 empty] | Dose 1 | Mag 1 | 10 | 15
Dose 1 | Mag 2 | 10 |  |  |  | Dose 1 | Mag 2 | 10 | 7.5 | [6 empty] | Dose 1 | Mag 2 | 10 | 15 | [6 empty] | Dose 1 | Mag 2 | 10 | 15
Dose 1 | Mag 3 | 10 |  |  |  | Dose 1 | Mag 3 | 10 | 7.5 | [6 empty] | Dose 1 | Mag 3 | 10 | 15 | [6 empty] | Dose 1 | Mag 3 | 10 | 15
Dose 1 | Mag 4 | 10 |  |  |  | Dose 1 | Mag 4 | 10 | 7.5 | [6 empty] | Dose 1 | Mag 4 | 10 | 15 | [6 empty] | Dose 1 | Mag 4 | 10 | 15
Dose 2 | Normal | 10 |  |  |  | Dose 2 | Normal | 10 | 15 | [6 empty] | Dose 2 | Normal | 10 | 15 | [6 empty] | Dose 2 | Normal | 10 | 15
Dose 2 | Mag 1 | 10 |  |  |  | Dose 2 | Mag 1 | 10 | 15 | [6 empty] | Dose 2 | Mag 1 | 10 | 15 | [6 empty] | Dose 2 | Mag 1 | 10 | 15
Dose 2 | Mag 2 | 10 |  |  |  | Dose 2 | Mag 2 | 10 | 15 | [6 empty] | Dose 2 | Mag 2 | 10 | 15 | [6 empty] | Dose 2 | Mag 2 | 10 | 15
Dose 2 | Mag 3 | 10 |  |  |  | Dose 2 | Mag 3 | 10 | 15 | [6 empty] | Dose 2 | Mag 3 | 10 | 15 | [6 empty] | Dose 2 | Mag 3 | 10 | 15
Dose 2 | Mag 4 | 10 |  |  |  | Dose 2 | Mag 4 | 10 | 15 | [6 empty] | Dose 2 | Mag 4 | 10 | 15 | [6 empty] | Dose 2 | Mag 4 | 10 | 15
Dose 3 | Normal | 10 |  |  |  | Dose 3 | Normal | 10 | 10 | [6 empty] | Dose 3 | Normal | 10 | 15 | [6 empty] | Dose 3 | Normal | 10 | 15
Dose 3 | Mag 1 | 10 |  |  |  | Dose 3 | Mag 1 | 10 | 10 | [6 empty] | Dose 3 | Mag 1 | 10 | 15 | [6 empty] | Dose 3 | Mag 1 | 10 | 15
Dose 3 | Mag 2 | 10 |  |  |  | Dose 3 | Mag 2 | 10 | 10 | [6 empty] | Dose 3 | Mag 2 | 10 | 15 | [6 empty] | Dose 3 | Mag 2 | 10 | 15
Dose 3 | Mag 3 | 10 |  |  |  | Dose 3 | Mag 3 | 10 | 10 | [6 empty] | Dose 3 | Mag 3 | 10 | 15 | [6 empty] | Dose 3 | Mag 3 | 10 | 15
Dose 3 | Mag 4 | 10 |  |  |  | Dose 3 | Mag 4 | 10 | 10 | [6 empty] | Dose 3 | Mag 4 | 10 | 15 | [6 empty] | Dose 3 | Mag 4 | 10 | 15
Dose 1 | Normal | 20 |  |  |  | Dose 1 | Normal | 20 | 7.5 | [6 empty] | Dose 1 | Normal | 20 | 15 | [6 empty] | Dose 1 | Normal | 20 | 15
Dose 1 | Mag 1 | 20 |  |  |  | Dose 1 | Mag 1 | 20 | 7.5 | [6 empty] | Dose 1 | Mag 1 | 20 | 15 | [6 empty] | Dose 1 | Mag 1 | 20 | 15
Dose 1 | Mag 2 | 20 |  |  |  | Dose 1 | Mag 2 | 20 | 7.5 | [6 empty] | Dose 1 | Mag 2 | 20 | 15 | [6 empty] | Dose 1 | Mag 2 | 20 | 15
Dose 1 | Mag 3 | 20 |  |  |  | Dose 1 | Mag 3 | 20 | 7.5 | [6 empty] | Dose 1 | Mag 3 | 20 | 15 | [6 empty] | Dose 1 | Mag 3 | 20 | 15
Dose 1 | Mag 4 | 20 |  |  |  | Dose 1 | Mag 4 | 20 | 7.5 | [6 empty] | Dose 1 | Mag 4 | 20 | 15 | [6 empty] | Dose 1 | Mag 4 | 20 | 15
Dose 2 | Normal | 20 |  |  |  | Dose 2 | Normal | 20 | 15 | [6 empty] | Dose 2 | Normal | 20 | 15 | [6 empty] | Dose 2 | Normal | 20 | 15
Dose 2 | Mag 1 | 20 |  |  |  | Dose 2 | Mag 1 | 20 | 15 | [6 empty] | Dose 2 | Mag 1 | 20 | 15 | [6 empty] | Dose 2 | Mag 1 | 20 | 15
Dose 2 | Mag 2 | 20 |  |  |  | Dose 2 | Mag 2 | 20 | 15 | [6 empty] | Dose 2 | Mag 2 | 20 | 15 | [6 empty] | Dose 2 | Mag 2 | 20 | 15
Dose 2 | Mag 3 | 20 |  |  |  | Dose 2 | Mag 3 | 20 | 15 | [6 empty] | Dose 2 | Mag 3 | 20 | 15 | [6 empty] | Dose 2 | Mag 3 | 20 | 15
Dose 2 | Mag 4 | 20 |  |  |  | Dose 2 | Mag 4 | 20 | 15 | [6 empty] | Dose 2 | Mag 4 | 20 | 15 | [6 empty] | Dose 2 | Mag 4 | 20 | 15
Dose 3 | Normal | 20 |  |  |  | Dose 3 | Normal | 20 | 10 | [6 empty] | Dose 3 | Normal | 20 | 15 | [6 empty] | Dose 3 | Normal | 20 | 15
Dose 3 | Mag 1 | 20 |  |  |  | Dose 3 | Mag 1 | 20 | 10 | [6 empty] | Dose 3 | Mag 1 | 20 | 15 | [6 empty] | Dose 3 | Mag 1 | 20 | 15
Dose 3 | Mag 2 | 20 |  |  |  | Dose 3 | Mag 2 | 20 | 10 | [6 empty] | Dose 3 | Mag 2 | 20 | 15 | [6 empty] | Dose 3 | Mag 2 | 20 | 15
Dose 3 | Mag 3 | 20 |  |  |  | Dose 3 | Mag 3 | 20 | 10 | [6 empty] | Dose 3 | Mag 3 | 20 | 15 | [6 empty] | Dose 3 | Mag 3 | 20 | 15
Dose 3 | Mag 4 | 20 |  |  |  | Dose 3 | Mag 4 | 20 | 10 | [6 empty] | Dose 3 | Mag 4 | 20 | 15 | [6 empty] | Dose 3 | Mag 4 | 20 | 15
Dose 1 | Normal | 30 |  |  |  | Dose 1 | Normal | 30 | 7.5 | [6 empty] | Dose 1 | Normal | 30 | 15 | [6 empty] | Dose 1 | Normal | 30 | 15
Dose 1 | Mag 1 | 30 |  |  |  | Dose 1 | Mag 1 | 30 | 7.5 | [6 empty] | Dose 1 | Mag 1 | 30 | 15 | [6 empty] | Dose 1 | Mag 1 | 30 | 15
Dose 1 | Mag 2' | 30 |  |  |  | Dose 1 | Mag 2 | 30 | 7.5 | [6 empty] | Dose 1 | Mag 2 | 30 | 15 | [6 empty] | Dose 1 | Mag 2 | 30 | 15
Dose 1 | Mag 3 | 30 |  |  |  | Dose 1 | Mag 3 | 30 | 7.5 | [6 empty] | Dose 1 | Mag 3 | 30 | 15 | [6 empty] | Dose 1 | Mag 3 | 30 | 15
Dose 1 | Mag 4 | 30 |  |  |  | Dose 1 | Mag 4 | 30 | 7.5 | [6 empty] | Dose 1 | Mag 4 | 30 | 15 | [6 empty] | Dose 1 | Mag 4 | 30 | 15
Dose 2 | Normal | 30 |  |  |  | Dose 2 | Normal | 30 | 15 | [6 empty] | Dose 2 | Normal | 30 | 15 | [6 empty] | Dose 2 | Normal | 30 | 15
Dose 2 | Mag 1 | 30 |  |  |  | Dose 2 | Mag 1 | 30 | 15 | [6 empty] | Dose 2 | Mag 1 | 30 | 15 | [6 empty] | Dose 2 | Mag 1 | 30 | 15
Dose 2 | Mag 2 | 30 |  |  |  | Dose 2 | Mag 2 | 30 | 15 | [6 empty] | Dose 2 | Mag 2 | 30 | 15 | [6 empty] | Dose 2 | Mag 2 | 30 | 15
Dose 2 | Mag 3 | 30 |  |  |  | Dose 2 | Mag 3 | 30 | 15 | [6 empty] | Dose 2 | Mag 3 | 30 | 15 | [6 empty] | Dose 2 | Mag 3 | 30 | 15
Dose 2 | Mag 4 | 30 |  |  |  | Dose 2 | Mag 4 | 30 | 15 | [6 empty] | Dose 2 | Mag 4 | 30 | 15 | [6 empty] | Dose 2 | Mag 4 | 30 | 15
Dose 3 | Normal | 30 |  |  |  | Dose 3 | Normal | 30 | 10 | [6 empty] | Dose 3 | Normal | 30 | 15 | [6 empty] | Dose 3 | Normal | 30 | 15
Dose 3 | Mag 1 | 30 |  |  |  | Dose 3 | Mag 1 | 30 | 10 | [6 empty] | Dose 3 | Mag 1 | 30 | 15 | [6 empty] | Dose 3 | Mag 1 | 30 | 15
Dose 3 | Mag 2 | 30 |  |  |  | Dose 3 | Mag 2 | 30 | 10 | [6 empty] | Dose 3 | Mag 2 | 30 | 15 | [6 empty] | Dose 3 | Mag 2 | 30 | 15
Dose 3 | Mag 3 | 30 |  |  |  | Dose 3 | Mag 3 | 30 | 10 | [6 empty] | Dose 3 | Mag 3 | 30 | 15 | [6 empty] | Dose 3 | Mag 3 | 30 | 15
Dose 3 | Mag 4 | 30 |  |  |  | Dose 3 | Mag 4 | 30 | 10 | [6 empty] | Dose 3 | Mag 4 | 30 | 15 | [6 empty] | Dose 3 | Mag 4 | 30 | 15
Dose 1 | Normal | Max |  |  |  | Dose 1 | Normal | Max | 7.5 | [6 empty] | Dose 1 | Normal | Max | 15
Dose 2 | Normal | Max |  |  |  | Dose 2 | Normal | Max | 15 | [6 empty] | Dose 2 | Normal | Max | 15
Dose 3 | Normal | Max |  |  |  | Dose 3 | Normal | Max | 10 | [6 empty] | Dose 3 | Normal | Max | 15

## RTI_HVL
HVL
Measurements
Set kV (kV) | Set mA (mA) | mm Al | Exposure rate (mGy/min)
80
80
80
80
80
80

## RTI_AKAccuracy
Air Kerma Accuracy |  |  |  | KAP Accuracy
Measurements |  |  |  | Measurements
Set kV (kV) | Set mA (mA) | Indicated exposure rate (mGy/min) | Exposure rate (mGy/min) | Set kV (kV) | Set mA (mA) | Start KAP | End KAP | Indicated KAP | KAP exposure (mGycm²) | KAP exposure rate (mGycm²/s) | KAP exposure diff %
60 |  |  |  | 60
80 |  |  |  | 80
100 |  |  |  | 100
120 |  |  |  | 120

## RTI_Scatter
Measurements
Set text | Exposure rate (µGy/h)
Fluoro 1 Unshielded Eye
Fluoro 1 Unshielded Waist
Fluoro 2 Unshielded Eye
Fluoro 2 Unshielded Waist
Fluoro 1 Shielded Eye
Fluoro 1 Shielded Waist
Fluoro 2 Shielded Eye
Fluoro 2 Shielded Waist
Receptor Transmission

## RTI_ReceptorEER
Continuous |  |  |  |  | Pulse 1 | [7 empty] | Pulse 2 | [7 empty] | Digital Acq
Measurements |  |  |  |  | Measurements | [7 empty] | Measurements | [7 empty] | Measurements
Dose mode | Mag mode | kV | mA | Exposure rate (µGy/min) | Dose mode | Mag mode | Set frames/s (FPS) | kV | mA | Exposure rate (µGy/min) | Exposure/frame (nGy/frame) | Frames/s (FPS) | Dose mode | Mag mode | Set frames/s (FPS) | kV | mA | Exposure rate (µGy/min) | Exposure/frame (nGy/frame) | Frames/s (FPS) | Dose mode | Mag mode | Set frames/s (FPS) | kV | mA | Exposure rate (µGy/min) | Exposure/frame (nGy/frame) | Frames/s (FPS)
Dose 1 | Normal |  |  |  | Dose 1 | Normal | 15 | [5 empty] | Dose 1 | Normal | 15 | [5 empty] | Dose 1 | Normal | 15
Dose 1 | Mag 1 |  |  |  | Dose 1 | Mag 1 | 15 | [5 empty] | Dose 1 | Mag 1 | 15 | [5 empty] | Dose 1 | Mag 1 | 15
Dose 1 | Mag 2 |  |  |  | Dose 1 | Mag 2 | 15 | [5 empty] | Dose 1 | Mag 2 | 15 | [5 empty] | Dose 1 | Mag 2 | 15
Dose 1 | Mag 3 |  |  |  | Dose 1 | Mag 3 | 15 | [5 empty] | Dose 1 | Mag 3 | 15 | [5 empty] | Dose 1 | Mag 3 | 15
Dose 1 | Mag 4 |  |  |  | Dose 1 | Mag 4 | 15 | [5 empty] | Dose 1 | Mag 4 | 15 | [5 empty] | Dose 1 | Mag 4 | 15
Dose 2 | Normal |  |  |  | Dose 2 | Normal | 15 | [5 empty] | Dose 2 | Normal | 15 | [5 empty] | Dose 2 | Normal | 15
Dose 2 | Mag 1 |  |  |  | Dose 2 | Mag 1 | 15 | [5 empty] | Dose 2 | Mag 1 | 15 | [5 empty] | Dose 2 | Mag 1 | 15
Dose 2 | Mag 2 |  |  |  | Dose 2 | Mag 2 | 15 | [5 empty] | Dose 2 | Mag 2 | 15 | [5 empty] | Dose 2 | Mag 2 | 15
Dose 2 | Mag 3 |  |  |  | Dose 2 | Mag 3 | 15 | [5 empty] | Dose 2 | Mag 3 | 15 | [5 empty] | Dose 2 | Mag 3 | 15
Dose 2 | Mag 4 |  |  |  | Dose 2 | Mag 4 | 15 | [5 empty] | Dose 2 | Mag 4 | 15 | [5 empty] | Dose 2 | Mag 4 | 15
Dose 3 | Normal |  |  |  | Does 3 | Normal | 15 | [5 empty] | Does 3 | Normal | 15 | [5 empty] | Does 3 | Normal | 15
Dose 3 | Mag 1 |  |  |  | Does 3 | Mag 1 | 15 | [5 empty] | Does 3 | Mag 1 | 15 | [5 empty] | Does 3 | Mag 1 | 15
Dose 3 | Mag 2 |  |  |  | Does 3 | Mag 2 | 15 | [5 empty] | Does 3 | Mag 2 | 15 | [5 empty] | Does 3 | Mag 2 | 15
Dose 3 | Mag 3 |  |  |  | Does 3 | Mag 3 | 15 | [5 empty] | Does 3 | Mag 3 | 15 | [5 empty] | Does 3 | Mag 3 | 15
Dose 3 | Mag 4 |  |  |  | Does 3 | Mag 4 | 15 | [5 empty] | Does 3 | Mag 4 | 15 | [5 empty] | Does 3 | Mag 4 | 15

## DataPage
Sheet | FLUORO
SurveyID
HVL | 80 | TBD
Field size
Dose mode | Auto 1
Auto
SEE | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Max SEE | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dose mode | Auto 1
Auto
SEE | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Max SEE | max | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
IEER |  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
IEER |  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0
TO.N3
TO.10 | Field size | A | B | C | D | E | F | G | H | J | K | L | M
Resolution